## E05
| 機台基本資料 Basic Infn Of Tester | [8 empty] | 輪班工程師處理說明 Shifts Engineer disposition Explain
維修單編號 EQPRepair No. | 機台名稱 Machine Name | 作業批號 Runcard | 異常類型 Abnormal type | Probe Card 編號 Probe Card No. | 產品型號 Dveice No. | 產品批號 Lot No. | 作業刻號 Wafer ID | 異常現象 Abnormal status | 開單人員 Operator | 異常分類 Abnormal Type | 異常原因 Abnormal Reason | 處理開始時間 Process Start Time | 處理結束時間 Process End Time | 處理時間 Handle time | 處理結果 Result | 處理人員 Operator | 處理過程說明 Disposition Explain | Fuse Mode 是否啟閉 | Override 是否啟閉 | Function Key是否啟閉 | Temperature 是否啟閉 | PGM是否正確 | H/W Damage硬體是否損壞
EQPRepairJ22103341 | T6373-38 | CWC7J0A61-01.001 | OverKill | JB1GKS09 | PA6543B | J67H5 | 04 | TOO MANY O/S(overkill done apply) | A6296 | Over Kill | O/S誤宰 | 2022/08/01 02:46:44 | 2022/08/01 02:56:54 | 0.17 | OK | 17542 | 機台清針，起測一段PASS | NO | NO | NO | YES | YES | NO
EQPRepairJ22103361 | T6373-15 | CWC7J0L28-01 | OverKill | JB6K8S01 | NC01682B(COG48AA)-A-AT-N-C | GGC155.1-21B2C | 07 | TOO MANY OS.. NO PARTICLE-A6925 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/01 04:19:47 | 2022/08/01 04:36:26 | 0.28 | OK | 17542 | Remote處理 ----- 機台清針，調整針位，試打(166,2)無異常，點測PASS穩定，還機待觀察 | NO | NO | YES | YES | YES | NO
EQPRepairJ22103375 | T6373-135 | CWC7J0812-01 | OverKill | JB1UJS03 | PA5129LS_12U80H-LT | P062Y.1 | 09 | PARTICLES IN NEEDLE(coz too many O/S) | A6296 | Over Kill | O/S誤宰 | 2022/08/01 05:24:02 | 2022/08/01 06:35:15 | 1.19 | OK | 17674 | 1. 下卡磨針pass，溫針range y=10 (246,8)#05 | NO | NO | NO | YES | YES | NO
EQPRepairJ22103385 | T6373-70 | CWC7J0H52-01.001 | OverKill | JB1TAS28 | PA5208KF-LT | J696J | 06 | TOO MANY O/S(overkill done apply) | A6296 | Over Kill | O/S | 2022/08/01 06:00:29 | 2022/08/01 06:18:51 | 0.31 | OK | 15096 | 清針、調整針位pass. | NO | NO | NO | YES | YES | NO
EQPRepairJ22103395 | T6373-135 |  | OverKill | JB1UJS03 | PA5129LS_12U80H-LT | P062Y.1 |  | os | 17674 | Over Kill | O/S誤宰 | 2022/08/01 06:35:57 | 2022/08/01 06:44:10 | 0.14 | OK | 17674 | 1. 更換備卡#14 PASS，溫針 RANGE Y=10 (41,1)#09 | NO | NO | NO | YES | YES | NO
EQPRepairJ22103406 | T6373-43 | CWC6J0527-01.004 | OverKill | JB1UJS13 | PA5129LS_12U80H-LT | P049S.1 | 17 | O/S TOO MANY-A6916 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/01 07:59:55 | 2022/08/01 08:19:53 | 0.33 | OK | 07910 | 機台清針 pass p/m(141.4)#17 RY12 | NO | NO | NO | YES | YES | NO
EQPRepairJ22103408 | T6373-38 | CWC7J0A61-01.001 | OverKill | JB1GKS09 | PA6543B | J67H5 | 07 | TOO MANY O/S-A6895 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/01 08:01:37 | 2022/08/01 08:25:56 | 0.41 | OK | 08774 | 毛刷清針131.7 | NO | NO | NO | YES | YES | NO
EQPRepairJ22103414 | T6373-111 | CWC7J0A56-01 | OverKill | TB1R8S67 | PA5137F-LT1 | P094K.1 | 16 | CONTINOUS EDA 2H , O/S | A6897 | Over Kill | 2H | 2022/08/01 08:28:04 | 2022/08/01 08:46:52 | 0.31 | OK | 15897 | 少量O/S.清針調針.Range:8um.#16(17.5).續測. | NO | NO | NO | YES | YES | NO
EQPRepairJ22103428 | T6373-29 |  | Prober Down | KB2BNS09 | ND01733G(COF94AA)-A-C | H2UM03.1-2BC2 |  | AlarmCode:1057 AlarmMessage:MODEL NOT FOUND ( ROUGH ALIGNMENT ) !! | CPAUTO | Prober | Wafer Edge/Mode Not Found | 2022/08/01 09:35:34 | 2022/08/01 09:45:18 | 0.16 | OK | 17712 | 重抓ok，check die ok，#16(38.3)rg20/14 | NO | NO | NO | YES | YES | NO
EQPRepairJ22103453 | T6373-43 | CWC6J0527-01.004 | OverKill | JB1UJS13 | PA5129LS_12U80H-LT | P049S.1 | 18 | Already follow Overkilll handling procedures but still not effective too many o/s | A6897 | Over Kill | O/S誤宰 | 2022/08/01 11:00:50 | 2022/08/01 11:11:26 | 0.18 | OK | 20231 | 1.清針起測一段穩定pass，PM(138,1) | NO | NO | NO | YES | YES | NO
EQPRepairJ22103454 | T6373-34 | CWC7J0829-01 | OverKill | TB1R8S76 | PA5137F-LT1 | P07NW.1 | 20 | O/S-A6571 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/01 11:05:28 | 2022/08/01 11:24:21 | 0.31 | OK | 17712 | 機台清針pass，#20(119.6)rg9 | NO | NO | NO | YES | YES | NO
EQPRepairJ22103455 | T6373-38 | CWC7J0A61-01.001 | OverKill | JB1GKS09 | PA6543B | J67H5 | 08 | o/s-a6571 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/01 11:06:29 | 2022/08/01 11:32:33 | 0.43 | OK | 08774 | 1.執行soaktime 重新對針 機台清針 2.pm(4.2) | NO | NO | NO | YES | YES | NO
EQPRepairJ22103465 | T6373-29 | CWC7J0J48-01.001 | OverKill | KB2BNS09 | ND01733G(COF94AA)-A-C | H2UM03.1-2BC2 | 17 | 2H no corr-A6894 | CPAUTO | Over Kill | 2H | 2022/08/01 11:34:03 | 2022/08/01 11:48:03 | 0.23 | OK | 17712 | [CUSTOMER_HOLD_CRITERIA] 請續測 [HW_Bin33+34>55EA(3%)]134(HW_Bin33+34>55EA(3%), Please HOLD(不需重測，上傳異常單請附datalog及WV(FK=1+2+7))) | NO | NO | NO | YES | YES | NO
EQPRepairJ22103470 | T6373-70 | CWC7J0H52-01.001 | OverKill | JB1TAS28 | PA5208KF-LT | J696J | 09 | too many o/s-A6894 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/01 11:53:29 | 2022/08/01 12:45:47 | 0.87 | OK | 20231 | 1.清調針PASS，pm(148,1) | NO | NO | NO | YES | YES | NO
EQPRepairJ22103487 | T6373-151 | CWC7J0L75-01 | OverKill | JB2P2S08 | R63455-A1S-EZ-CP | NAHM75.00-B1C | 09 | 2H DO NOT RUN CORR DEVICE-A6916 | CPAUTO | Over Kill | 2H | 2022/08/01 13:02:41 | 2022/08/01 13:10:35 | 0.13 | OK | 20231 | [HW_BIN26>0.5%]8----- 1.查看EDA資料，不同貨批(NAHM85.00)不同機卡有相同問題，續測再觀察。 | NO | NO | NO | YES | YES | NO
EQPRepairJ22103503 | T6373-39 | CWC7J0F95-01 | OverKill | JB1TAS21 | PA5208KF2 | J68G6 | 03 | too many o/s, please checky have particles-A6986 | CPAUTO | Over Kill | O/D不足 | 2022/08/01 14:10:23 | 2022/08/01 14:21:48 | 0.19 | OK | 20231 | 1.清針起測一段穩定PASS，pm(47,0) | NO | NO | NO | YES | YES | NO
EQPRepairJ22103514 | T6373-70 | CWC7J0H52-01.001 | OverKill | JB1TAS28 | PA5208KF-LT | J696J | 10 | CONTINUOUS EDA 2R O/S (CORR.PASSED) | A6897 | Over Kill | 2R | 2022/08/01 15:23:24 | 2022/08/01 15:56:44 | 0.56 | OK | 20231 | [HW_BIN1>1%](Bin 2 > 8EA) 12(hold wafer)----- 1.查看EDA資料，#08 MAP偏左上零星BIN2，#09 MAP右上區域BIN2 2.重新對針，清調針，rg:8，起測穩定pass，PM(147,1) | NO | NO | NO | YES | YES | NO
EQPRepairJ22103525 | T6373-39 | CWC7J0F95-01 | OverKill | JB1TAS21 | PA5208KF2 | J68G6 | 04 | o/s處裡無效 | 07079 | Over Kill | O/S誤宰 | 2022/08/01 16:22:12 | 2022/08/01 16:39:22 | 0.29 | OK | 07910 | 機台清針 2分法PASS P/M(147.5)#4 RY14 | NO | NO | NO | YES | YES | NO
EQPRepairJ22103543 | T6373-38 | CWC7J0A61-01.001 | OverKill | JB1GKS09 | PA6543B | J67H5 | 11 | o/s-20022 | CPAUTO | Over Kill | 配件殘金 | 2022/08/01 18:00:47 | 2022/08/01 18:42:36 | 0.70 | OK | 17712 | S09殘金更換備卡S12 PASS，#11(47.7)rg11 | NO | NO | NO | YES | YES | NO
EQPRepairJ22103560 | T6373-128 | CWC7J0M33-01 | Corr Fail | VB1BVS01 | PA5270D-LT | U38F8.00 | 05 | Corr Fail-A6148 | CPAUTO | Corr | Other | 2022/08/01 20:14:06 | 2022/08/01 20:18:26 | 0.07 | OK | 03426 | Remote 處理----T32 產品無 Corr Wafer Run Corr 比對 Fail | NO | NO | NO | YES | YES | NO
EQPRepairJ22103580 | T6373-42 | CWC7J0L30-01.004 | OverKill | JB1GVS22 | NC01682B(COG48AA)-A-AT-N-C | GGG498.1-BC112 | 16 | Too many fail Bin29 pm checked(56.4)-A6525 | CPAUTO | Over Kill | Function誤宰 | 2022/08/01 22:05:03 | 2022/08/01 23:01:16 | 0.94 | OK | 13597 | 更換配件#26，RANGE 10 UM | NO | NO | NO | YES | YES | NO
EQPRepairJ22103588 | T6373-152 | CWC7J0H67-01 | OverKill | JB2P2S10 | R63455-A1S-EZ-CP | NAHM72.00-B1C | 03 | #1.2 EDA SHOW 2R O/S-A5717 | CPAUTO | Over Kill | Show R | 2022/08/01 23:09:49 | 2022/08/02 00:01:04 | 0.85 | OK | 17542 | Show R，O/S不穩，更換配件#09，上機PASS，溫針重動針，range:20，試打(56,8)無異常，起測pass穩定 | NO | NO | NO | YES | YES | NO
EQPRepairJ22103592 | T6373-34 | CWC7J0J51-01.001 | OverKill | TB1R8S76 | PA5137F-LT1 | P094H.1 | 02 | Too many O/S Bin3 pm checked(120.5) od12-A6525 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/01 23:17:21 | 2022/08/02 00:14:52 | 0.96 | OK | 17542 | Remote處理 ----- 機台清針，OD:10，起測一段pass | NO | NO | NO | YES | YES | NO
EQPRepairJ22103594 | T6373-128 | CWC7J0M33-01 | OverKill | VB1BVS01 | PA5270D-LT | U38F8.00 | 06 | eda show 2H wait for pass wafer-A6924 | CPAUTO | Over Kill | show h | 2022/08/01 23:18:45 | 2022/08/01 23:42:32 | 0.40 | OK | 15965 | [HW_BIN60<100%](Bin 61 < 798EA) 717 查看EDA 不同機有相同問題續測。 | NO | NO | NO | YES | NO | NO
EQPRepairJ22103595 | T6373-151 | CWC7J0L75-01 | OverKill | JB2P2S08 | R63455-A1S-EZ-CP | NAHM75.00-B1C | 13 | #11.12 EDA/2H (DO NOT RUN CORR)-A6925 | CPAUTO | Over Kill | SHOW H | 2022/08/01 23:19:40 | 2022/08/01 23:47:23 | 0.46 | OK | 15965 | [HW_BIN26>0.5%]7 不同機台有相同問題續測 | NO | NO | NO | YES | NO | NO
EQPRepairJ22103610 | T6373-37 | CWC7J0L27-01.003 | Prober Down | JB1GVS21 | NC01682A(COG48AA)-A-AT-N2-C | GGS286.1-BC212 | 12 | PUSH <DATA IN> SW AT WAFER THICKNESS POSITIONS! | 19697 | Prober | Wafer Edge/Mode Not Found | 2022/08/02 00:16:40 | 2022/08/02 00:27:45 | 0.18 | OK | 03426 | Remote---DATA IN OK | NO | NO | NO | YES | YES | NO
EQPRepairJ22103614 | T6373-12 | CWC8J0010-01 | OverKill | JBCK3S01 | NB01853E(COG48AA)-A-AT-X-C | GGR107.1-1B1C | 01 | show I NO CORR | A6899 | Over Kill | show I | 2022/08/02 00:42:34 | 2022/08/02 00:46:31 | 0.07 | OK | 15096 | #01單片 show I，續測觀察. | NO | NO | NO | YES | YES | NO
EQPRepairJ22103620 | T6373-136 | CWC7J0L29-01.004 | OverKill | JB1GVS23 | NC01682B(COG48AA)-A-AT-N-C | GGH464.1-BC112 | 24 | There was a particle | A6525 | Over Kill | 機台清針後Pass | 2022/08/02 01:05:20 | 2022/08/02 01:34:43 | 0.49 | OK | 13510 | 機台清針後Pass | NO | NO | NO | YES | YES | NO
EQPRepairJ22103633 | T6373-111 | CWC7J0A56-01 | OverKill | TB1R8S67 | PA5137F-LT1 | P094K.1 | 25 | Too many fail Bin9 and O/S Bin3 pm checked(87.3) od13-A6525 | CPAUTO | Over Kill | O/S | 2022/08/02 01:45:22 | 2022/08/02 02:38:48 | 0.89 | OK | 15096 | 機台清針、調整針位PASS. | NO | NO | NO | YES | YES | NO
EQPRepairJ22103641 | T6373-13 | CWC7J0143-01 | Setup Fail | JB8Q6S04 | NB01761A(COF94AA)-A-C | H2NC82.1-B | 07 | 0111 #7 id passed, set up fail, no other p/c, too many o/s, probe mark no good, p/c polish in 6th floor | A5548 | Setup Fail | O/S Fail | 2022/08/02 02:05:56 | 2022/08/02 03:11:32 | 1.09 | OK | 15096 | Setup Fail-----o/s，下卡磨針，上機pass，setup 04 | NO | NO | NO | YES | YES | NO
EQPRepairJ22103643 | T6373-12 | CWC8J0010-01 | OverKill | JBCK3S01 | NB01853E(COG48AA)-A-AT-X-C | GGR107.1-1B1C | 02 | O/S TOO MANY-19697 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/02 02:19:27 | 2022/08/02 02:44:30 | 0.42 | OK | 13510 | 1.重新對針PASS | NO | NO | NO | YES | YES | NO
EQPRepairJ22103654 | T6373-103 | CWC7J0K75-01.002 | OverKill | JB1GVS17 | NC01682A(COG48AA)-A-AT-N3-C | GGS287.1-B1C22 | 05 | too many o/s 18.5 pm check pc has particle-A6881 | CPAUTO | Over Kill | Function誤宰 | 2022/08/02 03:01:23 | 2022/08/02 04:04:34 | 1.05 | OK | 03426 | 更換配件S15 等待20分鐘 降溫上機PASS 17.1 169.1 | NO | NO | NO | YES | YES | NO
EQPRepairJ22103658 | T6373-128 | CWC7J0M33-01 | OverKill | VB1BVS01 | PA5270D-LT | U38F8.00 | 07 | O/S TOO MANY-19697 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/02 03:20:46 | 2022/08/02 03:37:31 | 0.28 | OK | 03426 | 針位調整 機台清針PASS | NO | NO | NO | YES | YES | NO
EQPRepairJ22103663 | T6373-70 | CWC6J0N19-01.001 | OverKill | JB1TAS28 | PA5208KF-LT | J66H6.10 | 11 | There was a particle | A6525 | Over Kill | 配件有異物 | 2022/08/02 03:54:52 | 2022/08/02 05:02:06 | 1.12 | OK | 17674 | 1. 配件有異物 下卡磨針PASS，溫針重對針再溫針，#11 (83,3) RANGE Y=10 (43,6)，ABN 單 | NO | NO | NO | YES | YES | NO
EQPRepairJ22103666 | T6373-38 | CWC7J0A61-01 | OverKill | JB1GKS12 | PA6543B | J67H5 | 15 | TOO MANY OS.. NO PARTICLE-A6925 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/02 04:02:21 | 2022/08/02 04:29:53 | 0.46 | OK | 03426 | 更換配件S10 | NO | NO | NO | YES | YES | NO
EQPRepairJ22103671 | T6373-86 | CWC7J0J58-01 | OverKill | JB1Q4S16 | PA6542C | J683L | 15 | Too many O/S Bin2.3 pm checked(260.2) od15-A6525 | CPAUTO | Over Kill | O/S | 2022/08/02 04:32:21 | 2022/08/02 04:53:57 | 0.36 | OK | 15096 | 清針 調針 pass. | NO | NO | NO | YES | YES | NO
EQPRepairJ22103678 | T6373-43 | CWC7J0830-01.002 | OverKill | JB1UJS13 | PA5129LS_12U80H-LT | P065S.1 | 04 | Too many O/S Bin32 pm checked(21.4) od15-A6525 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/02 04:59:14 | 2022/08/02 05:14:33 | 0.26 | OK | 17674 | 1. 重對針、清針PASS， RANGE Y=10 (24,4)#04 | NO | NO | NO | YES | YES | NO
EQPRepairJ22103683 | T6373-152 | CWC7J0H67-01 | OverKill | JB2P2S09 | R63455-A1S-EZ-CP | NAHM72.00-B1C | 05 | Continuous 2 wafers EDA shows H No corr. wafer-A6525 | CPAUTO | Over Kill | show h | 2022/08/02 05:20:37 | 2022/08/02 05:34:11 | 0.23 | OK | 15965 | [HW_BIN26>0.5%]12 先前show r已有更換備卡，BIN26 查看EDA 不同機卡有相同BIN 續測。 | NO | NO | NO | YES | YES | NO
EQPRepairJ22103686 | T6373-52 | CWC7J0K72-01 | OverKill | TB1D3S04 | R63319C0EQCPQV-P | AYP075000-B11C | 04 | too many o/s-A6497 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/02 05:25:57 | 2022/08/02 05:49:08 | 0.39 | OK | 17542 | 使用Clean Wafer 清針，起測一段PASS穩定 | NO | NO | NO | YES | YES | NO
EQPRepairJ22103689 | T6373-111 | CWC7J0H50-01 | OverKill | TB1R8S67 | PA5137F-LT1 | P097F.1 | 01 | too many o/s,already overkill not effective | A6404 | Over Kill | 針位調整 | 2022/08/02 05:33:31 | 2022/08/02 05:44:27 | 0.18 | OK | 03426 | 針位調整PASS | NO | NO | NO | YES | YES | NO
EQPRepairJ22103704 | T6373-13 | CWC7J0143-01 | OverKill | JB8Q6S04 | NB01761A(COF94AA)-A-C | H2NC82.1-BC1111 | 07 | PROBEMARK NG (66.8)ALL SITES,OD5UM,TOO MANY O/S-A5717 | CPAUTO | Over Kill | Other | 2022/08/02 06:21:57 | 2022/08/02 06:34:10 | 0.20 | OK | 13597 | 1.此批為回測批，BIN29 SHORT P-P、F-F產品問題 2.試打針點(172.8)正常，續測 | NO | NO | NO | YES | YES | NO
EQPRepairJ22103713 | T6373-52 | CWC7J0K72-01 | OverKill | TB1D3S04 | R63319C0EQCPQV-P | AYP075000-B11C | 05 | #3,4 CONTINUES EDA SHOW R NO CORROLATION WAFER PM CHECK 256.6 OD15-A7024 | CPAUTO | Over Kill | 2R | 2022/08/02 07:03:09 | 2022/08/02 07:33:43 | 0.51 | OK | 07910 | 2R，查EDA BIN10 分佈，機台清針、檢查針況針痕、2分法調針 PASS P/M(257.1)#5 | NO | NO | NO | YES | YES | NO
EQPRepairJ22103722 | T6373-29 | CWC7J0J48-01.001 | OverKill | KB2BNS09 | ND01733G(COF94AA)-A-C | H2UM03.1-2BC2 | 21 | cont. 2H; NO CORR-a6962 | CPAUTO | Over Kill | 2H | 2022/08/02 07:34:52 | 2022/08/02 07:39:47 | 0.08 | OK | 08774 | 2H 1.[CUSTOMER_HOLD_CRITERIA] [HW_Bin33+34>55EA(3%)]138(HW_Bin33+34>55EA(3%), Please HOLD(不需重測，上傳異常單請附datalog及WV(FK=1+2+7))) 2.請產線換批 | NO | NO | NO | YES | YES | NO
EQPRepairJ22103723 | T6373-128 | CWC7J0M33-01 | OverKill | VB1BVS01 | PA5270D-LT | U38F8.00 | 08 | #6.7 2H-20022 | CPAUTO | Over Kill | 2H | 2022/08/02 07:38:51 | 2022/08/02 07:47:20 | 0.14 | OK | 07910 | 2H，PASS BIN 卡關 [HW_BIN60<100%](Bin 61 < 798EA) 706 ，請續測 | NO | NO | NO | YES | YES | NO
EQPRepairJ22103731 | T6373-34 | CWC7J0J51-01.001 | OverKill | TB1R8S76 | PA5137F-LT1 | P094H.1 | 06 | CHECK O/S-A6894 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/02 08:10:03 | 2022/08/02 08:25:49 | 0.26 | OK | 17712 | 機台清針，確認針痕針壓，#6(43.6)RG9 | NO | NO | NO | YES | YES | NO
EQPRepairJ22103732 | T6373-135 | CWC7J0811-01 | OverKill | JB1UJS14 | PA5129LS_12U80H-LT | P060L.1 | 04 | #02,#03 EDA SHOW 2H-20021 | CPAUTO | Over Kill | 2H | 2022/08/02 08:13:42 | 2022/08/02 08:17:52 | 0.07 | OK | 17680 | remote----[HW_BIN60<95%](Bin 61 < 675EA) 673,keep test | NO | NO | NO | YES | YES | NO
EQPRepairJ22103759 | T6373-86 | CWC7J0J58-01 | OverKill | JB1Q4S16 | PA6542C | J683L | 17 | too mnay o/s-A6894 | CPAUTO | Over Kill | O/D不足 | 2022/08/02 09:31:11 | 2022/08/02 09:40:17 | 0.15 | OK | 13749 | Remote--- 1.針壓跑掉 清針重對針 PASS (43.5) Slot #17 | NO | NO | NO | YES | YES | NO
EQPRepairJ22103762 | T6373-111 | CWC7J0H50-01 | OverKill | TB1R8S67 | PA5137F-LT1 | P097F.1 | 03 | 2H O/S-a6571 | CPAUTO | Over Kill | 2H | 2022/08/02 09:37:50 | 2022/08/02 10:37:20 | 0.99 | OK | 19771 | 2H 更換備卡,上機固定pin,重mnt多次pass #3(43.6) Range 10 | NO | NO | NO | YES | YES | NO
EQPRepairJ22103765 | T6373-70 |  | Prober Down | TB1R8S85 | PA5223Q_12U60H-LT | P06NP.1 |  | AlarmCode:1060 AlarmMessage:CANNOT MEASURE INDEX SIZE !! | CPAUTO | Prober | Wafer Edge/Mode Not Found | 2022/08/02 09:45:42 | 2022/08/02 09:53:00 | 0.12 | OK | 16631 | 1.重抓位置/check die/針痕 | NO | NO | NO | YES | YES | NO
EQPRepairJ22103769 | T6373-29 | CWC7J0J48-01.001 | OverKill | KB2BNS09 | ND01733G(COF94AA)-A-C | H2UM03.1-2BC2 | 22 | too many o.s-A6894 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/02 10:14:25 | 2022/08/02 10:49:26 | 0.58 | OK | 08774 | Remote 處理----接手 1.執行soaktime 重新對針 機台清針 | NO | NO | NO | YES | YES | NO
EQPRepairJ22103784 | T6373-37 | CWC7J0L27-01.004 | Tester Down | JB1GVS21 | NC01682A(COG48AA)-A-AT-N2-C | GGS286.1-BC12 | 16 | AlarmCode:669 AlarmMessage:GP-IB COMMUNICATION TIMEOUT ERROR !! | CPAUTO | Tester | 離線(OCT/GPIB error) | 2022/08/02 11:41:36 | 2022/08/02 11:53:22 | 0.20 | OK | 17712 | T44型號產線切狀態導致機台Alarm，切Run續測 | NO | NO | NO | YES | YES | NO
EQPRepairJ22103785 | T6373-37 | CWC7J0L27-01.004 | Tester Down | JB1GVS21 | NC01682A(COG48AA)-A-AT-N2-C | GGS286.1-BC12 | 16 | AlarmCode:661 AlarmMessage:GP-IB COMMAND EXECUTION ERROR !! | CPAUTO | Tester | 離線(OCT/GPIB error) | 2022/08/02 11:58:23 | 2022/08/02 12:02:15 | 0.06 | OK | 17712 | GPIB，STDF型號，#16退片重新上貨 | NO | NO | NO | YES | YES | NO
EQPRepairJ22103792 | T6373-39 | CWC7J0M27-01 | OverKill | JB1BYS10 | PA6587B_15U90H | J696G | 04 | #02 EDA SHOW I(o/s)-20021 | CPAUTO | Over Kill | SHOW I | 2022/08/02 12:32:49 | 2022/08/02 12:37:46 | 0.08 | OK | 07910 | SHOW I ，續測 | NO | NO | NO | YES | YES | NO
EQPRepairJ22103798 | T6373-29 | CWC7J0J48-01.001 | OverKill | KB2BNS09 | ND01733G(COF94AA)-A-C | H2UM03.1-2BC2 | 23 | too many fail bin 33-A6894 | CPAUTO | Over Kill | Function誤宰 | 2022/08/02 12:54:47 | 2022/08/02 13:05:04 | 0.17 | OK | 17680 | remote----此批與其他批使用不同卡皆無法回收bin33,keep test | NO | NO | NO | YES | YES | NO
EQPRepairJ22103809 | T6373-52 | CWC7J0K72-01 | OverKill | TB1D3S04 | R63319C0EQCPQV-P | AYP075000-B11C | 10 | CONTINUOUS EDA SHOW 2R | A6897 | Over Kill | 2R | 2022/08/02 13:53:09 | 2022/08/02 14:13:18 | 0.34 | OK | 07910 | 2R，機台清針 ，檢查無卡毛異物黑點，機台清針 PASS P/M(280.5)#10 | NO | NO | NO | YES | YES | NO
EQPRepairJ22103814 | T6373-38 | CWC7J0A61-01 | OverKill | JB1GKS10 | PA6543B | J67H5 | 20 | Already follow Overkilll handling procedures but still not effective too many O/S. | A6897 | Over Kill | O/S誤宰 | 2022/08/02 14:03:47 | 2022/08/02 14:42:14 | 0.64 | OK | 20352 | 1.執行soaktime重新對針 2.二分法確認針痕與五點水平 YRG=6 P=(5.2) | NO | NO | NO | YES | YES | NO
EQPRepairJ22103819 | T6373-128 | CWC7J0M33-01 | Tester Down | VB1BVS01 | PA5270D-LT | U38F8.00 | 10 | AlarmCode:669 AlarmMessage:GP-IB COMMUNICATION TIMEOUT ERROR !! | CPAUTO | Product Issue | RTE | 2022/08/02 14:27:21 | 2022/08/02 16:07:43 | 1.67 | OK | 16631 | 1.#10(110.1)RTE，重新連線/退針壓/重LOAD程式無效，TESTER SHUTDOWN OK，起測BIN31 LEAKAGE過多下卡烘烤，上機PASS 171.3，#10 LOAD AGAIN | NO | NO | NO | YES | YES | NO
EQPRepairJ22103823 | T6373-21 | CWC7J0J54-01 | Tester Down | JB1TAS26 | PA5208KF-LT | J696H | 04 | online initialization-A6962 | CPAUTO | Tester | Auto Cal / Online Init / ABS Fail | 2022/08/02 14:38:47 | 2022/08/02 14:55:50 | 0.28 | OK | 17712 | ok | NO | NO | NO | YES | YES | NO
EQPRepairJ22103836 | T6373-121 | CWC6J0M58-01.006 | Prober Down | JB1QNS04 | NA01687G(COG48BA)-A-AT-N2-C | GGS269.1-BC22112 | 08 | CANNOT MEASURE INDEX SIZE-A6921 | CPAUTO | Prober | Wafer Edge/Mode Not Found | 2022/08/02 15:31:07 | 2022/08/02 15:41:37 | 0.18 | OK | 20231 | 1.重抓model OK 2.check die OK 3.PM(144,0) | NO | NO | NO | YES | YES | NO
EQPRepairJ22103888 | T6373-147 | CWC7J0L15-01 | OverKill | TB1R7S31 | PA5137E1-LT | P097C.1 | 15 | CONTINUOUS EDA SHOW 2H O/S-A6881 | CPAUTO | Over Kill | SHOW H | 2022/08/02 19:34:33 | 2022/08/02 20:53:49 | 1.32 | OK | 13597 | 下卡磨針，RANGE 10 UM | NO | NO | NO | YES | YES | NO
EQPRepairJ22103896 | T6373-128 | CWC7J0M33-01 | OverKill | VB1BVS01 | PA5270D-LT | U38F8.00 | 11 | #9,10 Continuous 2 wafers EDA shows "H" NO CORROLATION WAFER PM CHECK 155.4 OD5 | A7024 | Over Kill | Show H | 2022/08/02 19:57:52 | 2022/08/02 20:00:39 | 0.05 | OK | 17542 | Remote處理 ----- Show H，[HW_BIN60<100%](Bin 61 < 798EA) 688，不同批有相同問題(U38F5.00)，續測 | NO | NO | NO | YES | YES | NO
EQPRepairJ22103903 | T6373-39 |  | Prober Down | JB1BYS10 | PA6587B_15U90H | J696C |  | AlarmCode:1995 AlarmMessage:Faied to upload log file | CPAUTO | Prober | 網路異常(Spooling Error) | 2022/08/02 20:19:30 | 2022/08/02 20:29:18 | 0.16 | OK | 13597 | CHECK OK | NO | NO | NO | YES | YES | NO
EQPRepairJ22103914 | T6373-38 | CWC7J0A61-01 | OverKill | JB1GKS10 | PA6543B | J67H5 | 24 | Continuous 2 wafers EDA shows H O/S Bin30-A6525 | CPAUTO | Over Kill | Show H | 2022/08/02 21:27:58 | 2022/08/02 22:01:26 | 0.56 | OK | 17542 | Show H，O/S過多，更換配件#09，上機pass，溫針重對針，range:8，試打(131,0)無異常，起測pass | NO | NO | NO | YES | YES | NO
EQPRepairJ22103926 | T6373-111 |  | Prober Down | TB1R8S82 | PA5137F-LT1 | P07NW.1 |  | laoder-A6899 | CPAUTO | Prober | Loader Down | 2022/08/02 22:07:50 | 2022/08/02 22:45:03 | 0.62 | OK | 15096 | Loader init ok | NO | NO | NO | YES | YES | NO
EQPRepairJ22103927 | T6373-128 | CWC7J0M33-01 | OverKill | VB1BVS01 | PA5270D-LT | U38F8.00 | 11 | Already follow Overkilll handling procedures but still not effective too manyBIN31 | A6899 | Over Kill | Function誤宰 | 2022/08/02 22:14:06 | 2022/08/02 23:10:25 | 0.94 | OK | 13597 | 點測PASS，起測PASS，先還機觀察 | NO | NO | NO | YES | YES | NO
EQPRepairJ22103937 | T6373-39 | CWC8J0088-01 | OverKill | JB1BYS10 | PA6587B_15U90H | J696C | 19 | too many open short | A6786 | Over Kill | O/S誤宰 | 2022/08/02 23:17:30 | 2022/08/03 00:58:32 | 1.68 | OK | 13597 | 更換配件#12，RANGE 10 UM | NO | NO | NO | YES | YES | NO
EQPRepairJ22103938 | T6373-07 | CWC7J0F26-01.002 | Prober Down | JB1TAS15 | PA5208K-LT | QT6L8 | 25 | CANNOT MEASURE INDEX SIZE! | A6296 | Prober | Wafer Edge/Mode Not Found | 2022/08/02 23:19:41 | 2022/08/02 23:27:48 | 0.14 | OK | 15096 | 重抓model ok，check die ok. | NO | NO | NO | YES | YES | NO
EQPRepairJ22103951 | T6373-34 |  | Prober Down | TB1R7S28 | PA5137E1-LT | P097A.1 |  | AlarmCode:1057 AlarmMessage:MODEL NOT FOUND ( ROUGH ALIGNMENT ) !! | CPAUTO | Prober | Wafer Edge/Mode Not Found | 2022/08/03 00:09:16 | 2022/08/03 00:20:59 | 0.20 | OK | 03426 | 重抓Model OK | NO | NO | NO | YES | YES | NO
EQPRepairJ22103959 | T6373-38 | CWC7J0A61-01 | OverKill | JB1GKS09 | PA6543B | J67H5 | 25 | TOO MANY O/S-19697 | CPAUTO | Over Kill | 針位調整 | 2022/08/03 00:21:31 | 2022/08/03 00:25:20 | 0.06 | OK | 03426 | 針位調整PASS | NO | NO | NO | YES | YES | NO
EQPRepairJ22103962 | T6373-29 | CWC7J0H27-01.001 | Prober Down | KB2BNS09 | ND01733D(COF94AA)-A-C | H2PU18.3-21BC12 | 17 | model not found | A6296 | Prober | Wafer Edge/Mode Not Found | 2022/08/03 00:37:14 | 2022/08/03 00:45:41 | 0.14 | OK | 03426 | 重抓Model OK | NO | NO | NO | YES | YES | NO
EQPRepairJ22103964 | T6373-21 | CWC7J0J54-01 | OverKill | JB1TAS26 | PA5208KF-LT | J696H | 07 | Too many O/S Bin2 pm checked(59.5) od10-A6525 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/03 00:46:34 | 2022/08/03 00:48:38 | 0.03 | OK | 03426 | 使用Clean Wafer 清針 | NO | NO | NO | YES | YES | NO
EQPRepairJ22103967 | T6373-16 | CWC7J0G38-01 | OverKill | JB1TAS16 | PA5208K-LT | QT7YW | 22 | too many o/s-A6148 | CPAUTO | Over Kill | O/D不足 | 2022/08/03 00:54:53 | 2022/08/03 01:21:03 | 0.44 | OK | 17542 | 溫針重對針，range:20，試打(148,5)無異常，起測一段pass穩定 | NO | NO | NO | YES | YES | NO
EQPRepairJ22103980 | T6373-07 | CWC7J0F26-01.002 | OverKill | JB1TAS15 | PA5208K-LT | QT77K | 25 | O/S-19697 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/03 01:24:47 | 2022/08/03 01:55:55 | 0.52 | OK | 17542 | 機台清針，重對針，range:8，試打(17,5)無異常，起測一段PASS穩定 | NO | NO | NO | YES | YES | NO
EQPRepairJ22103987 | T6373-128 | CWC7J0M33-01 | OverKill | VB1BVS01 | PA5270D-LT | U38F8.00 | 12 | BIN31 TOO MANY-19697 | CPAUTO | Probe Card | bin31 leak | 2022/08/03 01:42:09 | 2022/08/03 03:20:53 | 1.65 | OK | 13597 | P/C Down----P/C:VB1BVS01送修無備卡掛配件當 1.測試中bin31 leak不穩，日班已下卡烘烤，上機測試一樣bin31過多 2.抬頭空跑測試項無fail 3.故此配件先行送修無備卡配件當 | NO | NO | NO | YES | YES | NO
EQPRepairJ22103993 | T6373-38 | CWC7J0A61-01 | OverKill | JB1GKS09 | PA6543B | J67H5 | 25 | too many o/s,already overkill not effective | A6404 | Over Kill | O/S誤宰 | 2022/08/03 02:20:18 | 2022/08/03 02:47:58 | 0.46 | OK | 17542 | 二分法調針，range:14，試打(19,6)無異常，起測一段PASS穩定 | NO | NO | NO | YES | YES | NO
EQPRepairJ22104010 | T6373-128 |  | Probe Card Down | VB1BVS01 | PA5270D-LT | U38F8.00 |  | P/C Down----P/C:VB1BVS01送修無備卡掛配件當 1.測試中bin31 leak不穩，日班已下卡烘烤，上機測試一樣bin31過多 2.抬頭空跑測試項無fail 3.故此配件先行送修無備卡配件當 | 13597 | Probe Card | P/C Down | 2022/08/03 03:21:18 | 2022/08/03 08:17:38 | 4.94 | OK | 13222 | P/C Down----P/C:VB1BVS01送修無備卡掛配件當 1.測試中bin31 leak不穩，日班已下卡烘烤，上機測試一樣bin31過多 2.抬頭空跑測試項無fail 3.故此配件先行送修無備卡配件當 4.產線 通知改測 上機pass ragne=10 還機掛setup04 | NO | NO | NO | YES | YES | NO
EQPRepairJ22104018 | T6373-38 | CWC7J0724-01 | OverKill | JB1GKS09 | PA6543B | J67H7 | 16 | TOO MANY O/S-19697 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/03 04:16:44 | 2022/08/03 05:05:36 | 0.81 | OK | 17542 | 1.換批起測O/S不過 2.機台溫針，重對針，點測PASS，Data-In後溫針並執行Soaktime 3.再次重對針，range:12，試打(173,5)無異常，起測一段pass穩定 | NO | NO | NO | YES | YES | NO
EQPRepairJ22104046 | T6373-42 | CWC7J0A01-01.001 | Setup Fail | JB8Q6S05 | NA01761D(COF94AA)-A-Y-C | GGR094.1-1B | 18 | 0528 # 18 ID OK.setup failed remount testerhead and pull pin but still o/s no other p/c..pc polish @6f..too large by AE | A6788 | Setup Fail | O/S Fail | 2022/08/03 06:49:15 | 2022/08/03 08:53:31 | 2.07 | OK | 15362 | JB8S6S05下卡磨針,上機PASS,二分法X/Y RANGE:10/8,P/M(165,8) | NO | NO | NO | YES | YES | NO
EQPRepairJ22104048 | T6373-41 | CWC8J0085-01.001 | OverKill | JB4R2S01 | NA01248F(COG4)-A-AT-X2-C | NT2YA-B2C | 15 | eda show i no corr-A6924 | CPAUTO | Over Kill | Show I | 2022/08/03 06:52:37 | 2022/08/03 06:57:43 | 0.09 | OK | 17542 | Remote處理 ----- Show I，T44單片Show I，請先行續測待後續觀察 | NO | NO | NO | YES | YES | NO
EQPRepairJ22104050 | T6373-21 | CWC7J0J54-01 | OverKill | JB1TAS26 | PA5208KF-LT | J696H | 10 | eda show 2R o/s-A6924 | CPAUTO | Over Kill | EDA 2R | 2022/08/03 06:57:42 | 2022/08/03 07:58:22 | 1.01 | OK | 20230 | 1.EDA 2R 2.更換備卡S28,YRG=11,XRG=44,PM(45.6) 3.PASS | NO | NO | NO | YES | YES | NO
EQPRepairJ22104060 | T6373-152 | CWC7J0L75-01 | OverKill | JB2P2S09 | R63455-A1S-EZ-CP | NAHM75.00-B1C | 20 | eda show 2H NO CORR WAFER-A6858 | CPAUTO | Over Kill | 2H | 2022/08/03 07:48:37 | 2022/08/03 08:04:32 | 0.27 | OK | 13222 | 2H 同批同機卡相同問題 續測 | NO | NO | NO | YES | YES | NO
EQPRepairJ22104064 | T6373-41 | CWC8J0085-01.001 | OverKill | JB4R2S01 | NA01248F(COG4)-A-AT-X2-C | NT2YA-B2C | 15 | # 12 already follow overkill, too many o/s-A6143 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/03 08:07:31 | 2022/08/03 08:28:39 | 0.35 | OK | 10210 | 重新對針(34.7) | NO | NO | NO | YES | YES | NO
EQPRepairJ22104093 | T6373-35 | CWC7J0L29-01.001 | OverKill | JB1GVS24 | NC01682B(COG48AA)-A-AT-N-C | GGH464.1-BC2 | 06 | TOO MANY FAILBIN-A6920 | CPAUTO | Over Kill | Function誤宰 | 2022/08/03 11:04:18 | 2022/08/03 11:51:15 | 0.78 | OK | 07764 | #06 BIN18 LEAK FAIL 重新對針抓取CONTACT (17.5) 調整風管PASS | NO | NO | NO | YES | YES | NO
EQPRepairJ22104102 | T6373-128 | CEC6J0112-01 | OverKill | JB1BYS10 | PA6587B_15U90H | J65P8 | 03 | # 03 已依誤宰處理流程處置無效 (O/S 過多) | 05131 | Over Kill | 機台清針後Pass | 2022/08/03 11:47:19 | 2022/08/03 12:06:12 | 0.31 | OK | 07764 | Remote處理----open 清針pass | NO | NO | NO | YES | YES | NO
EQPRepairJ22104161 | T6373-07 | CWC7J0G37-01 | Prober Down | JB1TAS20 | PA5208K-LT | QT6L8 | 05 | MODEL NOT FOUND-A6317 | CPAUTO | Prober | Wafer Mode Not Found | 2022/08/03 17:01:47 | 2022/08/03 17:16:18 | 0.24 | OK | 15362 | 重新抓取,CHECK DIE OK | NO | NO | NO | YES | YES | NO
EQPRepairJ22104163 | T6373-70 | CWC7J0J88-01.001 | OverKill | TB1R8S85 | PA5223Q_12U60H-LT | P06UK.1 | 11 | O/S TOO MANY-A6317 | CPAUTO | Over Kill | o/s | 2022/08/03 17:15:06 | 2022/08/03 18:00:49 | 0.76 | OK | 13222 | o/s 重新對針機台清針OK (18.1) 起測時 DATA IN座標異常 確定無DATA IN到wafer座標，確認此為非燒CODE型號 WAFER ALIGNMENT j重測 | NO | NO | NO | YES | YES | NO
EQPRepairJ22104201 | T6373-38 | CWC7J0A61-01.003 | OverKill | JB1GKS09 | PA6543B | J67H5 | 13 | too many o/s | A6941 | Over Kill | O/S誤宰 | 2022/08/03 20:40:22 | 2022/08/03 21:05:23 | 0.42 | OK | 17091 | 固定pin清針後確認針壓ok 二分法調針r:-10~4 起測回收os ok | NO | NO | NO | YES | YES | NO
EQPRepairJ22104251 | T6373-15 | CWC8J0118-01 | OverKill | JB6K8S01 | NC01682A(COG48AA)-A-AT-N3-C | GGR320.1-B1C | 19 | too many o/s | A6898 | Over Kill | O/D不足 | 2022/08/03 23:33:34 | 2022/08/03 23:46:45 | 0.22 | OK | 20182 | 清針PASS PM(137,6) | NO | NO | NO | YES | YES | NO
EQPRepairJ22104257 | T6373-39 | CWC8J0168-01 | OverKill | JB1BYS12 | PA6587B_15U90H | J6969 | 03 | #3 Wafer 1.2 eda show H (OS) | A6896 | Over Kill | 2H | 2022/08/04 00:00:22 | 2022/08/04 00:38:02 | 0.63 | OK | 20182 | 清針，調針RG-10 PM(45,7) 起測一段穩定 | NO | NO | NO | YES | YES | NO
EQPRepairJ22104264 | T6373-34 | CWC7J0J51-01.001 | OverKill | TB1R8S76 | PA5137F-LT1 | P094H.1 | 07 | too many bin 9 | A6898 | Over Kill | Function誤宰 | 2022/08/04 00:21:22 | 2022/08/04 00:47:24 | 0.43 | OK | 17091 | bin9 前去點測皆pass，機台清針確認針壓Ok 二分法調針 起測一段穩定還機觀察 | NO | NO | NO | YES | YES | NO
EQPRepairJ22104266 | T6373-128 | CWC7J0M33-01 | OverKill | VB1BVS01 | PA5270D-LT | U38F8.00 | 03 | TOO MANY FAIL BIN . | A6941 | Over Kill | Function誤宰 | 2022/08/04 00:38:10 | 2022/08/04 00:55:30 | 0.29 | OK | 20277 | Remote---- 1.Bin31 leakage Range:10 PASS | NO | NO | NO | YES | YES | NO
EQPRepairJ22104274 | T6373-128 | CWC7J0M33-01 | OverKill | VB1BVS01 | PA5270D-LT | U38F8.00 | 12 | BIN31 | 03491 | Probe Card | P/C Down---- | 2022/08/04 01:32:46 | 2022/08/04 01:51:33 | 0.31 | OK | 20277 | P/C Down----P/C:VB1BVS01送修無備卡掛配件當 1.測試中bin31 leak不穩 2.先前有相同送修紀錄 3.故配件送修無備卡配件當 | NO | NO | NO | YES | YES | NO
EQPRepairJ22104280 | T6373-128 |  | Probe Card Down | VB1BVS01 | PA5270D-LT | U38F8.00 |  | P/C Down----P/C:VB1BVS01送修無備卡掛配件當 1.測試中bin31 leak不穩 2.先前有相同送修紀錄 3.故配件送修無備卡配件當 | 20277 | Probe Card | Bin31 Leakage | 2022/08/04 01:52:03 | 2022/08/04 06:05:31 | 4.22 | OK | 20277 | 輪工處理---- 1.Bin31 Leakage 送修，配件烘烤上機PASS，待觀察 Range:X>20，Y:14 | NO | NO | NO | YES | YES | NO
EQPRepairJ22104294 | T6373-70 | CWC7J0J88-01.001 | OverKill | TB1R8S85 | PA5223Q_12U60H-LT | P06UK.1 | 17 | too many o/s | A6675 | Over Kill | O/S誤宰 | 2022/08/04 03:13:43 | 2022/08/04 03:28:00 | 0.24 | OK | 16157 | 機台清針後Pass | NO | NO | NO | YES | YES | NO
EQPRepairJ22104305 | T6373-16 | CWC7J0H15-01 | OverKill | JB1TAS25 | PA5208K-LT | QT84Q | 10 | o/s-03491 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/04 04:13:38 | 2022/08/04 04:29:48 | 0.27 | OK | 20277 | Remote---- 1.機台清針 Range:14 起測一段PASS穩定，還機生產 | NO | NO | NO | YES | YES | NO
EQPRepairJ22104311 | T6373-34 | CWC7J0J51-01.001 | OverKill | TB1R8S76 | PA5137F-LT1 | P094H.1 | 09 | too many o/s | A6941 | Over Kill | O/S誤宰 | 2022/08/04 04:32:36 | 2022/08/04 04:49:14 | 0.28 | OK | 20182 | 1.清針 重對針PASS 2.二分法RG-10 PM(43,6) | NO | NO | NO | YES | YES | NO
EQPRepairJ22104343 | T6373-15 | CWC8J0135-01 | OverKill | JB6K8S01 | NC01682A(COG48AA)-A-AT-N3-C | GGP621.1-BC | 03 | eda show I DUTTDIFF NO CORR | A6675 | Over Kill | EDA SHOW I | 2022/08/04 07:16:16 | 2022/08/04 08:49:58 | 1.56 | OK | 20230 | 1.EDA SHOW I 2.更換備卡S03,上機驗證PASS,RG=20,PM(45.6) 3.PASS | NO | NO | NO | YES | YES | NO
EQPRepairJ22104359 | T6373-147 | CWC7J0F92-01 | OverKill | TB1R8S67 | PA5137F-LT1 | P094F.1 | 09 | too many os | A6957 | Over Kill | O/S誤宰 | 2022/08/04 08:18:25 | 2022/08/04 08:38:43 | 0.34 | OK | 17229 | 1.針尖輕微黑點,開啟cleaning wafer清針檢查針位針壓ok關閉cleaninwafer起測一段穩定pass. | NO | NO | NO | YES | YES | NO
EQPRepairJ22104375 | T6373-70 | CWC7J0J88-01.001 | OverKill | TB1R8S85 | PA5223Q_12U60H-LT | P06UK.1 | 22 | #22 EDA CONTINUES SHOW 2R OS-A6132 | CPAUTO | Over Kill | 2R | 2022/08/04 09:14:05 | 2022/08/04 09:48:01 | 0.57 | OK | 15362 | 2R O/S,清調針,P/M(16,5),續測觀察 | NO | NO | NO | YES | YES | NO
EQPRepairJ22104384 | T6373-13 | CWC8J0122-01 | OverKill | JB6H1S05 | NC01690G(COG48AA)-A-AT-N4-C | GGG485.1-BC | 04 | eda show 2 h-A6293 | CPAUTO | Over Kill | Show H | 2022/08/04 09:55:30 | 2022/08/04 10:00:09 | 0.08 | OK | 10210 | Bin26 Show H 請產線依EDA執行 | NO | NO | NO | YES | YES | NO
EQPRepairJ22104387 | T6373-147 | CWC7J0F92-01 | OverKill | TB1R8S67 | PA5137F-LT1 | P094F.1 | 10 | Already follow Overkilll handling procedures but still not effective too many o/s | A6957 | Over Kill | o/s | 2022/08/04 09:57:38 | 2022/08/04 10:15:18 | 0.29 | OK | 13222 | 機台清針 加針壓od:5>10 | NO | NO | NO | YES | YES | NO
EQPRepairJ22104394 | T6373-15 | CWC8J0135-01 | OverKill | JB6K8S03 | NC01682A(COG48AA)-A-AT-N3-C | GGP621.1-BC | 05 | #3 EDA 秀 I NO CORR | 05131 | Over Kill | Show I | 2022/08/04 10:22:22 | 2022/08/04 10:38:51 | 0.27 | OK | 10210 | 2I 請產線續測 | NO | NO | NO | YES | YES | NO
EQPRepairJ22104399 | T6373-70 | CWC7J0J88-01.001 | OverKill | TB1R8S85 | PA5223Q_12U60H-LT | P06UK.1 | 23 | #23 2R-09052 | CPAUTO | Over Kill | 2R | 2022/08/04 10:54:11 | 2022/08/04 11:05:03 | 0.18 | OK | 10210 | #22已改善，請產線續測待觀察 | NO | NO | NO | YES | YES | NO
EQPRepairJ22104411 | T6373-135 |  | Prober Down | VB4BKS01 | PA5129LS_12U80H-LT | P078R.1 |  | O/S SYSTEM ERROR | A6957 | Prober | O/S SYSTEM CALL ERROR | 2022/08/04 11:38:59 | 2022/08/04 13:36:29 | 1.96 | OK | 13222 | Prober down---O/S SYSTEM CALL ERROR 1.重開機多次 無效 2. 重插拔CPU Board 無效 3. 轉掛 Prober down | NO | NO | NO | YES | YES | NO
EQPRepairJ22104413 | T6373-103 | CWC7J0L30-01 | OverKill | JB1GVS15 | NC01682B(COG48AA)-A-AT-N-C | GGG498.1-BC111 | 25 | too many o/s-A6785 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/04 11:54:32 | 2022/08/04 12:14:38 | 0.33 | OK | 07764 | open 調針位/清針 | NO | NO | NO | YES | YES | NO
EQPRepairJ22104419 | T6373-01 |  | Tester Down | JB1TAS23 | PA5208K-LT | QT7LG |  | cbvp cannot open | A6920 | Product Issue | Other | 2022/08/04 12:23:01 | 2022/08/04 13:42:48 | 1.33 | OK | 07764 | 常日處理 ok | NO | NO | NO | YES | YES | NO
EQPRepairJ22104421 | T6373-37 |  | Tester Down | JB1GVS21 | NC01682A(COG48AA)-A-AT-N3-C | GGR320.1-B1 |  | cbvp cannot open | A6920 | Product Issue | Other | 2022/08/04 12:27:49 | 2022/08/04 13:39:21 | 1.19 | OK | 07764 | 常日處理 ok | NO | NO | NO | YES | YES | NO
EQPRepairJ22104422 | T6373-70 | CWC7J0J88-01.001 | OverKill | TB1R8S85 | PA5223Q_12U60H-LT | P06UK.1 | 25 | #25 pls check and confirm pm too many os-A6132 | CPAUTO | Over Kill | o/s | 2022/08/04 12:29:57 | 2022/08/04 13:03:55 | 0.57 | OK | 13222 | 更換配件#79上機pass (157.4) | NO | NO | NO | YES | YES | NO
EQPRepairJ22104424 | T6373-25 |  | Setup Fail | JB1NUS02 | ILIRJE6ADCM-2CP | QSR8Q |  | tester螢幕沒有畫面 | 11316 | Tester | 螢幕黑頻 | 2022/08/04 12:35:15 | 2022/08/04 13:50:35 | 1.26 | OK | 20230 | 通知常日班,更換備品ok | NO | NO | NO | YES | YES | NO
EQPRepairJ22104428 | T6373-35 |  | Tester Down | JB1GVS24 | NC01682A(COG48AA)-A-AT-N3-C | GGP620.1-2BC22 |  | Cant open cbvp/mass production..-A6143 | CPAUTO | Tester | Other | 2022/08/04 12:52:16 | 2022/08/04 13:35:02 | 0.71 | OK | 08850 | do init datarecover,Run ok | NO | NO | NO | YES | YES | NO
EQPRepairJ22104432 | T6373-147 | CWC7J0F92-01 | OverKill | TB1R8S67 | PA5137F-LT1 | P094F.1 | 11 | TOO MANY OS | A6957 | Over Kill | O/S誤宰 | 2022/08/04 13:00:12 | 2022/08/04 14:07:04 | 1.11 | OK | 13581 | 1.原配件殘金，更換#83配件上機BIN3 SHORT ,清調針無效，再更換#77配件上機pass，故#83配件送修 | NO | NO | NO | NO | YES | NO
EQPRepairJ22104439 | T6373-42 | CWC8J0118-01 | Prober Down | JB1GVS20 | NC01682A(COG48AA)-A-AT-N3-C | GGP620.1-2B | 02 | cannot measure index size-A7086 | CPAUTO | Prober | Wafer Edge/Mode Not Found | 2022/08/04 13:30:47 | 2022/08/04 13:39:05 | 0.14 | OK | 10210 | 按START OK | NO | NO | NO | YES | YES | NO
EQPRepairJ22104442 | T6373-70 | CWC7J0J88-01.001 | OverKill | TB1R8S79 | PA5223Q_12U60H-LT | P06UK.1 | 25 | #25 PLS CHECK AND CON FIRM PM TOO MANY FAIL BIN 9 AND OS-A6132 | CPAUTO | Over Kill | bin9 | 2022/08/04 13:33:32 | 2022/08/04 15:09:50 | 1.61 | OK | 13222 | 配件#79 上機測試一行後bin9 過多 清、調針無效 更換配件#85 上機固定pin bin3 short 再次更換配件#78 上機pass 配件#79、85 送修 | NO | NO | NO | YES | YES | NO
EQPRepairJ22104443 | T6373-135 |  | Prober Down | VB4BKS01 | PA5129LS_12U80H-LT | P078R.1 |  | Prober down---O/S SYSTEM CALL ERROR 1.重開機多次 無效 2. 重插拔CPU Board 無效 3. 轉掛 Prober down | 13222 | Prober | O/S SYSTEM CALL ERROR | 2022/08/04 13:37:15 | 2022/08/04 18:01:44 | 4.41 | OK | 13222 | Prober down---O/S SYSTEM CALL ERROR 1.重開機多次 無效 2. 重插拔CPU Board 無效 3. 轉掛 Prober down 4. 請常日設備 郎哥 更換硬碟OK 上機RESETUP 上機PASS (138.4) | NO | NO | NO | YES | YES | NO
EQPRepairJ22104450 | T6373-12 | CWC7J0861-01.002 | Prober Down | JB1PNS08 | NA01761D(COF94AA)-A-Y-C | GGN637.1-BC221 | 01 | AlarmCode:1057 AlarmMessage:MODEL NOT FOUND ( ROUGH ALIGNMENT ) !! | CPAUTO | Prober | Mode Not Found | 2022/08/04 14:21:38 | 2022/08/04 14:37:14 | 0.26 | OK | 20075 | 1.想要執行WAFER ALIGMENT時誤觸重新對針 2.重新對針RNAGE 10UM,MARK 218,4 3.檢查CHECK DIE OK | NO | NO | NO | YES | YES | NO
EQPRepairJ22104453 | T6373-26 | CWC1J1J97-01 | OverKill | JB1NUS03 | ILIRJE6ADCM-2CP | QSR8Q | 06 | too many failbin-A6920 | CPAUTO | Over Kill | Function誤宰 | 2022/08/04 14:48:13 | 2022/08/04 15:08:17 | 0.33 | OK | 20230 | 1.FUNCTION BIN 37 2.CLEAN WAFER,二分法調針YRG=21.XRG=44,PM(48.6) 3.PASS | NO | NO | NO | YES | YES | NO
EQPRepairJ22104459 | T6373-12 | CWC7J0861-01.002 | Prober Down | JB1PNS08 | NA01761D(COF94AA)-A-Y-C | GGN637.1-BC221 | 01 | AlarmCode:1054 AlarmMessage:AUTO FOCUS CHECK ERROR !! | CPAUTO | Prober | Wafer Edge/Mode Not Found | 2022/08/04 15:18:18 | 2022/08/04 15:54:24 | 0.60 | OK | 10210 | 重抓MODE，請產線重測#01 | NO | NO | NO | YES | YES | NO
EQPRepairJ22104475 | T6373-43 | CWC7J0312-01.001 | Prober Down | JB1UJS03 | PA5129LS_12U80H-LT | P044K.1 | 17 | cannot measure index size-A7086 | CPAUTO | Prober | Wafer Edge/Mode Not Found | 2022/08/04 17:19:15 | 2022/08/04 17:33:14 | 0.23 | OK | 13222 | REMOTE---Wafer Edge/Mode Not Found OK | NO | NO | NO | YES | YES | NO
EQPRepairJ22104484 | T6373-111 |  | Tester Down | TB1R8S82 | PA5137F-LT1 | P07HJ.1 |  | CANNOT LOAD CBVP PRODUCTION | A6317 | Product Issue | 自動化異常 | 2022/08/04 18:16:39 | 2022/08/04 18:24:24 | 0.13 | OK | 10210 | 無法開啟自動化，待產品確認 | NO | NO | NO | YES | YES | NO
EQPRepairJ22104485 | T6373-111 |  | Product Issue | TB1R8S82 | PA5137F-LT1 | P07HJ.1 |  | 無法開啟自動化，待產品確認 | 10210 | Product Issue | 自動化異常 | 2022/08/04 18:24:52 | 2022/08/04 18:57:50 | 0.55 | OK | 10210 | 產品處理後OK | NO | NO | NO | YES | YES | NO
EQPRepairJ22104486 | T6373-42 |  | Prober Down | JB1GVS20 | NC01682A(COG48AA)-A-AT-N3-C | GGR320.1-B1C1 |  | AlarmCode:1057 AlarmMessage:MODEL NOT FOUND ( ROUGH ALIGNMENT ) !! | CPAUTO | Prober | Wafer Edge/Mode Not Found | 2022/08/04 18:26:26 | 2022/08/04 18:39:25 | 0.22 | OK | 13222 | Wafer Edge/Mode Not Found OK | NO | NO | NO | YES | YES | NO
EQPRepairJ22104497 | T6373-26 | CWC1J1J97-01 | OverKill | JB1NUS03 | ILIRJE6ADCM-2CP | QSR8Q | 08 | #6.7 EDA 2H no corr. | A6920 | Over Kill | 2H | 2022/08/04 19:01:14 | 2022/08/04 19:33:42 | 0.54 | OK | 17091 | 2H同批與其他批不同機卡組合一樣有BIN27 37問題 續測觀察 | NO | NO | NO | YES | YES | NO
EQPRepairJ22104498 | T6373-12 | CWC7J0A01-01.001 | Other | JB1PNS08 | NA01761D(COF94AA)-A-Y-C | GGN637.1-BC221 | 18 | PLEASE CONFIRM PROBEMARK.PM NOT GOOOD | A6564 | Other | 確認針位 | 2022/08/04 19:03:35 | 2022/08/04 19:24:38 | 0.35 | OK | 17184 | 1.調針(55.0) | NO | NO | NO | NO | YES | NO
EQPRepairJ22104504 | T6373-37 | CWC8J0118-01.001 | OverKill | JB1GVS21 | NC01682A(COG48AA)-A-AT-N3-C | GGR320.1-B1C2 | 16 | # 16 pm checked shift ,adjust pm up 6 um too many fail bin29-A6143 | CPAUTO | Over Kill | Function誤宰 | 2022/08/04 19:14:51 | 2022/08/04 19:38:24 | 0.39 | OK | 15492 | 1.針位貼下，清、調針 PASS，RANGE:16 2.P/M(169，5)OK | NO | NO | NO | YES | YES | NO
EQPRepairJ22104505 | T6373-135 | CWC8J0087-01 | OverKill | VB4BKS01 | PA5129LS_12U80H-LT | P078R.1 | 07 | o/s-03491 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/04 19:15:54 | 2022/08/04 19:55:12 | 0.66 | OK | 20182 | 1.重對針PASS 二分法RG-10 PM(115,1) | NO | NO | NO | YES | YES | NO
EQPRepairJ22104510 | T6373-25 | CWC1J1J97-01.002 | OverKill | JB1NUS02 | ILIRJE6ADCM-2CP | QSR8Q | 18 | EDA 2H, NO CORR WAFER-A6599 | CPAUTO | Over Kill | 2H | 2022/08/04 19:28:24 | 2022/08/04 20:28:53 | 1.01 | OK | 17091 | 2H BIN27偏多，更換配件#01 全面BIN26 FUNCTION，與前次送修紀錄相同，磨原配件#02 上機PASS RANGE:-8~8 PM(151，5) | NO | NO | NO | YES | YES | NO
EQPRepairJ22104520 | T6373-128 | CWC7J0M33-01.001 | OverKill | VB1BVS01 | PA5270D-LT | U38F8.00 | 16 | too many fail bin 31-A6673 | CPAUTO | Over Kill | Function誤宰 | 2022/08/04 19:59:18 | 2022/08/04 20:15:42 | 0.27 | OK | 15492 | 1.BIN31，針髒清針 PASS | NO | NO | NO | YES | YES | NO
EQPRepairJ22104524 | T6373-26 | CWC1J1J97-01 | OverKill | JB1NUS03 | ILIRJE6ADCM-2CP | QSR8Q | 08 | TOO MANY OPEN/SHORT (PM UP AND DOWN)-A6599 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/04 20:15:10 | 2022/08/04 20:44:32 | 0.49 | OK | 17091 | 清針後重新對針 二分法R:-8~8 點測穩定還機 | NO | NO | NO | YES | YES | NO
EQPRepairJ22104531 | T6373-11 | CWC8J0120-01 | OverKill | JB1GVS17 | NC01682A(COG48AA)-A-AT-N3-C | GGR011.1-1B11C11 | 09 | Already follow Overkilll handling procedures but still not effective too many O/S. | A6944 | Over Kill | O/S誤宰 | 2022/08/04 20:52:18 | 2022/08/04 21:18:01 | 0.43 | OK | 17091 | REMOTE---機台清針確認針壓OK 二分法R:2~18 PM(43，6)點測穩定還機 | NO | NO | NO | YES | YES | NO
EQPRepairJ22104536 | T6373-86 | CWC7J0814-01.006 | Prober Down | JB1Q4S16 | PA6542C | J683K | 05 | cannot measure index size-A6944 | CPAUTO | Prober | Wafer Edge/Mode Not Found | 2022/08/04 21:24:06 | 2022/08/04 21:36:35 | 0.21 | OK | 10897 | Mode Not Found/check die/針位OK | NO | NO | NO | NO | YES | NO
EQPRepairJ22104539 | T6373-128 | CWC7J0M33-01.001 | OverKill | VB1BVS01 | PA5270D-LT | U38F8.00 | 16 | reprobing fail bin 31 still did no passed-A6673 | CPAUTO | Over Kill | 測試中bin31 | 2022/08/04 21:39:02 | 2022/08/05 00:26:02 | 2.78 | OK | 20182 | 1.測試中bin31 leak不穩，下卡烘烤，接手上機PASS 2.二分法RG-10 PM(40,5) | NO | NO | NO | YES | YES | NO
EQPRepairJ22104550 | T6373-39 | CWC8J0168-01.001 | OverKill | JB1BYS12 | PA6587B_15U90H | J6969 | 25 | too many o/s PC has particle | A6673 | Over Kill | 配件卡毛 | 2022/08/04 22:05:19 | 2022/08/04 22:17:30 | 0.20 | OK | 16157 | O/S 卡毛，前蓋清針pass | NO | NO | NO | YES | YES | NO
EQPRepairJ22104567 | T6373-39 | CWC8J0168-01.001 | OverKill | JB1BYS12 | PA6587B_15U90H | J6969 | 25 | already follow overkill procedure but still not effective. | A7048 | Over Kill | O/S誤宰 | 2022/08/04 23:07:15 | 2022/08/04 23:20:53 | 0.23 | OK | 20182 | 清針 調針PASS PM(46,7) | NO | NO | NO | YES | YES | NO
EQPRepairJ22104576 | T6373-28 | CWC8J0221-01.001 | OverKill | JB8Y9S06 | ILISLQPEDQD-2CP | ILA13191.1 | 17 | O/S Too Many | A6318 | Over Kill | O/S誤宰 | 2022/08/04 23:35:51 | 2022/08/04 23:53:34 | 0.30 | OK | 20182 | 1針位偏下 調整PASS PM(67,13) | NO | NO | NO | YES | YES | NO
EQPRepairJ22104577 | T6373-12 | CWC7J0A01-01.001 | OverKill | JB1PNS08 | NA01761D(COF94AA)-A-Y-C | GGR094.1-1BC21 | 18 | after mbd still too many o/s .. reprobing cannot pass | A6944 | Over Kill | O/S誤宰 | 2022/08/04 23:37:58 | 2022/08/04 23:59:30 | 0.36 | OK | 17184 | Remote--1.清針&調針(168.8) | NO | NO | NO | YES | YES | NO
EQPRepairJ22104581 | T6373-103 | CWC8J0137-01.001 | OverKill | JB1GVS15 | NC01682B(COG48AA)-A-AT-N-C | GGC155.1-221BC2 | 10 | too many fail bin 29 | A6857 | Over Kill | 配件殘金 | 2022/08/04 23:43:02 | 2022/08/05 00:29:36 | 0.78 | OK | 15492 | 1.下卡磨針，二分法(0，-10)、range:12 2.P/M(169，5)OK | NO | NO | NO | YES | YES | NO
EQPRepairJ22104582 | T6373-21 | CWC7J0J52-01 | OverKill | JB1TAS28 | PA5208KF-LT | J68QW | 01 | o/s too many-A6318 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/04 23:56:50 | 2022/08/05 00:08:33 | 0.20 | OK | 15492 | 1.清、調針 PASS 2.P/M(148，5)OK | NO | NO | NO | YES | YES | NO
EQPRepairJ22104584 | T6373-86 | CWC7J0J58-01.004 | OverKill | JB1Q4S16 | PA6542C | J683L | 01 | already follow ovwekill but still too many o/s .. reprobinfg cannot pass-A6944 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/05 00:09:04 | 2022/08/05 00:28:14 | 0.32 | OK | 10110 | Remote---short調整針位後pass | NO | NO | NO | YES | YES | NO
EQPRepairJ22104596 | T6373-28 | CWC8J0221-01.001 | OverKill | JB8Y9S06 | ILISLQPEDQD-2CP | ILA13191.1 | 17 | too many open/short-A6599 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/05 00:39:55 | 2022/08/05 01:11:14 | 0.52 | OK | 15492 | 1.清針、二分法(1，6)、RANGE:16 2.P/M(222，9)OK | NO | NO | NO | YES | YES | NO
EQPRepairJ22104597 | T6373-13 | CWC8J0190-01 | OverKill | JB6H1S05 | NC01690F(COG46BA)-A-AT-C | GGH632.1-BC | 09 | eda show i no corr wafer | A6944 | Over Kill | SHOW I | 2022/08/05 00:52:48 | 2022/08/05 01:01:25 | 0.14 | OK | 17184 | 1.Dut4 補掛，換片無此問題續觀察 | NO | NO | NO | NO | YES | NO
EQPRepairJ22104609 | T6373-11 | CWC8J0135-01.001 | OverKill | JB1GVS17 | NC01682A(COG48AA)-A-AT-N3-C | GGP621.1-BC21 | 01 | already follow overkill but still too many o/s ad failbin-A6944 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/05 01:19:11 | 2022/08/05 01:59:27 | 0.67 | OK | 17091 | 更換配件#22 Pass r:-6~10 pm(138，6) | NO | NO | NO | YES | YES | NO
EQPRepairJ22104612 | T6373-128 | CWC7J0M33-01.001 | OverKill | VB1BVS01 | PA5270D-LT | U38F8.00 | 17 | # 15-16 eda show 2H (NO CORR WAFER) | A6673 | Over Kill | 2h | 2022/08/05 01:34:41 | 2022/08/05 01:40:42 | 0.10 | OK | 20182 | 2h 不同批機台有相同問題，續測 | NO | NO | NO | YES | YES | NO
EQPRepairJ22104615 | T6373-136 | CWC8J0191-01 | OverKill | JB6H1S04 | NC01690G(COG48AA)-A-AT-N-C | GGN489.1-BC | 18 | #18 Wafer 11.12 eda show H (O/s bin 57) | A6896 | Over Kill | Show H | 2022/08/05 01:40:05 | 2022/08/05 01:47:29 | 0.12 | OK | 20182 | Show H 續測 | NO | NO | NO | YES | YES | NO
EQPRepairJ22104618 | T6373-12 | CWC8J0136-01 | Setup Fail | JB6K8S02 | NC01682A(COG48AA)-A-AT-N3-C | GGP622.1-B | 08 | setup fail 對針異常，待輪工確認 | 10851 | Setup Fail | 針位異常 needle abnormal | 2022/08/05 01:43:18 | 2022/08/05 02:27:58 | 0.74 | OK | 17091 | remote----對針異常1-3、4-13 by pass無效，全by後再次重新對針，自動對針Pass，請產線掛setup02 | NO | NO | NO | YES | YES | NO
EQPRepairJ22104625 | T6373-28 | CWC8J0221-01.001 | OverKill | JB8Y9S06 | ILISLQPEDQD-2CP | ILA13191.1 | 18 | eda SHOWS 2R OPEN/SHORT-A6599 | CPAUTO | Over Kill | 2R | 2022/08/05 02:09:39 | 2022/08/05 02:17:06 | 0.12 | OK | 15492 | 1.已MBD 處理 OK，下一片測試狀況良好，無 O/S 2.續測觀察 | NO | NO | NO | YES | YES | NO
EQPRepairJ22104645 | T6373-39 |  | Prober Down | JB1BYS12 | PA6587B_15U90H | J696F |  | AlarmCode:1995 AlarmMessage:Faied to upload log file | CPAUTO | Prober | Other | 2022/08/05 03:28:35 | 2022/08/05 03:38:47 | 0.17 | OK | 16157 | OK | NO | NO | NO | YES | YES | NO
EQPRepairJ22104651 | T6373-25 | CWC1J1J97-01.002 | OverKill | JB1NUS02 | ILIRJE6ADCM-2CP | QSR8Q | 21 | EDA CONT. 2H NO CORR WAFER-A6599 | CPAUTO | Over Kill | 2H | 2022/08/05 04:05:59 | 2022/08/05 04:10:37 | 0.08 | OK | 15492 | 1.同批不同機、卡，皆有相同之問題 2.續測觀察 | NO | NO | NO | YES | YES | NO
EQPRepairJ22104652 | T6373-103 | CWC8J0137-01.001 | OverKill | JB1GVS15 | NC01682B(COG48AA)-A-AT-N-C | GGC155.1-221BC2 | 12 | too many fail bin 29-A6857 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/05 04:08:59 | 2022/08/05 05:37:38 | 1.48 | OK | 20182 | 1.下卡磨針，上機pass 2.二分法RG-10 PM(18,5) | NO | NO | NO | YES | YES | NO
EQPRepairJ22104653 | T6373-39 | CWC7J0K25-01 | OverKill | JB1BYS12 | PA6587B_15U90H | J696F | 14 | # 13 EDA SHOW I (WARNING O/S MODE TEST) | A6673 | Over Kill | I | 2022/08/05 04:14:45 | 2022/08/05 04:19:02 | 0.07 | OK | 15492 | 1.[!!! WARNING !!!OPEN AND SHORT MODE TEST] | NO | NO | NO | YES | YES | NO
EQPRepairJ22104669 | T6373-37 | CWC8J0118-01.001 | OverKill | JB1GVS21 | NC01682A(COG48AA)-A-AT-N3-C | GGR320.1-B1C2 | 19 | O/S Too Many-A6318 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/05 05:36:48 | 2022/08/05 05:49:45 | 0.22 | OK | 17091 | 清針後確認針壓OK 二分法R:16~32 | NO | NO | NO | YES | YES | NO
EQPRepairJ22104692 | T6373-152 | CWC7J0H67-01 | Prober Down | JB2P2S09 | R63455-A1S-EZ-CP | NAHM72.00-B1C | 08 | AlarmCode:8241 AlarmMessage:CLEANING UNIT COUNT EXCEEDED LIMIT !! COUNTER CLEAR OR CLEANING UNIT CHANGE !! | CPAUTO | Prober | 清針次數到達 | 2022/08/05 07:04:49 | 2022/08/05 07:37:02 | 0.54 | OK | 13222 | UNIT清針次數到達，確認OK 還機生產 | NO | NO | NO | YES | YES | NO
EQPRepairJ22104693 | T6373-37 | CWC8J0118-01.001 | OverKill | JB1GVS21 | NC01682A(COG48AA)-A-AT-N3-C | GGR320.1-B1C2 | 20 | EDA 2R-A6318 | CPAUTO | Over Kill | EDA 2R | 2022/08/05 07:05:03 | 2022/08/05 07:53:54 | 0.81 | OK | 20230 | 1.EDA 2R 2.更換備卡S26,YRG=15,XRG=38,PM(46.6) 3.PASS | NO | NO | NO | YES | YES | NO
EQPRepairJ22104697 | T6373-26 | CWC1J1J97-01 | OverKill | JB1NUS03 | ILIRJE6ADCM-2CP | QSR8Q | 13 | EDA CONT. 2H NO CORR WAFER-A6599 | CPAUTO | Over Kill | Function誤宰 | 2022/08/05 07:10:15 | 2022/08/05 07:30:46 | 0.34 | OK | 07764 | 2H 同批於ND3-25測試相同 , 續測 | NO | NO | NO | YES | YES | NO
EQPRepairJ22104706 | T6373-86 | CWC7J0J58-01.004 | OverKill | JB1Q4S16 | PA6542C | J683L | 05 | TOO MANY O/S-A6131 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/05 08:11:01 | 2022/08/05 08:25:22 | 0.24 | OK | 10897 | 清調針PASS,RANGE 12,MARK 174,2 | NO | NO | NO | YES | YES | NO
EQPRepairJ22104722 | T6373-86 | CWC7J0J58-01.004 | OverKill | JB1Q4S16 | PA6542C | J683L | 05 | TOO MANY O/S-A6131 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/05 09:21:42 | 2022/08/05 09:40:26 | 0.31 | OK | 20075 | 重新對針pass,range 10,mark -1,2 | NO | NO | NO | YES | YES | NO
EQPRepairJ22104733 | T6373-135 | CWC7J0L69-01 | OverKill | VB4BKS01 | PA5129LS_12U80H-LT | P078M.1 | 03 | MBD 2H NO CORR WAFER | A6957 | Over Kill | 2H | 2022/08/05 10:22:00 | 2022/08/05 10:31:17 | 0.15 | OK | 13222 | 兩片不同FAIL BIN 續測觀察 | NO | NO | NO | YES | YES | NO
EQPRepairJ22104736 | T6373-103 | CWC8J0137-01.001 | OverKill | JB1GVS15 | NC01682B(COG48AA)-A-AT-N-C | GGC155.1-221BC2 | 13 | CONTINOUS FAIL BIN29-A5813 | CPAUTO | Over Kill | Function誤宰 | 2022/08/05 10:54:32 | 2022/08/05 11:37:49 | 0.72 | OK | 07764 | bin29 掛run 維修 | NO | NO | NO | YES | YES | NO
EQPRepairJ22104738 | T6373-111 | CWC7J0761-01 | OverKill | TB1R8S82 | PA5137F-LT1 | P07HJ.1 | 15 | too many o/s. has particle p/c-A7086 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/05 11:04:35 | 2022/08/05 11:48:31 | 0.73 | OK | 20075 | 換卡#S79range 12,mark 162,3 | NO | NO | NO | YES | YES | NO
EQPRepairJ22104741 | T6373-13 |  | Tester Down | JB6H1S05 | NC01690F(COG46BA)-A-AT-C | GGH632.1-BC1 |  | keyboard NG | A6920 | Tester | Monitor /滑鼠/鍵盤異常 | 2022/08/05 11:22:17 | 2022/08/05 12:24:20 | 1.03 | OK | 20075 | 1.Tester keyboard NG,插拔Remote & USB 無效 2.開main frame側門插拔USB驗證, 側門螺絲孔卡死,無法開啟 3.轉掛測試機當，請常日協助確認，謝謝 | NO | NO | NO | YES | YES | NO
EQPRepairJ22104743 | T6373-152 | CWC7J0H67-01 | OverKill | JB2P2S09 | R63455-A1S-EZ-CP | NAHM72.00-B1C | 10 | too many O/S-A6873 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/05 11:26:34 | 2022/08/05 12:08:30 | 0.70 | OK | A6143 | 1.DUT2 O/S檢查針痕,貼下邊邊,CLEAN UNIT調整針位起測PASS. | NO | NO | NO | YES | YES | NO
EQPRepairJ22104745 | T6373-15 | CWC8J0189-01 | OverKill | JB6K8S03 | NC01682A(COG48AA)-A-AT-N2-C | GGS284.1-B1C | 15 | #14 EDA 秀I NO CORR | 05131 | Over Kill | O/S誤宰 | 2022/08/05 11:31:44 | 2022/08/05 11:58:07 | 0.44 | OK | 07764 | EDA單片 I , 確認針痕續測 #15(175.4) | NO | NO | NO | YES | YES | NO
EQPRepairJ22104752 | T6373-103 | CWC8J0137-01.001 | OverKill | JB1GVS15 | NC01682B(COG48AA)-A-AT-N-C | GGC155.1-221BC2 | 13 | bin29 fail-07764 | CPAUTO | Over Kill | Function誤宰 | 2022/08/05 12:17:52 | 2022/08/05 12:54:26 | 0.61 | OK | 07764 | bin29 fail 更換備卡p/c#17 pass #13(168.1)(17.1) | NO | NO | NO | YES | YES | NO
EQPRepairJ22104754 | T6373-13 |  | Tester Down | JB6H1S05 | NC01690F(COG46BA)-A-AT-C | GGH632.1-BC1 |  | 1.Tester keyboard NG,插拔Remote & USB 無效 2.開main frame側門插拔USB驗證, 側門螺絲孔卡死,無法開啟 3.轉掛測試機當，請常日協助確認，謝謝 | 20075 | Tester | Other | 2022/08/05 12:25:21 | 2022/08/05 14:21:33 | 1.94 | OK | 15362 | tester鍵盤無反應,請查日班協助確認,更換鍵盤轉接器後Ok | NO | NO | NO | YES | YES | NO
EQPRepairJ22104755 | T6373-15 | CWC8J0189-01 | Prober Down | JB6K8S03 | NC01682A(COG48AA)-A-AT-N2-C | GGS284.1-B1C | 16 | WAITIG FOR WHILE | 05131 | Prober | 面板Down (Monitor Down) | 2022/08/05 12:28:56 | 2022/08/05 13:05:05 | 0.60 | OK | 20230 | 改掛面板Down | NO | NO | NO | YES | YES | NO
EQPRepairJ22104759 | T6373-15 |  | Prober Down | JB6K8S03 | NC01682A(COG48AA)-A-AT-N2-C | GGS284.1-B1C |  | 面板當 | 20230 | Prober | 面板當 | 2022/08/05 13:05:48 | 2022/08/05 14:29:23 | 1.39 | OK | 20230 | 1.面板當 2.PROBER 重開OK,二分法調針RG=21,PM(45.6) 3.PASS | NO | NO | NO | YES | YES | NO
EQPRepairJ22104765 | T6373-11 | CWC8J0135-01.001 | OverKill | JB1GVS22 | NC01682A(COG48AA)-A-AT-N3-C | GGP621.1-BC21 | 08 | #6.7 EDA 2R o/s.... #8 too many o/s...please check(136.6) pm shift-A6920 | CPAUTO | Over Kill | 機台清針後Pass | 2022/08/05 13:45:46 | 2022/08/05 14:04:31 | 0.31 | OK | 20075 | 機台清針後Pass,range 12,mark 182,3 | NO | NO | NO | YES | YES | NO
EQPRepairJ22104775 | T6373-13 |  | Tester Down | JB6H1S05 | NC01690F(COG46BA)-A-AT-C | GGH632.1-BC1 |  | GP-IB-A6920 | CPAUTO | Tester | 離線(OCT/GPIB error) | 2022/08/05 14:30:37 | 2022/08/05 14:36:19 | 0.10 | OK | 20075 | data recover ok | NO | NO | NO | YES | YES | NO
EQPRepairJ22104778 | T6373-42 | CWC8J0118-01 | OverKill | JB1GVS20 | NC01682A(COG48AA)-A-AT-N3-C | GGR320.1-B1C1 | 12 | # 12 too many fail bin 29-A6143 | CPAUTO | Over Kill | BIN29 | 2022/08/05 14:46:03 | 2022/08/05 15:09:16 | 0.39 | OK | 13222 | BIN29 機台清針OK | NO | NO | NO | YES | YES | NO
EQPRepairJ22104780 | T6373-26 | CWC1J1A46-01.001 | OverKill | JB1NUS03 | ILIRJE6ADCM-2CP | QSR8P | 25 | TOO MANY O/S-A6564 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/05 15:09:43 | 2022/08/05 16:07:20 | 0.96 | OK | 20230 | 1.O/S 2.機台清針,重新對針,二分法調針RG=17,PM(47.6) 3.PASS | NO | NO | NO | YES | YES | NO
EQPRepairJ22104784 | T6373-42 | CWC8J0118-01 | OverKill | JB1GVS20 | NC01682A(COG48AA)-A-AT-N3-C | GGR320.1-B1C1 | 12 | # 12 continuous alarm fail bin 29-A6143 | CPAUTO | Over Kill | BIN29 | 2022/08/05 15:24:42 | 2022/08/05 15:55:41 | 0.52 | OK | 13222 | BIN29 更換配件#21 上機PASS 預溫 RANGE=18 (17.5) | NO | NO | NO | YES | YES | NO
EQPRepairJ22104789 | T6373-28 | CWC8J0221-01.001 | OverKill | JB1LWS17 | ILISLQPEDQD-2CP | ILA13191.1 | 16 | TOO MANY FAIL BIN40-A6564 | CPAUTO | Over Kill | BIN 40 | 2022/08/05 15:46:28 | 2022/08/05 17:13:15 | 1.45 | OK | 20075 | 1.換P/C,重開TESTER 無效 2.詢問LEADER,換板子k機,轉掛TESTER DOWN | NO | NO | NO | YES | YES | NO
EQPRepairJ22104791 | T6373-70 | CWC7J0M28-01.001 | OverKill | JB1Q4S15 | PA6542C | J683J | 19 | #19 PLS CHECK AND CONFIRM PM TOO BIG..TOO MANY OS-A6132 | CPAUTO | Over Kill | 機台清針後Pass | 2022/08/05 15:59:48 | 2022/08/05 16:25:37 | 0.43 | OK | 13222 | 機台清針後Pass | NO | NO | NO | YES | YES | NO
EQPRepairJ22104801 | T6373-135 | CWC7J0L69-01 | OverKill | VB4BKS01 | PA5129LS_12U80H-LT | P078M.1 | 19 | #19 連續2片EDA秀"R" o/s 過多 | 05131 | Over Kill | r | 2022/08/05 16:53:49 | 2022/08/05 17:04:47 | 0.18 | OK | 17229 | 1.查看eda o/s皆在同區域點測當片同區域穩定pass,檢查針痕ok5點pass. | NO | NO | NO | YES | YES | NO
EQPRepairJ22104804 | T6373-35 | CWC6J0996-01.001 | Corr Fail | JB1GVS16 | NC01682A(COG48AA)-A-AT-N3-C | GGG836.1-B1 | 02 | Corr Fail | A6660 | Corr | Other | 2022/08/05 16:58:09 | 2022/08/05 17:18:17 | 0.34 | OK | 07764 | pass | NO | NO | NO | YES | YES | NO
EQPRepairJ22104807 | T6373-11 | CWC8J0135-01.001 | Other | JB1GVS22 | NC01682A(COG48AA)-A-AT-N3-C | GGP621.1-BC21 | 09 | please check (5.2)pm shift.... already adjast up | A6920 | Over Kill | O/S誤宰 | 2022/08/05 17:09:38 | 2022/08/05 17:36:23 | 0.45 | OK | 15362 | 清調針,y range:12,p/m(46,6) | NO | NO | NO | YES | YES | NO
EQPRepairJ22104809 | T6373-28 | CWC8J0221-01.001 | Tester Down | JB1LWS17 | ILISLQPEDQD-2CP | ILA13191.1 | 14 | 1.換P/C,重開TESTER 無效 2.詢問LEADER,換板子k機,轉掛TESTER DOWN | 20075 | Tester | bin40 | 2022/08/05 17:14:15 | 2022/08/05 20:46:19 | 3.53 | OK | 17091 | 1.測試中跑bin40 function fail ,更換備卡與重新load程式無效 2.點測pass都變跑bin40 fail ,查看fail log 為PPS fail ,改掛Tester down 3.PPS#1 SWAP #2 , 執行ABS.INIT FAIL 顯示LCD#09 FAIL 4.LCD#09 SWAP #08 ,執行ABS.INIT LCD #06 Fail 5.單Cal Contact LCD#09 Fail，嘗試更換Diag Board Cal Contact PASS 6.重Cal ABS.PASS，上機PASS 二分法RANGE X:-16~10 Y:-6~18 PM(178，13) | NO | NO | NO | YES | YES | NO
EQPRepairJ22104816 | T6373-43 | CWC6J0527-01.004 | Corr Fail | JB1UJS03 | PA5129LS_12U80H-LT | P049S.1 | 24 | run corr ok-A6143 | CPAUTO | Corr | CORR FAIL | 2022/08/05 18:17:18 | 2022/08/05 18:44:22 | 0.45 | OK | 20230 | Corr Fail ---- 上機回收PASS，重新比對Corr Pass | NO | NO | NO | YES | YES | NO
EQPRepairJ22104821 | T6373-111 | CWC7J0761-01 | OverKill | TB1R8S74 | PA5137F-LT1 | P07HJ.1 | 19 | EDA SHOW 2H ALREADY CORR PASSED | A6858 | Over Kill | EDA SHOW 2H | 2022/08/05 19:04:19 | 2022/08/05 19:29:15 | 0.42 | OK | 10897 | corr pass先行續測 | NO | NO | NO | YES | YES | NO
EQPRepairJ22104829 | T6373-135 | CWC7J0L69-01 | OverKill | VB4BKS01 | PA5129LS_12U80H-LT | P078M.1 | 23 | too many o/s-A6941 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/05 19:29:29 | 2022/08/05 19:55:25 | 0.43 | OK | 20277 | Remote---- 1.針位貼下 Range:14 機台清針 回收PASS穩定，還機生產 | NO | NO | NO | YES | YES | NO
EQPRepairJ22104830 | T6373-12 | CWC8J0136-01 | OverKill | JB6K8S02 | NC01682A(COG48AA)-A-AT-N3-C | GGP622.1-BC1 | 25 | TOO MANY FAIL BIN 26-A6810 | CPAUTO | Over Kill | Function | 2022/08/05 19:40:17 | 2022/08/05 19:55:58 | 0.26 | OK | 10897 | BIIN26不過，驗證PASS die P->P F->F點測多顆相同，清調針無效，疑似產品問題續測觀察 | NO | NO | NO | YES | YES | NO
EQPRepairJ22104848 | T6373-128 | CWC7J0M33-01.001 | OverKill | VB1BVS01 | PA5270D-LT | U38F8.00 | 18 | fail bin too many-A6318 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/05 20:59:17 | 2022/08/06 00:53:07 | 3.90 | OK | 20182 | 1.請配件室確認低溫封口是否正常，重新鎖緊後烘烤 2.Bin31 Leakage 下卡烘烤，接手上機PASS 3.二分法RG-10 PM(156.1) | NO | NO | NO | YES | YES | NO
EQPRepairJ22104863 | T6373-34 | CWC7J0K25-01.001 | OverKill | JB1BYS10 | PA6587B_15U90H | J696F | 08 | too many o/s-A6898 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/05 21:59:39 | 2022/08/05 22:19:26 | 0.33 | OK | 17091 | Remote----機台清針 PASS | NO | NO | NO | YES | YES | NO
EQPRepairJ22104869 | T6373-28 | CWC8J0221-01.001 | Prober Down | JB1LWS15 | ILISLQPEDQD-2CP | ILA13191.1 | 17 | pre alignment error LOAD PORT1 SAFETY ERROR-A6675 | CPAUTO | Prober | loadport | 2022/08/05 22:41:50 | 2022/08/05 22:49:12 | 0.12 | OK | 17091 | init ok確認eda正常 | NO | NO | NO | YES | YES | NO
EQPRepairJ22104873 | T6373-103 | CWC8J0137-01.001 | OverKill | JB1GVS17 | NC01682B(COG48AA)-A-AT-N-C | GGC155.1-221BC2 | 18 | ALWAYS ALARM OS PC HAS PARTICLE | A6142 | Over Kill | 配件殘金 | 2022/08/05 23:09:02 | 2022/08/05 23:47:37 | 0.64 | OK | 15492 | 1.下卡磨針，二分法(-10，-15)、RANGE:10 2.P/M(181，4)OK | NO | NO | NO | YES | YES | NO
EQPRepairJ22104881 | T6373-151 | CWC8J0038-01 | OverKill | JB2P2S08 | R63455-A1S-EZ-CP | NAPY65.00-B1C | 01 | too many fail bin-A6941 | CPAUTO | Over Kill | Function誤宰 | 2022/08/05 23:38:06 | 2022/08/06 00:31:43 | 0.89 | OK | 15492 | 1.重測批，O/S 針髒，清針 PASS 2.#02 區域性BIN35 ，續測待換批觀察 | NO | NO | NO | YES | YES | NO
EQPRepairJ22104887 | T6373-70 | CWC7J0M28-01.001 | OverKill | JB1Q4S15 | PA6542C | J683J | 25 | o/s-03491 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/06 00:00:36 | 2022/08/06 00:15:11 | 0.24 | OK | 15492 | REMOTE--- 1.零星O/S，清針 PASS 2.P/M(128，5)OK | NO | NO | NO | YES | YES | NO
EQPRepairJ22104888 | T6373-42 | CWC8J0135-01.003 | OverKill | JB1GVS21 | NC01682A(COG48AA)-A-AT-N3-C | GGP621.1-BC12 | 19 | too many fail bin..-A6941 | CPAUTO | Over Kill | Function誤宰 | 2022/08/06 00:09:12 | 2022/08/06 00:45:48 | 0.61 | OK | 20277 | Remote---- 1.機台清針 Range:14 點測PASS穩定，還機生產待觀察 | NO | NO | NO | YES | YES | NO
EQPRepairJ22104894 | T6373-70 | CWC7J0M28-01.001 | OverKill | JB1Q4S15 | PA6542C | J683J | 25 | o/s-03491 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/06 00:35:44 | 2022/08/06 01:14:05 | 0.64 | OK | 15492 | 1.下卡磨針，上機PASS，RANGE:8 | NO | NO | NO | YES | YES | NO
EQPRepairJ22104897 | T6373-47 | CWC8J0125-01.002 | Probe Card Down | JB1HKS08 | NA01687G(COG48BA)-A-AT-N2-C | GGP242.1-B1C221 | 17 | JB1HKS08 超過pm次數 | 03405 | Probe Card | 定期PM | 2022/08/06 00:52:55 | 2022/08/06 01:25:11 | 0.54 | OK | 17091 | PM 更換配件#09 PASS PM(170，3) | NO | NO | NO | YES | YES | NO
EQPRepairJ22104902 | T6373-33 | CWC7J0A59-01 | OverKill | JB1TAS14 | PA5208KF2 | J67LK | 02 | too many o/s-A6810 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/06 01:15:06 | 2022/08/06 01:42:27 | 0.46 | OK | 17184 | 1.清針 | NO | NO | NO | NO | YES | NO
EQPRepairJ22104908 | T6373-37 | CWC8J0136-01.001 | OverKill | JB1GVS26 | NC01682A(COG48AA)-A-AT-N3-C | GGP622.1-BC2 | 01 | too many fail bin-A6898 | CPAUTO | Over Kill | 配件殘金 | 2022/08/06 01:46:26 | 2022/08/06 03:00:24 | 1.23 | OK | 17091 | 夏卡磨針 上機pass r:-10~4 pm(137，6) | NO | NO | NO | YES | YES | NO
EQPRepairJ22104911 | T6373-151 | CWC8J0038-01 | OverKill | JB2P2S08 | R63455-A1S-EZ-CP | NAPY65.00-B1C | 02 | EDA yields < 40% | A6318 | Over Kill | EDA yields | 2022/08/06 02:17:40 | 2022/08/06 02:52:08 | 0.57 | OK | 20182 | 1.查看EDA 其它機台不同批有相同問題，續測下一片待觀察 | NO | NO | NO | YES | YES | NO
EQPRepairJ22104920 | T6373-12 |  | Corr Fail | JB1GVS20 | NC01682A(COG48AA)-A-AT-N3-C | GGG836.1-B1 |  | corr fail | A6938 | Corr | corr fail | 2022/08/06 02:53:44 | 2022/08/06 02:56:43 | 0.05 | OK | 10897 | 這次良率較高判定PASS | NO | NO | NO | NO | YES | NO
EQPRepairJ22104927 | T6373-12 | CRC7J0009-01 | Prober Down | TB1QIS01 | ILI8LFZAEEB-2CP | ILR31190.1 | 25 | model not found (rough alignment) | A6446 | Prober | Wafer Edge/Mode Not Found | 2022/08/06 03:25:53 | 2022/08/06 03:40:18 | 0.24 | OK | 10897 | Mode Not Found/check die/針位確認 | NO | NO | NO | NO | YES | NO
EQPRepairJ22104932 | T6373-12 | CRC7J0009-01 | Tester Down | JB1GVS20 | NC01682A(COG48AA)-A-AT-N3-C | GFK688.1-BC12111C | 25 | AlarmCode:665 AlarmMessage:GP-IB STOP COMMAND RECEIVED !! | CPAUTO | Tester | 離線(OCT/GPIB error) | 2022/08/06 03:48:39 | 2022/08/06 04:06:41 | 0.30 | OK | 17091 | REBOOT OK | NO | NO | NO | YES | YES | NO
EQPRepairJ22104936 | T6373-128 | CWC7J0M33-01.001 | OverKill | VB1BVS01 | PA5270D-LT | U38F8.00 | 18 | Retest Fail-> O/S | A6318 | Probe Card | P/C Down---- | 2022/08/06 04:03:01 | 2022/08/06 04:34:13 | 0.52 | OK | 20277 | P/C Down----P/C:VB1BVS01送修無備卡掛配件當 1.EDA片片Show H Bin31 2.先前已下卡烘烤，仍無法久測 3.故配件送修無備卡配件當 | NO | NO | NO | YES | YES | NO
EQPRepairJ22104946 | T6373-128 |  | Probe Card Down | VB1BVS01 | PA5270D-LT | U38F8.00 |  | P/C Down----P/C:VB1BVS01送修無備卡掛配件當 1.EDA片片Show H Bin31 2.先前已下卡烘烤，仍無法久測 3.故配件送修無備卡配件當 | 20277 | Probe Card | 配件當 | 2022/08/06 04:34:38 | 2022/08/06 06:17:48 | 1.72 | OK | 20277 | 輪工處理---- 1.配件為修OK 上機PASS Range:14 | NO | NO | NO | YES | YES | NO
EQPRepairJ22104950 | T6373-86 | CWC7J0J58-01.004 | OverKill | JB1Q4S16 | PA6542C | J683L | 19 | too many os | A6446 | Over Kill | O/S誤宰 | 2022/08/06 05:15:46 | 2022/08/06 05:42:47 | 0.45 | OK | 17184 | 1.清針&調針(43.5) | NO | NO | NO | NO | YES | NO
EQPRepairJ22104977 | T6373-34 | CWC7J0K25-01.001 | OverKill | JB1BYS10 | PA6587B_15U90H | J696F | 12 | Too many o/s | A6898 | Over Kill | O/S誤宰 | 2022/08/06 06:51:19 | 2022/08/06 06:59:21 | 0.13 | OK | 15492 | 1.清、調針 PASS 2.P/M(132，7)OK | NO | NO | NO | YES | YES | NO
EQPRepairJ22104982 | T6373-38 | CWC6J0815-01.001 | OverKill | JB1GKS09 | PA6543B | J6630 | 17 | eda 2h, already corr passed | A6898 | Over Kill | 2H | 2022/08/06 07:13:42 | 2022/08/06 08:36:17 | 1.38 | OK | 07910 | 2H，下卡磨針 取卡上機PASS P/M(172.5)#18 RY8 | NO | NO | NO | YES | YES | NO
EQPRepairJ22104994 | T6373-28 |  | Prober Down | JB1LWS15 | ILISLQPEDQD-2CP | ILA13191.1 |  | loadport error-A6962 | CPAUTO | Prober | Loader Down | 2022/08/06 08:17:43 | 2022/08/06 08:22:53 | 0.09 | OK | 08774 | INIT OK | NO | NO | NO | YES | YES | NO
EQPRepairJ22104999 | T6373-103 | CWC8J0118-01.001 | Prober Down | JB1GVS17 | NC01682A(COG48AA)-A-AT-N3-C | GGR320.1-B1C21 | 14 | AlarmCode:4827 AlarmMessage:ARM @@VACUUM ERROR !! | CPAUTO | Prober | 手臂異常(Arm Error) | 2022/08/06 08:27:52 | 2022/08/06 09:12:49 | 0.75 | OK | 08774 | 1. #14測試中ARM1 VACUUM ERROR ，重開ok。 2. 燒CODE型號無法接測，請協助將Wafer #14 Hold To TPD。 3. 請產線&DN協助後續處置，謝謝!! | NO | NO | NO | YES | YES | NO
EQPRepairJ22105000 | T6373-136 | CWC8J0250-01 | OverKill | JB6H1S04 | NC01690G(COG48AA)-A-AT-N4-C | GGM127.1-BC1 | 24 | o/s-17436 | CPAUTO | Over Kill | 配件殘金 | 2022/08/06 08:29:58 | 2022/08/06 09:29:38 | 0.99 | OK | 17712 | S04殘金，更換S06，#24(164.4)，確認針痕針壓OK | NO | NO | NO | YES | YES | NO
EQPRepairJ22105002 | T6373-128 | CWC7J0M33-01.001 | OverKill | VB1BVS01 | PA5270D-LT | U38F8.00 | 19 | #19 O/S過多 | 19617 | Over Kill | O/D不足 | 2022/08/06 08:34:37 | 2022/08/06 08:47:53 | 0.22 | OK | 17712 | 溫針重對，#19(40.5)rg16 | NO | NO | NO | YES | YES | NO
EQPRepairJ22105005 | T6373-111 | CWC7J0J51-01.001 | OverKill | TB1R8S74 | PA5137F-LT1 | P094H.1 | 22 | TOO MANY O/S-A6894 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/06 08:48:26 | 2022/08/06 09:53:43 | 1.09 | OK | 20231 | 1.首片測試全面BIN2 open，固定pin不過，調針位、加針壓、重新對針無效，重MNT後變為其他pin 2.重MNT 8次pass，二分法座標(83,3)，rg:8，，PM(144,1) | NO | NO | NO | YES | YES | NO
EQPRepairJ22105018 | T6373-103 | CWC8J0118-01.001 | Prober Down | JB1GVS17 | NC01682A(COG48AA)-A-AT-N3-C | GGR320.1-B1C21 | 15 | NEEDLE POSITION-A6921 | CPAUTO | Prober | Other | 2022/08/06 09:22:11 | 2022/08/06 09:34:30 | 0.21 | OK | 08774 | 重新對針S/U 138.6 | NO | NO | NO | YES | YES | NO
EQPRepairJ22105034 | T6373-15 |  | Setup Fail | JB6K8S04 | NC01682A(COG48AA)-A-AT-N3-C | GGR319.1-B |  | setup fail , already change p/c due to o/s fail JB6K8S04 (DUT2 O/S) JB6K8S01 (DUT1+4 O/S)-11933 | CPAUTO | Setup Fail | O/S Fail | 2022/08/06 10:06:36 | 2022/08/06 10:17:29 | 0.18 | OK | 17712 | Dut1 4 OS，重MNT P/C Pass | NO | NO | NO | YES | YES | NO
EQPRepairJ22105060 | T6373-128 | CWC7J0M33-01.001 | OverKill | VB1BVS01 | PA5270D-LT | U38F8.00 | 20 | #20 已依誤宰處理流程處置無效 (Bin31 過多) | 19617 | Over Kill | Function誤宰 | 2022/08/06 11:38:17 | 2022/08/06 12:17:11 | 0.65 | OK | 17712 | Bin31 Leak，夜班已烘烤上機仍測試不穩，再次送修無備卡 | NO | NO | NO | YES | YES | NO
EQPRepairJ22105062 | T6373-15 | CWC7J0797-01.001 | Corr Fail | JB6K8S04 | NC01682A(COG48AA)-A-AT-N3-C | GGR319.1-B | 14 | corr-11933 | CPAUTO | Corr | Corr Fail | 2022/08/06 11:53:06 | 2022/08/06 11:56:48 | 0.06 | OK | 17712 | 比對Corr Wafer測試較佳PASS | NO | NO | NO | YES | YES | NO
EQPRepairJ22105064 | T6373-11 | CWC8J0135-01.004 | OverKill | JB1GVS22 | NC01682A(COG48AA)-A-AT-N3-C | GGP621.1-BC112 | 02 | too many os-A6957 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/06 11:56:21 | 2022/08/06 13:16:24 | 1.33 | OK | 13749 | 1.SLOT # o/s 過多不穩 針尖黑點+金凸，更換備卡#S22->#S23 OK. 2.預溫重對針，點測PASS 溫針退針壓 OK 3.pm(45.6) x=0 y=7 | NO | NO | NO | YES | YES | NO
EQPRepairJ22105071 | T6373-128 |  | Probe Card Down | VB1BVS01 | PA5270D-LT | U38F8.00 |  | Bin31 Leak 夜班已烘烤，送修無備卡 | 17712 | Probe Card | Function Fail | 2022/08/06 12:18:04 | 2022/08/06 15:53:30 | 3.59 | OK | 15897 | 改測接手.溫針重對.Pass.Run CORR. | YES | NO | NO | YES | YES | NO
EQPRepairJ22105080 | T6373-43 | CWC7J0811-01.008 | OverKill | JB1UJS03 | PA5129LS_12U80H-LT | P060L.1 | 04 | o/s-20022 | CPAUTO | Over Kill | 機台清針後Pass | 2022/08/06 13:10:23 | 2022/08/06 13:29:15 | 0.31 | OK | 17712 | 機台清針，觀察一段ok，#4(139.4)rg10 | NO | NO | NO | YES | YES | NO
EQPRepairJ22105085 | T6373-13 | CWC8J0251-01 | OverKill | JB6H1S05 | NC01690I(COG46BA)-A-AT-C | GGS960.1-BC | 14 | Continuous 2 wafers EDA shows "H" O/S | A6957 | Over Kill | 2H | 2022/08/06 13:25:21 | 2022/08/06 13:46:01 | 0.34 | OK | 13749 | 1.Slot #17 EDA 多片Show H 為BIN25 Function & BIN83 Short 過多穩定 造成 PP、FF 區域性. 2.請產線先型Run Corr or Change Lot Test Check，以利釐清問題點. | NO | NO | NO | YES | YES | NO
EQPRepairJ22105092 | T6373-141 | CWC8J0124-01.001 | OverKill | JB1HKS06 | NA01687G(COG48BA)-A-AT-N2-C | GGS268.1-B2C2 | 18 | 2R O/S | A6962 | Over Kill | 2R | 2022/08/06 13:37:07 | 2022/08/06 13:57:49 | 0.35 | OK | 07910 | 2R 檢視針況正常，機台清針 PASS P/M(171.3)#18 RY14 | NO | NO | NO | YES | YES | NO
EQPRepairJ22105126 | T6373-135 | CWC7J0M32-01.002 | OverKill | VB4BKS01 | PA5129LS_12U80H-LT | P060Q.1 | 02 | o/s-dut 4-12254 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/06 15:30:20 | 2022/08/06 15:53:14 | 0.38 | OK | 17680 | 調整針位range:13um,position(0,14),mark(115,5)，OD-10um，起測穩定 | NO | NO | NO | YES | YES | NO
EQPRepairJ22105129 | T6373-16 | CMC7J0212-01 | OverKill | JB1TAS25 | PA5208K-LT | QT79P | 09 | o/s-20022 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/06 15:42:46 | 2022/08/06 16:04:47 | 0.37 | OK | 17680 | 零星o/s，clean unit，起測穩定 | NO | NO | NO | YES | YES | NO
EQPRepairJ22105131 | T6373-12 |  | Setup Fail | JB8KZS14 | ND01657B(COF94AA)-A-C | S0V457.00-B |  | too large needle --wait engr repair-A6668 | CPAUTO | Prober | Other | 2022/08/06 15:45:07 | 2022/08/06 16:10:28 | 0.42 | OK | 08774 | ND01657 改機自動對針NG 約14根PIN 1.重對無效，確認差距過大，高溫 2.BY PASS 後LOAD WAFER進去，再重新對針1PIN NG 3.溫針，再次重新對針ok 續掛s/u | NO | NO | NO | YES | YES | NO
EQPRepairJ22105133 | T6373-47 | CWC8J0125-01 | OverKill | JB1HKS09 | NA01687G(COG48BA)-A-AT-N2-C | GGP242.1-B1C11 | 22 | O/S-20022 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/06 15:48:19 | 2022/08/06 16:06:09 | 0.30 | OK | 17680 | remote----換片o/s,調整針位position(15,12),mark(36,0)，OD+5um，點測穩定 | NO | NO | NO | YES | YES | NO
EQPRepairJ22105134 | T6373-28 | CWC8J0221-01.001 | OverKill | JB1LWS15 | ILISLQPEDQD-2CP | ILA13191.1 | 25 | #25 fail bin 23-20021 | CPAUTO | Probe Card | Function Fail更換P/C | 2022/08/06 15:49:57 | 2022/08/06 16:34:20 | 0.74 | OK | 08774 | Function Fail更換P/C 104.6 | NO | NO | NO | YES | YES | NO
EQPRepairJ22105135 | T6373-128 |  | Other | VB1BVS01 | PA5270D-LT | U38F8.00 |  | CORR 不為N.無法掛CORR.故MBD RUN... | 15897 | Over Kill | Function誤宰 | 2022/08/06 15:57:00 | 2022/08/06 16:06:32 | 0.16 | OK | 15897 | 輪工處理---- 1.PA5270D-LT/CP1_Bin31改測PA6587B/CP3.疑似機台問題. 2.CORR 1st.連續性Bin12 Function.點測、回收皆Pass.還機生產續測觀察. | NO | NO | NO | YES | YES | NO
EQPRepairJ22105147 | T6373-121 |  | Setup Fail | JB1QNS01 | NA01687G(COG48BA)-A-AT-X1-C | GGS268.1-B1 |  | 使用改機自動化tester產生異常訊息Error writing "stdout" I/O error | 08341 | Other | Auto Setup App NG | 2022/08/06 16:27:30 | 2022/08/06 16:45:12 | 0.30 | OK | 13749 | 1.產線改機使用Auto Setup App NG (Error writing "stdout" I/O error. 2.輪工確認關閉程式後重新開啟Auto Setup APP 無效. 3.使用cbsetup Load 正常，還機產線 先行使用cbseup | NO | NO | NO | YES | YES | NO
EQPRepairJ22105156 | T6373-141 | CWC8J0124-01.001 | Prober Down | JB1HKS06 | NA01687G(COG48BA)-A-AT-N2-C | GGS268.1-B2C21 | 21 | AlarmCode:8732 AlarmMessage:It failed in lock/unlock of the manual ring lever. | CPAUTO | Prober | lock/unlock ring | 2022/08/06 17:02:18 | 2022/08/06 17:15:33 | 0.22 | OK | 07910 | #21 多次lock/unlock ring，退片調整 2.STDF 型號 #21 重新上貨 | NO | NO | NO | YES | YES | NO
EQPRepairJ22105158 | T6373-128 | CWC6J0A73-01.005 | OverKill | JB1BYS11 | PA6587B_15U90H | J65P5 | 12 | #12 已依誤宰處理流程處置無效 (O/S 過多) | 19617 | Over Kill | 配件有異物 | 2022/08/06 17:11:45 | 2022/08/06 18:23:24 | 1.19 | OK | 15897 | 下卡請配件室清針.溫針重對Range:16um.#12(19.6). | NO | NO | NO | YES | YES | NO
EQPRepairJ22105163 | T6373-141 | CWC8J0124-01.001 | Other | JB1HKS06 | NA01687G(COG48BA)-A-AT-N2-C | GGS268.1-B2C21 | 21 | Other | 20021 | Prober | LOCK/UNLOCK RING | 2022/08/06 17:27:55 | 2022/08/06 18:34:17 | 1.11 | OK | 07910 | 1.調整檔片位置 OK，重下發現殘金下卡磨針 2.取卡上機PASS，P/M(172.1)#18 RY6 | NO | NO | NO | YES | YES | NO
EQPRepairJ22105173 | T6373-38 | CWC6J0815-01.001 | OverKill | JB1GKS09 | PA6543B | J6630 | 23 | #21,22 EDA SHOW 2H-20021 | CPAUTO | Over Kill | 2H | 2022/08/06 17:56:46 | 2022/08/06 18:06:29 | 0.16 | OK | 08774 | 2H 1.[CUSTOMER_HOLD_CRITERIA] [HW_BIN17>1%](Bin 18 > 11EA) 15(Hold Wafer) [HW_BIN20>0.3%](Bin 21 > 3EA) 5(Hold Wafer) [HW_BIN21>1%](Bin 22 > 11EA) 42(Hold Wafer) [HW_BIN22>0.5%](Bin 23 > 5EA) 13(Hold Wafer) [HW_BIN60<95.5%](Bin 61 < 1081EA) 966(Hold Wafer) [HW_BIN8>0.3%](Bin 9 > 3EA) 10(Hold Wafer) 2.續測RUN CORR | NO | NO | NO | YES | YES | NO
EQPRepairJ22105199 | T6373-16 | CMC7J0212-01 | OverKill | JB1TAS25 | PA5208K-LT | QT79P | 10 | O/S-20022 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/06 19:13:24 | 2022/08/06 20:13:46 | 1.01 | OK | 17674 | 1. 更換備卡#19 PASS， 溫針 RANGE Y=10 (68,6)#10 | NO | NO | NO | YES | YES | NO
EQPRepairJ22105209 | T6373-29 | CWC8J0135-01.002 | OverKill | JB1GVS19 | NC01682A(COG48AA)-A-AT-N3-C | GGP621.1-BC22 | 13 | EDA SHOW 2H (NO CORR)-19697 | CPAUTO | Over Kill | Show H | 2022/08/06 20:15:11 | 2022/08/06 21:55:15 | 1.67 | OK | 17542 | Show H，Bin8 Function Fail過多，更換配件#18，上機PASS，range:16，試打(17,5)無異常，點測pass穩定，OD:15 | NO | NO | YES | YES | YES | NO
EQPRepairJ22105211 | T6373-16 | CMC7J0212-01 | Other | JB1TAS25 | PA5208K-LT | QT79P | 10 | p/c not available | A6336 | Other | Other | 2022/08/06 20:21:08 | 2022/08/06 20:48:48 | 0.46 | OK | 17674 | 1. P/C #19 ABN單未RLS，更換備卡#22 PASS，溫針 RANGE Y=10 (72,6)#10 | NO | NO | NO | YES | YES | NO
EQPRepairJ22105236 | T6373-135 |  | Prober Down | JB1UJS13 | PA5129LS_12U80H-LT | P062Y.1 |  | cannot measure index size-A6148 | CPAUTO | Prober | Wafer Edge/Mode Not Found | 2022/08/06 21:29:54 | 2022/08/06 21:35:16 | 0.09 | OK | 11348 | 重抓OK | NO | NO | YES | YES | YES | NO
EQPRepairJ22105240 | T6373-111 | CWC7J0829-01 | OverKill | TB1R8S74 | PA5137F-LT1 | P07NW.1 | 01 | TOO MANY OPEN SHORT-A6786 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/06 21:42:32 | 2022/08/06 21:59:26 | 0.28 | OK | 20276 | OD過重 15um>10um range 10um P/M(118.0) | NO | NO | NO | YES | YES | NO
EQPRepairJ22105244 | T6373-38 | CWC6J0815-01.001 | OverKill | JB1GKS09 | PA6543B | J6630 | 24 | too many OS.. PC has particle-A6925 | CPAUTO | Over Kill | 配件卡毛 | 2022/08/06 21:48:39 | 2022/08/06 22:44:11 | 0.93 | OK | 03426 | 更換配件S12 | NO | NO | NO | YES | YES | NO
EQPRepairJ22105249 | T6373-29 | CWC8J0135-01.002 | OverKill | JB1GVS18 | NC01682A(COG48AA)-A-AT-N3-C | GGP621.1-BC22 | 13 | O/S TOO MANY-19697 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/06 22:30:23 | 2022/08/07 00:33:13 | 2.05 | OK | 17542 | 1.測試中O/S不過，確認針痕時發生wafer表面有小冰晶 2.機台升溫 -40度 >> 100度，11:00到達溫度，11:30降回-40度 3.重對針，點測PASS，Data-In後，Model由 模糊 >> 清晰，確認點測PASS高度不同 4.重對針，點測PASS，Data-In後，Model正常，確認點測PASS高度相同，range:16，試打(18,5)無異常，還機待觀察 | NO | NO | YES | YES | YES | NO
EQPRepairJ22105256 | T6373-38 | CWC6J0815-01.001 | Probe Card Down | JB1GKS09 | PA6543B | J6630 | 24 | 此P/C無法MOVE IN | 19697 | Probe Card | 配件未release | 2022/08/06 22:51:15 | 2022/08/06 23:47:23 | 0.94 | OK | 17542 | 配件未release --- 1.更換配件#08，上機pass，range:12，試打(162,1)無異常，起測pass | NO | NO | NO | YES | YES | NO
EQPRepairJ22105258 | T6373-42 |  | Prober Down | JB6H1S07 | NC01690J(COG48AA)-A-AT-N5-C | GGH883.1-BC |  | AlarmCode:1057 AlarmMessage:MODEL NOT FOUND ( ROUGH ALIGNMENT ) !! | CPAUTO | Prober | Wafer Edge/Mode Not Found | 2022/08/06 22:59:29 | 2022/08/06 23:16:07 | 0.28 | OK | 15096 | Remote處理----重抓model ok ，check die ok. | NO | NO | NO | YES | YES | NO
EQPRepairJ22105272 | T6373-34 | CWC7J0G36-01.002 | OverKill | JB1BYS10 | PA6587B_15U90H | J696A | 22 | too many os.. no particle-A6925 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/06 23:55:57 | 2022/08/07 01:24:01 | 1.47 | OK | 17542 | 1.下卡磨針，上機點測發生Online INIT 2.-50um溫針重對針，range:10，試打(135,0)無異常，起測一段PASS穩定 | NO | NO | NO | YES | YES | NO
EQPRepairJ22105276 | T6373-141 | CWC8J0124-01.001 | OverKill | JB1HKS06 | NA01687G(COG48BA)-A-AT-N2-C | GGS268.1-B2C21 | 25 | TOO MANY OS.. NO PARTICLE-A6925 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/07 00:29:29 | 2022/08/07 01:30:03 | 1.01 | OK | 03426 | 更換配件S08 O/S Fail 機台清針 重MNT 換Pin無效 更換S10 PASS | NO | NO | NO | YES | YES | NO
EQPRepairJ22105281 | T6373-147 | CWC7J0H96-01 | OverKill | TB1R8S77 | PA5137F-LT1 | P094J.1 | 13 | too many o/s-A6148 | CPAUTO | Over Kill | 機台清針後Pass | 2022/08/07 00:45:03 | 2022/08/07 01:10:24 | 0.42 | OK | 17674 | 1. 機台清針後Pass (115,0)#13 | NO | NO | NO | YES | YES | NO
EQPRepairJ22105283 | T6373-21 | CWC7J0J52-01 | OverKill | JB1TAS28 | PA5208KF-LT | J68QW | 22 | too many os.. no particle-A6925 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/07 00:48:40 | 2022/08/07 01:47:14 | 0.98 | OK | 17542 | 更換配件#26，上機PASS，溫針重對針，range:16，試打(147,5)無異常，起測PASS穩定 | NO | NO | NO | YES | YES | NO
EQPRepairJ22105285 | T6373-47 | CWC8J0134-01.002 | OverKill | JB1HKS09 | NA01687G(COG48BA)-A-AT-N2-C | GGS940.1-BC22 | 17 | o/s,od:13p/m(13.1)-A6336 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/07 00:51:14 | 2022/08/07 01:39:41 | 0.81 | OK | 17674 | 1. 退針壓pass，重對針，#17 (92,2)溫針 range y=10 (144,0) | NO | NO | NO | YES | YES | NO
EQPRepairJ22105287 | T6373-152 | CWC7J0H67-01 | OverKill | JB2P2S09 | R63455-A1S-EZ-CP | NAHM72.00-B1C | 15 | CON TINUOUS EDA SHOW 2H NO CORR WAFER-A6881 | CPAUTO | Over Kill | 2H | 2022/08/07 00:51:53 | 2022/08/07 01:15:25 | 0.39 | OK | 17674 | 1. 2H 不同FAIL BIN 過多，不同批不同機卡相同問題，NAHM75.00-2、NAPY65.00 | NO | NO | NO | YES | YES | NO
EQPRepairJ22105301 | T6373-128 | CWC7J0K25-01.002 | OverKill | JB1BYS11 | PA6587B_15U90H | J696F | 17 | too manyo/s 161,1 pm cheked up/down od11-A6881 | CPAUTO | Over Kill | 機台清針後Pass | 2022/08/07 01:40:54 | 2022/08/07 01:53:04 | 0.20 | OK | 17674 | Remote 處理---- 1. 機台清針後Pass (117,1)#17 | NO | NO | NO | YES | YES | NO
EQPRepairJ22105303 | T6373-70 | CWC8J0217-01 | OverKill | JB1Q4S15 | PA6542C | J68S2 | 20 | too many o/s-A6148 | CPAUTO | Over Kill | 配件殘金 | 2022/08/07 01:45:00 | 2022/08/07 02:49:34 | 1.08 | OK | 17674 | 1. OS 配件殘金 (金絲)，清針參數依 Memo單設定，配件下卡磨針pass，溫針、重對針再溫針 range y=10 (66,4)#20 | NO | NO | NO | YES | YES | NO
EQPRepairJ22105314 | T6373-141 |  | Other | JB1HKS10 | NA01687G(COG48BA)-A-AT-N2-C | GGS268.1-B2C21 |  | CAL機 | 19697 | Tester | Cal | 2022/08/07 02:26:51 | 2022/08/07 02:41:24 | 0.24 | OK | 03426 | Online K機，K機後PASS | NO | NO | NO | YES | YES | NO
EQPRepairJ22105322 | T6373-141 | CWC8J0124-01.001 | Tester Down | JB1HKS10 | NA01687G(COG48BA)-A-AT-N2-C | GGS268.1-B2C21 | 16 | AlarmCode:665 AlarmMessage:GP-IB STOP COMMAND RECEIVED !! | CPAUTO | Tester | 離線(OCT/GPIB error) | 2022/08/07 02:54:10 | 2022/08/07 02:58:02 | 0.06 | OK | 17542 | 離線(OCT/GPIB error) --- T44 STDF，起測異常，wafer alignment退片重上 | NO | NO | NO | YES | YES | NO
EQPRepairJ22105324 | T6373-47 | CWC8J0134-01.002 | OverKill | JB1HKS09 | NA01687G(COG48BA)-A-AT-N2-C | GGS940.1-BC22 | 18 | too many o/s 170.3 pm checked up/down-A6881 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/07 03:00:10 | 2022/08/07 03:48:36 | 0.81 | OK | 17674 | 1. 更換備卡#07 PASS， #18 (92,2)溫針 RANGE Y=10 (34,4) | NO | NO | NO | YES | YES | NO
EQPRepairJ22105333 | T6373-128 | CWC7J0K25-01.002 | OverKill | JB1BYS11 | PA6587B_15U90H | J696F | 18 | CONTINUOUS EDA SHOW 2R O/S-A6881 | CPAUTO | Over Kill | 2R | 2022/08/07 03:50:37 | 2022/08/07 04:16:07 | 0.42 | OK | 17674 | 1. 2R O/D不足，重對針、溫針 PASS RANGE Y=8 (52,0)#18，還機觀察一片 | NO | NO | NO | YES | YES | NO
EQPRepairJ22105343 | T6373-34 | CWC7J0G36-01.002 | OverKill | JB1BYS10 | PA6587B_15U90H | J696A | 23 | too many o/s-A6924 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/07 04:25:46 | 2022/08/07 04:35:16 | 0.16 | OK | 17542 | 機台清針，起測一段PASS穩定 | NO | NO | NO | YES | YES | NO
EQPRepairJ22105350 | T6373-37 |  | Prober Down | JB1GVS26 | NC01682A(COG48AA)-A-AT-N3-C | GGP622.1-BC2 |  | AlarmCode:1060 AlarmMessage:CANNOT MEASURE INDEX SIZE !! | CPAUTO | Prober | Wafer Edge/Mode Not Found | 2022/08/07 04:52:17 | 2022/08/07 05:00:54 | 0.14 | OK | 17542 | Wafer Edge/Mode Not Found -- Re-start OK，Check Die OK，P/M(17,5)OK | NO | NO | NO | YES | YES | NO
EQPRepairJ22105360 | T6373-136 | CWC8J0295-01 | OverKill | JB6H1S06 | NC01690G(COG48AA)-A-AT-N4-C | GGG486.1-BC | 12 | #9 eda show I ( no corr wafer )-A6148 | CPAUTO | Over Kill | Show H | 2022/08/07 05:52:36 | 2022/08/07 06:08:31 | 0.27 | OK | 03426 | Show H 更換配件S05 | NO | NO | NO | YES | YES | NO
EQPRepairJ22105362 | T6373-39 | CWC7J0M27-01.001 | OverKill | JB1BYS12 | PA6587B_15U90H | J696G | 16 | too many o/s 6.5 pm check up down OD 13-A6881 | CPAUTO | Over Kill | 機台清針後Pass | 2022/08/07 05:53:52 | 2022/08/07 06:07:27 | 0.23 | OK | 17542 | 機台清針後Pass | NO | NO | NO | YES | YES | NO
EQPRepairJ22105387 | T6373-27 | CWC8J0271-01 | OverKill | JB8Y9S02 | ILISLQPEDQD-2CP | ILA13192.1 | 06 | o/s too many-19697 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/07 07:17:24 | 2022/08/07 07:36:28 | 0.32 | OK | 07910 | 機台清針 pass | NO | NO | NO | YES | YES | NO
EQPRepairJ22105389 | T6373-39 | CWC7J0M27-01.001 | OverKill | JB1BYS12 | PA6587B_15U90H | J696G | 17 | Continuous wafers EDA shows "I" | A6881 | Over Kill | 2I | 2022/08/07 07:18:59 | 2022/08/07 08:14:29 | 0.93 | OK | 15897 | O/S Mode teset.check ok.測試狀況不佳.黑點磨針溫針重對Range:14um.#17(19.6). | NO | NO | NO | YES | YES | NO
EQPRepairJ22105397 | T6373-147 | CWC7J0L14-01.002 | Prober Down | TB1R8S77 | PA5137E1-LT | P097A.1 | 01 | AlarmCode:1191 AlarmMessage:NEEDLE POSITION NOT FOUND !! | CPAUTO | Over Kill | 配件殘金 | 2022/08/07 07:57:01 | 2022/08/07 09:12:27 | 1.26 | OK | 17712 | S77殘金，更換在庫卡S81上機易跳Bin27，更換在庫卡S85Pass，S81送修 | NO | NO | NO | YES | YES | NO
EQPRepairJ22105400 | T6373-16 | CMC7J0212-01 | OverKill | JB1TAS21 | PA5208K-LT | QT79P | 13 | 2R O/S-A6571 | CPAUTO | Over Kill | 2R | 2022/08/07 08:12:28 | 2022/08/07 09:49:40 | 1.62 | OK | 15897 | 輪工處理---- 1.2R更換配件_JB1TAS16_Bin8 Function.#22_Bin13 Function.皆送修. 2.原卡磨針上機Pass.溫針重對Range:16um.#13(17.5). | NO | NO | NO | YES | YES | NO
EQPRepairJ22105406 | T6373-38 | CMC8J0010-01 | OverKill | JB1GKS08 | PA6543B | J6630 | 06 | EDA SHOW 2H (CORR PASS)-17436 | CPAUTO | Over Kill | 2H | 2022/08/07 08:38:37 | 2022/08/07 08:46:10 | 0.13 | OK | 07910 | 2H，CORR PASS ，續測 | NO | NO | NO | YES | YES | NO
EQPRepairJ22105408 | T6373-128 | CWC7J0K25-01.002 | OverKill | JB1BYS11 | PA6587B_15U90H | J696F | 20 | o/s-A6571 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/07 08:41:25 | 2022/08/07 08:58:07 | 0.28 | OK | 20231 | 1.清針起測一段穩定PASS，pm(135,0) | NO | NO | NO | YES | YES | NO
EQPRepairJ22105411 | T6373-34 | CWC8J0168-01 | Tester Down | JB1BYS10 | PA6587B_15U90H | J696A | 01 | failed to load main progam-17436 | CPAUTO | Tester | 無法load 程式 | 2022/08/07 08:46:24 | 2022/08/07 09:47:13 | 1.01 | OK | 07910 | 1.load 程式秀Because depended file Global.tdl failed to loader 2.查看concel io98 Fail，T/H 重MOUNT，重LOAD 程式正常 3.#2 C-LOT上貨重新對針 PASS，退片請產線重上 | NO | NO | NO | YES | YES | NO
EQPRepairJ22105422 | T6373-32 | CWC6J0996-01.001 | Corr Fail | JB1GVS19 | NC01682A(COG48AA)-A-AT-N3-C | GGG836.1-B1 | 02 | Corr Fail | 11595 | Corr | Other | 2022/08/07 09:14:03 | 2022/08/07 09:22:32 | 0.14 | OK | 08774 | Corr Fail---- Corr 良率比 Corr Wafer 好，前流程 Fail Bin 回收成 Pass，手動比對 Corr Pass | NO | NO | NO | YES | YES | NO
EQPRepairJ22105428 | T6373-152 | CWC7J0H67-01 | OverKill | JB2P2S09 | R63455-A1S-EZ-CP | NAHM72.00-B1C | 18 | TOO MANY O/S-A6897 | CPAUTO | Over Kill | 針位調整 | 2022/08/07 09:35:27 | 2022/08/07 10:37:00 | 1.03 | OK | 15897 | 針況、壓ok.貼下、上調Pass.Range:12um.#18(27.7) | NO | NO | NO | YES | YES | NO
EQPRepairJ22105430 | T6373-38 | CMC8J0010-01 | OverKill | JB1GKS08 | PA6543B | J6630 | 06 | 配件有異物 | 17436 | Over Kill | O/S誤宰 | 2022/08/07 09:39:31 | 2022/08/07 09:47:21 | 0.13 | OK | 08774 | 毛刷清針 48.7 | NO | NO | NO | YES | YES | NO
EQPRepairJ22105472 | T6373-141 | CWC8J0134-01.001 | OverKill | JB1HKS10 | NA01687G(COG48BA)-A-AT-N2-C | GGS940.1-BC21 | 12 | o/s-14212 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/07 12:35:35 | 2022/08/07 12:54:03 | 0.31 | OK | 07910 | 機台清針 2分法pass p/m(164.1)#12 RY18 | NO | NO | NO | YES | YES | NO
EQPRepairJ22105512 | T6373-13 | CWC8J0251-01 | OverKill | JB6H1S05 | NC01690I(COG46BA)-A-AT-C | GGS960.1-BC | 17 | EDA SHOW I ; NO CORR-A6962 | CPAUTO | Over Kill | SHOW I | 2022/08/07 14:48:13 | 2022/08/07 14:58:33 | 0.17 | OK | 20231 | #15 [HW_BIN83>=12EA(1.5%)]22(HW_BIN83>=12EA(1.5%)) 1.#15 SHOW I，不同貨批(GGP052.1)不同機卡有相同問題，#16 SHOW N，續測再觀察 | NO | NO | NO | YES | YES | NO
EQPRepairJ22105515 | T6373-141 | CWC8J0134-01.001 | OverKill | JB1HKS10 | NA01687G(COG48BA)-A-AT-N2-C | GGS940.1-BC21 | 13 | #7.12 2R #13 o/s-14212 | CPAUTO | Over Kill | 2R | 2022/08/07 14:52:30 | 2022/08/07 16:01:43 | 1.15 | OK | 08774 | 2R 下卡磨針 | NO | NO | NO | YES | YES | NO
EQPRepairJ22105521 | T6373-128 | CWC7J0K25-01.002 | OverKill | JB1BYS11 | PA6587B_15U90H | J696F | 23 | o/s-A6571 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/07 15:34:45 | 2022/08/07 16:12:47 | 0.63 | OK | 16631 | 1.調整針位，機台清針 | NO | NO | NO | YES | YES | NO
EQPRepairJ22105533 | T6373-27 |  | Prober Down | JB8Y9S02 | ILISLQPEDQD-2CP | ILA13202.1 |  | AlarmCode:1060 AlarmMessage:CANNOT MEASURE INDEX SIZE !! | CPAUTO | Prober | CANOT MEASURE INDEX SIZE | 2022/08/07 16:46:20 | 2022/08/07 16:55:30 | 0.15 | OK | 07910 | 按START 機台自動抓取 OK | NO | NO | NO | YES | YES | NO
EQPRepairJ22105542 | T6373-47 | CWC8J0124-01 | OverKill | JB1HKS07 | NA01687G(COG48BA)-A-AT-N2-C | GGS268.1-B2C1 | 19 | O/S TOO MANY.-A6893 | CPAUTO | Over Kill | 機台清針後Pass | 2022/08/07 17:17:58 | 2022/08/07 17:36:06 | 0.30 | OK | 17712 | Wafer Clean，#19(170.3)RG8 | NO | NO | NO | YES | YES | NO
EQPRepairJ22105564 | T6373-128 | CWC7J0K25-01.002 | OverKill | JB1BYS11 | PA6587B_15U90H | J696F | 24 | o/s-A6571 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/07 18:41:43 | 2022/08/07 20:02:28 | 1.35 | OK | 17674 | 1. 下卡磨針Pass，溫針 range y=10 (41,7)#24 | NO | NO | NO | YES | YES | NO
EQPRepairJ22105567 | T6373-152 | CWC7J0H67-01 | OverKill | JB2P2S09 | R63455-A1S-EZ-CP | NAHM72.00-B1C | 21 | continuous eda 2h | A6897 | Over Kill | 2h | 2022/08/07 18:50:27 | 2022/08/07 18:53:49 | 0.06 | OK | 17680 | 其他批使用不同機卡仍無法回收bin26,keep test lot:NAHM98.00 #04 | NO | NO | NO | YES | YES | NO
EQPRepairJ22105569 | T6373-103 | CWC8J0135-01.001 | OverKill | JB1GVS17 | NC01682A(COG48AA)-A-AT-N3-C | GGP621.1-BC21 | 05 | SHOW 2H, O/S FAIL ALARM | A6954 | Over Kill | 針位調整 | 2022/08/07 18:56:53 | 2022/08/07 19:26:07 | 0.49 | OK | 03426 | 針位調整PASS | NO | NO | NO | YES | YES | NO
EQPRepairJ22105573 | T6373-47 | CWC8J0134-01.003 | OverKill | JB1HKS07 | NA01687G(COG48BA)-A-AT-N2-C | GGS940.1-BC12 | 01 | o/s-A6571 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/07 19:05:17 | 2022/08/07 20:03:27 | 0.97 | OK | 17674 | 1. 更換備卡#09 Pass，#01 (92,2)溫針 range y=10 (164,1) | NO | NO | NO | YES | YES | NO
EQPRepairJ22105579 | T6373-15 | CWC8J0352-01 | Setup Fail | JB8Q6S04 | NB01761A(COF94AA)-A-C | H2NC84.1-B | 01 | 19:17 setup fail due to Too large needle position offset! | A6053 | Setup Fail | 對針異常 | 2022/08/07 19:41:35 | 2022/08/07 20:39:57 | 0.97 | OK | 15096 | Setup Fail-----對針異常7-12、8-8， by pass進入溫針後，重新自動對針ok. | NO | NO | NO | YES | YES | NO
EQPRepairJ22105581 | T6373-103 | CWC8J0135-01.001 | OverKill | JB1GVS17 | NC01682A(COG48AA)-A-AT-N3-C | GGP621.1-BC21 | 05 | TOO MANY OS-A6965 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/07 19:44:45 | 2022/08/07 20:08:44 | 0.40 | OK | 17542 | 使用clean wafer清針，調整針位，試打(43,0)(142,6)無異常，點測PASS穩定 | NO | NO | YES | YES | YES | NO
EQPRepairJ22105592 | T6373-86 | CWC7J0814-01 | OverKill | JB1Q4S16 | PA6542C | J683K | 22 | too many o/s 126,5 pm checked up/down od15-A6881 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/07 20:28:44 | 2022/08/07 20:50:59 | 0.37 | OK | 11348 | Remote 處理----機台清針，重新對針Range:10um | NO | NO | NO | YES | YES | NO
EQPRepairJ22105604 | T6373-25 | CWC8J0267-01 | OverKill | JB1LWS14 | ILISLQPEDQD-2CP | ILA13196.1 | 18 | ALREADY CORR PASSED BUT STILL EDA SHOW H-A6724 | CPAUTO | Over Kill | Show H | 2022/08/07 21:08:38 | 2022/08/07 21:14:53 | 0.10 | OK | 03426 | Show H 請產線續測 | NO | NO | NO | YES | YES | NO
EQPRepairJ22105616 | T6373-28 |  | Prober Down | JB8Y9S06 | ILISLQPEDQD-2CP | ILA13202.1 |  | AlarmCode:1060 AlarmMessage:CANNOT MEASURE INDEX SIZE !! | CPAUTO | Prober | Wafer Edge/Mode Not Found | 2022/08/07 21:41:58 | 2022/08/07 21:49:28 | 0.13 | OK | 03426 | 重抓Model OK | NO | NO | NO | YES | YES | NO
EQPRepairJ22105617 | T6373-103 | CWC8J0135-01.001 | OverKill | JB1GVS17 | NC01682A(COG48AA)-A-AT-N3-C | GGP621.1-BC21 | 06 | TOO MANY OS-A6965 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/07 21:42:57 | 2022/08/07 22:52:55 | 1.17 | OK | 03426 | 更換配件S16 4.2 182.2 | NO | NO | NO | YES | YES | NO
EQPRepairJ22105621 | T6373-11 | CWC8J0353-01 | OverKill | JBCK3S01 | NB01853E(COG48AA)-A-AT-X-C | GGR104.1-B111C | 09 | EDA shows " I " [HW_P/DUTDIFF>10%] No Corr. Wafer | A6525 | Over Kill | SHOW I | 2022/08/07 22:00:16 | 2022/08/07 22:06:43 | 0.11 | OK | 13597 | 無固定DUT FAIL，續測 | NO | NO | NO | YES | YES | NO
EQPRepairJ22105625 | T6373-28 | CWC8J0288-01.001 | OverKill | JB8Y9S06 | ILISLQPEDQD-2CP | ILA13202.1 | 21 | TOO MANY O/S-A6926 | CPAUTO | Over Kill | 針位調整 | 2022/08/07 22:42:31 | 2022/08/07 22:56:30 | 0.23 | OK | 03426 | 針位調整PASS | NO | NO | NO | YES | YES | NO
EQPRepairJ22105640 | T6373-28 | CWC8J0288-01.001 | OverKill | JB8Y9S06 | ILISLQPEDQD-2CP | ILA13202.1 | 21 | O/S HAVE PROBEMARK SHIFT..PLS CONFIRM..-A6724 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/07 23:52:50 | 2022/08/08 00:31:40 | 0.65 | OK | 17542 | 溫針 & 執行soaktime，重新對針，range: X=28 Y=22，落點(2,1)，試打(11,10)(23,9)無異常，起測一段PASS | NO | NO | NO | YES | YES | NO
EQPRepairJ22105656 | T6373-141 | CWC8J0134-01 | OverKill | JB1HKS10 | NA01687G(COG48BA)-A-AT-N2-C | GGS940.1-BC11 | 08 | Check PM (38,0) pm up and down.o/s too many-A6724 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/08 00:52:14 | 2022/08/08 01:40:38 | 0.81 | OK | 03426 | 更換配件S07 165.1 9.3 | NO | NO | NO | YES | YES | NO
EQPRepairJ22105670 | T6373-135 | CEC7J0106-01 | OverKill | JB1UJS13 | PA5129LS_12U80H-LT | P08WT.1 | 11 | too many o/s pc has particles-A7033 | CPAUTO | Over Kill | 配件卡毛 | 2022/08/08 02:07:05 | 2022/08/08 02:48:57 | 0.70 | OK | 17674 | 1. 配件卡毛 下卡磨針Pass、溫針 range y=8 (35,5)#11 | NO | NO | NO | YES | YES | NO
EQPRepairJ22105678 | T6373-152 | CWC7J0H67-01 | OverKill | JB2P2S09 | R63455-A1S-EZ-CP | NAHM72.00-B1C | 23 | #21.22 CONTINOUS 2R O/S-A6336 | CPAUTO | Over Kill | 單片I、R | 2022/08/08 02:27:34 | 2022/08/08 02:56:03 | 0.47 | OK | 17674 | 1 單片I、R ，清針調針Pass (123,5)#23 ，還機觀察一片 | NO | NO | NO | YES | YES | NO
EQPRepairJ22105681 | T6373-47 | CWC8J0133-01.002 | OverKill | JB1HKS09 | NA01687G(COG48BA)-A-AT-N2-C | GGS416.1-BC22 | 02 | o/s,od:7,p/m(37.0)-A6336 | CPAUTO | Over Kill | 針位調整 | 2022/08/08 02:31:58 | 2022/08/08 02:48:33 | 0.28 | OK | 03426 | 針位調整點側PASS PASS | NO | NO | NO | YES | YES | NO
EQPRepairJ22105689 | T6373-86 | CWC7J0814-01 | Prober Down | JB1Q4S16 | PA6542C | J683K | 25 | failed lock/unlock of manual ring lever-A6881 | CPAUTO | Prober | Other | 2022/08/08 03:00:57 | 2022/08/08 03:24:59 | 0.40 | OK | 11348 | 抬頭調整ok | NO | NO | NO | YES | YES | NO
EQPRepairJ22105693 | T6373-35 | CWC8J0135-01.004 | OverKill | JB1GVS24 | NC01682A(COG48AA)-A-AT-N3-C | GGP621.1-BC112 | 03 | too many o/s-A6926 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/08 03:18:46 | 2022/08/08 05:04:23 | 1.76 | OK | 17542 | 低溫型號，更換配件#22，上機PASS，range:8，試打(17,5)無異常 | NO | NO | YES | YES | YES | NO
EQPRepairJ22105695 | T6373-47 | CWC8J0133-01.002 | OverKill | JB1HKS09 | NA01687G(COG48BA)-A-AT-N2-C | GGS416.1-BC22 | 03 | #1,2 eda show 2R (o/s)-A6881 | CPAUTO | Over Kill | 2R | 2022/08/08 03:24:46 | 2022/08/08 04:12:50 | 0.80 | OK | 17674 | 1. 2R 更換備卡#06 PASS，#03 (92,2)溫針 RANGE Y=10 (170,3) | NO | NO | NO | YES | YES | NO
EQPRepairJ22105698 | T6373-38 |  | Prober Down | JB1GKS10 | PA6543B | J6631 |  | AlarmCode:1060 AlarmMessage:CANNOT MEASURE INDEX SIZE !! | CPAUTO | Prober | Wafer Edge/Mode Not Found | 2022/08/08 03:34:49 | 2022/08/08 03:49:34 | 0.25 | OK | 03426 | 重抓Model | NO | NO | NO | YES | YES | NO
EQPRepairJ22105702 | T6373-42 | CWC8J0358-01 | OverKill | JB6H1S07 | NC01690F(COG46BA)-A-AT-C | GGY684.1-B1C | 17 | #13.14 continous 2H o/s-A6336 | CPAUTO | Over Kill | 2H | 2022/08/08 04:00:04 | 2022/08/08 04:03:50 | 0.06 | OK | 17674 | 1. 2H [HW_BIN28>=16EA(2%)]17(換機卡回收無效請續下流程) ，下一片 N 機台清針還機觀察 | NO | NO | NO | YES | YES | NO
EQPRepairJ22105707 | T6373-26 | CWC8J0267-01.001 | OverKill | JB1LWS17 | ILISLQPEDQD-2CP | ILA13196.1 | 07 | O/S TOO MANY-A6724 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/08 04:16:59 | 2022/08/08 04:38:39 | 0.36 | OK | 17542 | 使用clean wafer清針，起測一段pass穩定 | NO | NO | NO | YES | YES | NO
EQPRepairJ22105720 | T6373-121 | CWC8J0124-01.001 | OverKill | JB1QNS01 | NA01687G(COG48BA)-A-AT-N2-C | GGS268.1-B2C21 | 16 | Continuous 2 wafers EDA shows "H" No corr. wafer | A6525 | Over Kill | Show H | 2022/08/08 05:19:30 | 2022/08/08 05:28:45 | 0.15 | OK | 17542 | Remote處理 ----- Show H，[HW_Bin5>=7.00ea]10，不同機卡有相同問題(GGS269.1-T)，故續測 | NO | NO | NO | YES | YES | NO
EQPRepairJ22105721 | T6373-13 | CWC8J0361-01 | OverKill | JB6H1S05 | NC01690I(COG46BA)-A-AT-C | GGS956.1-BC | 15 | EDA shows " I " No Corr. wafer | A6525 | Over Kill | show I | 2022/08/08 05:25:53 | 2022/08/08 05:29:12 | 0.06 | OK | 15096 | #14單片show I，續測觀察. | NO | NO | NO | YES | YES | NO
EQPRepairJ22105730 | T6373-37 | CWC8J0136-01.001 | OverKill | JB1GVS26 | NC01682A(COG48AA)-A-AT-N3-C | GGP622.1-BC21 | 23 | 2 continues eda show H-A6926 | CPAUTO | Over Kill | Show H | 2022/08/08 05:56:45 | 2022/08/08 06:44:25 | 0.79 | OK | 17542 | Show H，Function Fail Bin29過多，更換配件#23，上機PASS，溫針重對針，range:12，試打(17,5)無異常，點測PASS穩定 | NO | NO | YES | YES | YES | NO
EQPRepairJ22105749 | T6373-38 | CWC6J0J45-01.001 | OverKill | JB1GKS10 | PA6543B | J6631 | 04 | too many o/s 2H-RETEST LOT-A6926 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/08 06:40:13 | 2022/08/08 07:57:44 | 1.29 | OK | 07910 | 機台清針 ，重新對針 PASS P/M(173.2)#4 RY12 | NO | NO | NO | YES | YES | NO
EQPRepairJ22105755 | T6373-39 | CWC8J0168-01.001 | OverKill | JB1BYS12 | PA6587B_15U90H | J6969 | 18 | #16.17 continous 2I..open and short mode test-A6336 | CPAUTO | Over Kill | 2I | 2022/08/08 07:10:39 | 2022/08/08 08:16:39 | 1.10 | OK | 17712 | 2I Open & Short Mode，確認針尖黑點下卡磨針#18(18.6)RG16 | NO | NO | NO | YES | YES | NO
EQPRepairJ22105764 | T6373-29 | CEC7J0111-01.002 | Corr Fail | JB1EBS16 | NA01644G(COG88AA)-A-AT-V-C | GGV062.1-2BC2211 | 19 | t44 corr compare fail-A6526 | CPAUTO | Corr | Other | 2022/08/08 08:17:49 | 2022/08/08 08:25:48 | 0.13 | OK | 17680 | remote----手動比對pass | NO | NO | NO | YES | YES | NO
EQPRepairJ22105767 | T6373-90 | CRC7J0020-01 | Corr Fail | JB1GVS17 | NC01682A(COG48AA)-A-AT-N3-C | GCPG57.1-BC112C | 18 | corr fail-A6526 | CPAUTO | Corr | Other | 2022/08/08 08:22:50 | 2022/08/08 08:27:36 | 0.08 | OK | 17680 | remote----手動比對pass | NO | NO | NO | YES | YES | NO
EQPRepairJ22105773 | T6373-16 | CWC7J0H53-01 | OverKill | JB1TAS21 | PA5208KF-LT | J67HK | 11 | too many o/s-A6668 | CPAUTO | Over Kill | 配件殘金 | 2022/08/08 08:40:25 | 2022/08/08 09:38:17 | 0.96 | OK | 17712 | 下卡磨針，#11(146.5)RG16 | NO | NO | NO | YES | YES | NO
EQPRepairJ22105774 | T6373-90 | CWC7J0797-01.001 | Prober Down | JB1GVS17 | NC01682A(COG48AA)-A-AT-N3-C | GCPG57.1-BC112C | 14 | model not found-A6526 | CPAUTO | Prober | Wafer Edge/Mode Not Found | 2022/08/08 08:44:14 | 2022/08/08 08:46:59 | 0.05 | OK | 08774 | 重抓 | NO | NO | NO | YES | YES | NO
EQPRepairJ22105775 | T6373-70 | CWC8J0216-01 | Prober Down | JB1Q4S15 | PA6542C | J683M | 21 | AlarmCode:1995 AlarmMessage:Faied to upload log file | CPAUTO | Prober | Other | 2022/08/08 08:47:49 | 2022/08/08 08:54:46 | 0.12 | OK | 15897 | Check setting ok.換片起測Open.OD:10um ok.#21(13.4). | NO | NO | NO | YES | YES | NO
EQPRepairJ22105779 | T6373-90 | CWC7J0797-01.001 | Prober Down | JB1GVS17 | NC01682A(COG48AA)-A-AT-N3-C | GGR319.1-BC2111 | 14 | AlarmCode:8732 AlarmMessage:It failed in lock/unlock of the manual ring lever. | CPAUTO | Prober | lock/unlock | 2022/08/08 08:55:45 | 2022/08/08 09:06:26 | 0.18 | OK | 07910 | 微調位置ok | NO | NO | NO | YES | YES | NO
EQPRepairJ22105781 | T6373-26 | CWC8J0267-01.001 | OverKill | JB1LWS17 | ILISLQPEDQD-2CP | ILA13196.1 | 09 | O/S-17436 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/08 09:04:22 | 2022/08/08 09:25:09 | 0.35 | OK | 07910 | 機台清針 2分法PASS P/M(242.4)#9 RY22 | NO | NO | NO | YES | YES | NO
EQPRepairJ22105789 | T6373-141 | CWC8J0125-01.001 | OverKill | JB1HKS07 | NA01687G(COG48BA)-A-AT-N2-C | GGP242.1-B1C21 | 18 | o/s-17436 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/08 09:22:02 | 2022/08/08 09:37:17 | 0.25 | OK | 07910 | 機台清針 2分法PASS P/M(170.3)#14 RY17 | NO | NO | NO | YES | YES | NO
EQPRepairJ22105791 | T6373-103 | CWC8J0135-01.001 | OverKill | JB1GVS16 | NC01682A(COG48AA)-A-AT-N3-C | GGP621.1-BC21 | 10 | 2H O/S (181.4)PM CHECK-11933 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/08 09:39:44 | 2022/08/08 09:44:57 | 0.09 | OK | 08774 | Remote 處理-------清針 確認針壓針位 續測觀察一片 | NO | NO | NO | YES | YES | NO
EQPRepairJ22105794 | T6373-35 | CWC8J0136-01.003 | OverKill | JB1GVS22 | NC01682A(COG48AA)-A-AT-N3-C | GGP622.1-BC222 | 10 | too many o/s-A6526 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/08 09:54:24 | 2022/08/08 11:08:49 | 1.24 | OK | 17680 | REMOTE----調整針位POSITION (-10,5),MARK(140,0)，點測穩定 | NO | NO | NO | YES | YES | NO
EQPRepairJ22105796 | T6373-128 | CWC7J0K25-01.001 | OverKill | JB1BYS11 | PA6587B_15U90H | J696F | 06 | 06 too many O/S-11595 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/08 09:56:40 | 2022/08/08 10:20:43 | 0.40 | OK | 15897 | 清針重對Range:16um.#06(19.6). | NO | NO | NO | YES | YES | NO
EQPRepairJ22105804 | T6373-86 | CWC7J0C21-01.009 | OverKill | JB1Q4S16 | PA6542C | J68S1 | 05 | o/s-12254 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/08 10:26:49 | 2022/08/08 16:22:09 | 5.92 | OK | 16631 | 1.<低溫型號>輪工宥承處理O/S不穩，更換備卡2張各有固定一串OPEN & 易跳SHORT，後續輪工品睿接手處理重MNT & 重新對針PASS，起測一段OK，先交回續測觀察 | NO | NO | NO | YES | YES | NO
EQPRepairJ22105810 | T6373-111 | CWC7J0672-01 | OverKill | TB1R8S74 | PA5137F-LT1 | P07NT.1 | 07 | o/s-A6921 | CPAUTO | Over Kill | O/D不足 | 2022/08/08 10:55:26 | 2022/08/08 11:27:30 | 0.53 | OK | 15897 | 清針、重對Range:10um.#07(17.5) | NO | NO | NO | YES | YES | NO
EQPRepairJ22105813 | T6373-141 | CWC8J0125-01.001 | OverKill | JB1HKS07 | NA01687G(COG48BA)-A-AT-N2-C | GGP242.1-B1C21 | 18 | o/s-17436 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/08 11:09:59 | 2022/08/08 12:22:56 | 1.22 | OK | 07910 | 下卡磨針 取卡上機PASS P/M(173.3)#18 RY6 | NO | NO | NO | YES | YES | NO
EQPRepairJ22105817 | T6373-21 | CWC7J0K21-01 | OverKill | JB1TAS26 | PA5208KF-LT | J696N | 11 | EDA SHOW 2R.-A6526 | CPAUTO | Over Kill | 2R | 2022/08/08 11:26:30 | 2022/08/08 12:20:19 | 0.90 | OK | 13749 | 1.EDA SHOW 2R O/S 針尖黑點. 2.更換備卡#S26->#S19 OK. 3.上機預溫 重對針點測PASS 溫針退針壓 OK PM(46.4) X=0 Y=-1 slot #11 | NO | NO | NO | YES | YES | NO
EQPRepairJ22105818 | T6373-152 |  | Prober Down | JB2P2S09 | R63455-A1S-EZ-CP | NAPY65.00-B1C |  | ERROR OCCURS IN NETWORK PROCESS-A6986 | CPAUTO | Product Issue | 無法 Download Map | 2022/08/08 11:30:51 | 2022/08/08 11:46:22 | 0.26 | OK | 17712 | NAHN27.00-2 無法下載MAP，查看edatrans無資料，請DN 產品協助確認 | NO | NO | NO | YES | YES | NO
EQPRepairJ22105820 | T6373-26 | CWC8J0267-01.001 | OverKill | JB1LWS17 | ILISLQPEDQD-2CP | ILA13196.1 | 10 | EDA SHOW 2R(O/S)-17436 | CPAUTO | Over Kill | 2R | 2022/08/08 11:35:12 | 2022/08/08 11:45:52 | 0.18 | OK | 07910 | 2R，檢查針況正常，機台清針 pass ，觀察下片 | NO | NO | NO | YES | YES | NO
EQPRepairJ22105825 | T6373-152 |  | Product Issue | JB2P2S09 | R63455-A1S-EZ-CP | NAPY65.00-B1C |  | 無法下載MAP，查看edatrans無資料，請DN 產品協助確認 | 17712 | Product Issue | 無法 Download Map | 2022/08/08 11:47:00 | 2022/08/08 12:14:54 | 0.47 | OK | 17712 | 產品協助確認後ok。 | NO | NO | NO | YES | YES | NO
EQPRepairJ22105834 | T6373-37 | CWC8J0119-01.001 | Corr Fail | JB1GVS23 | NC01682A(COG48AA)-A-AT-N3-C | GGP622.1-BC21 | 09 | Corr Fail | 17436 | Corr | Corr Fail | 2022/08/08 12:40:23 | 2022/08/08 12:52:45 | 0.21 | OK | 16631 | 1.Corr Fail，pass | NO | NO | NO | YES | YES | NO
EQPRepairJ22105835 | T6373-152 | CWC8J0165-01 | Other | JB2P2S09 | R63455-A1S-EZ-CP | NAHN27.00-B2C | 16 | please save datalog | A6668 | Other | Other | 2022/08/08 12:41:10 | 2022/08/08 12:51:27 | 0.17 | OK | 17712 | 設定Datalog路徑 | NO | NO | NO | YES | YES | NO
EQPRepairJ22105844 | T6373-136 | CWC8J0251-01 | Prober Down | JB6H1S06 | NC01690I(COG46BA)-A-AT-C | GGS960.1-BC1 | 25 | AlarmCode:8732 AlarmMessage:It failed in lock/unlock of the manual ring lever. | CPAUTO | Prober | Manual ring level | 2022/08/08 13:29:43 | 2022/08/08 13:33:55 | 0.07 | OK | 17712 | 重新LOCK/UNLOCK多次OK | NO | NO | NO | YES | YES | NO
EQPRepairJ22105850 | T6373-136 | CWC8J0394-01 | Prober Down | JB6H1S06 | NC01690G(COG48AA)-A-AT-N4-C | GGM297.1-BC | 01 | AlarmCode:8732 AlarmMessage:It failed in lock/unlock of the manual ring lever. | CPAUTO | Prober | Manual ring level | 2022/08/08 14:16:21 | 2022/08/08 15:24:34 | 1.14 | OK | 17712 | 更換Lock桿卡榫，配件殘金磨針，#1(127.7) | NO | NO | NO | YES | YES | NO
EQPRepairJ22105856 | T6373-27 | CWC8J0288-01 | OverKill | JB1LWS15 | ILISLQPEDQD-2CP | ILA13202.1 | 01 | o/s-17436 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/08 14:46:26 | 2022/08/08 15:14:28 | 0.47 | OK | 08774 | Remote 處理----重新對針 210.2 | NO | NO | NO | YES | YES | NO
EQPRepairJ22105862 | T6373-121 | CWC8J0124-01.002 | Other | JB1QNS01 | NA01687G(COG48BA)-A-AT-N2-C | GGS268.1-B2C211 | 21 | STDF-MES WITHOUT AUTO RUN-12254 | CPAUTO | Other | Other | 2022/08/08 14:58:30 | 2022/08/08 16:03:20 | 1.08 | OK | 16631 | 1.請產線手動掛run，重新上貨 | NO | NO | NO | YES | YES | NO
EQPRepairJ22105867 | T6373-103 | CWC8J0085-01.004 | Setup Fail | JB1CES09 | NA01248F(COG4)-A-AT-X2-C | NT2YA-B | 01 | 低溫機，固定PIN FAIL(Short bin 3)，調針清針皆無效，無備卡，下卡磨針待輪工上機確認。 | 08438 | Setup Fail | O/S Fail | 2022/08/08 15:19:34 | 2022/08/08 18:30:26 | 3.18 | OK | 17680 | range:11um,position(-5,0),mark(68,0) | NO | NO | NO | YES | YES | NO
EQPRepairJ22105872 | T6373-34 | CWC7J0M27-01 | OverKill | JB1BYS10 | PA6587B_15U90H | J696G | 04 | EDA SHOW 2H NO RUN CORR.-A6526 | CPAUTO | Over Kill | 2H | 2022/08/08 15:37:15 | 2022/08/08 15:48:16 | 0.18 | OK | 20231 | #02 [HW_BIN29>0.4%](Bin 30 > 4EA) 5(Hold Wafer) #03 [HW_BIN60<97%](Bin 61 < 1107EA) 1095----- 1.兩片不同原因，續測再觀察。 | NO | NO | NO | YES | YES | NO
EQPRepairJ22105873 | T6373-27 | CWC8J0288-01 | OverKill | JB1LWS15 | ILISLQPEDQD-2CP | ILA13202.1 | 01 | TOO MANY FAIL BIN 23-A6526 | CPAUTO | Probe Card | bin24 | 2022/08/08 15:39:38 | 2022/08/08 18:40:39 | 3.02 | OK | 07910 | 1.原配件(JB1LWS15)BIN23，磨針後跑BIN24 2.與量產機ND3-28(JB1LWS16)交互驗證上ND3-27 PASS 3.P/C(#16)再上原機PASS ，故P/C(#15)送修無備卡，暫掛配件當 | NO | NO | NO | YES | YES | NO
EQPRepairJ22105875 | T6373-47 | CWC8J0133-01.002 | OverKill | JB1HKS06 | NA01687G(COG48BA)-A-AT-N2-C | GGS416.1-BC22 | 09 | too many O/S, previous eda already 2R,O/S-A6668 | CPAUTO | Over Kill | 2R | 2022/08/08 15:50:27 | 2022/08/08 17:41:22 | 1.85 | OK | 17712 | 下卡磨針，#9(35.4)RG18 | NO | NO | NO | YES | YES | NO
EQPRepairJ22105882 | T6373-90 | CWC5J0942-01.009 | Prober Down | JB1GVS17 | NC01682A(COG48AA)-A-AT-N2-C | GFYR90.1-BC211121 | 14 | AlarmCode:8732 AlarmMessage:It failed in lock/unlock of the manual ring lever. | CPAUTO | Prober | MANUAL RING ERROR | 2022/08/08 16:12:22 | 2022/08/08 16:26:23 | 0.23 | OK | 17680 | 微調LOCK桿,手動測試多次皆正常 | NO | NO | NO | YES | YES | NO
EQPRepairJ22105886 | T6373-86 | CWC7J0C21-01.009 | OverKill | JB1Q4S16 | PA6542C | J68S1 | 05 | still o/s probemark too big near too probemarkshift 165.2-A6894 | CPAUTO | Prober | Other | 2022/08/08 16:30:54 | 2022/08/08 18:19:09 | 1.80 | OK | 20231 | 1.測試中open不穩掛當，更換P/C #14-->#22上機，停留在cardchange end畫面20分鐘後才cardchange end開始降溫 2.調針位完rg:14，針痕確認OK，點測上下左右中時發現有輕微結霜，起測不穩 3.升溫100度烘烤30分鐘，轉掛prober down。 | NO | NO | NO | YES | YES | NO
EQPRepairJ22105891 | T6373-28 | CWC8J0288-01.001 | Other | JB1LWS16 | ILISLQPEDQD-2CP | ILA13202.1 | 14 | ae change pc-11933 | CPAUTO | Probe Card | Other | 2022/08/08 17:36:11 | 2022/08/08 18:11:02 | 0.58 | OK | 07910 | 1.配件(JB1LWS16)借ND3-27 上機驗證PASS 2.再將卡於ND3-28 重下針壓PASS P/M(221.11)#14 Range:x(-14~12)y(1~-16) | NO | NO | NO | YES | YES | NO
EQPRepairJ22105894 | T6373-21 | CWC7J0A61-01 | Prober Down | JB1GKS08 | PA6543B | J67H7 | 22 | MODEL NOT FOUND-17436 | CPAUTO | Prober | Wafer Edge/Mode Not Found | 2022/08/08 17:52:31 | 2022/08/08 18:11:55 | 0.32 | OK | 17680 | retry & check die ok,mark(47,7) | NO | NO | NO | YES | YES | NO
EQPRepairJ22105900 | T6373-86 |  | Prober Down | JB1Q4S16 | PA6542C | J68S1 |  | 換卡時有等20分鐘，發現仍結霜，100度烘烤30分鐘，轉掛prober down | 20231 | Prober | Other | 2022/08/08 18:20:55 | 2022/08/08 19:13:33 | 0.88 | OK | 20231 | 1.烘烤30分鐘後降溫重新對針，上機測試PASS，二分法座標(76,2)，RG:14，pm(12,1)，起測穩定 | NO | NO | NO | YES | YES | NO
EQPRepairJ22105901 | T6373-21 |  | Prober Down | JB1GKS08 | PA6543B | J67H5 |  | AlarmCode:1060 AlarmMessage:CANNOT MEASURE INDEX SIZE !! | CPAUTO | Prober | Wafer Edge/Mode Not Found | 2022/08/08 18:21:19 | 2022/08/08 18:30:08 | 0.15 | OK | 15897 | 重抓Module ok#23(20.6)(58.3). | NO | NO | NO | YES | YES | NO
EQPRepairJ22105906 | T6373-70 | CWC7J0M28-01 | OverKill | JB1Q4S16 | PA6542C | J683J | 02 | ALWAYS ALARM, TOO MANY O/S, PM CHECK UP/DOWN (128,5)-A6986 | CPAUTO | Over Kill | O/D不足 | 2022/08/08 18:35:55 | 2022/08/08 19:03:25 | 0.46 | OK | 15897 | 跳針壓、重對Range:14um.#02(13.4). | NO | NO | NO | YES | YES | NO
EQPRepairJ22105907 | T6373-27 |  | Probe Card Down | JB1LWS15 | ILISLQPEDQD-2CP | ILA13202.1 |  | 1.原配件(JB1LWS15)BIN23，磨針後跑BIN24 2.與量產機ND3-28(JB1LWS16)交互驗證上ND3-27 PASS 3.P/C(#16)再上原機PASS ，故P/C(#15)送修無備卡，暫掛配件當 | 07910 | Probe Card | bin24 | 2022/08/08 18:41:12 | 2022/08/08 20:30:44 | 1.83 | OK | 13597 | 1.原配件(JB1LWS15)BIN23，磨針後跑BIN24 2.與量產機ND3-28(JB1LWS16)交互驗證上ND3-27 PASS 3.P/C(#16)再上原機PASS ，故P/C(#15)送修無備卡，暫掛配件當 4.站長通知idle | NO | NO | NO | YES | YES | NO
EQPRepairJ22105932 | T6373-121 | CWC8J0125-01.004 | Other | JB1QNS01 | NA01687G(COG48BA)-A-AT-N2-C | GGS940.1-BC122 | 12 | STDF-NO AUTO MOVE IN-A6881 | CPAUTO | Other | 無自動進站 | 2022/08/08 20:53:19 | 2022/08/08 20:58:19 | 0.08 | OK | 11348 | T44 STDF 上貨起無自動進站，退片請產線重上 | NO | NO | NO | YES | YES | NO
EQPRepairJ22105943 | T6373-07 | CWC6J0R33-01.001 | OverKill | JB1TAS20 | PA5208K-LT | QT6L7 | 02 | TOO MANY O/S PM CHECK 19.5 OD10-A7024 | CPAUTO | Over Kill | O/S | 2022/08/08 21:31:03 | 2022/08/08 21:54:57 | 0.40 | OK | 15096 | 邊緣o/s，針壓不足，重新對針pass. | NO | NO | NO | YES | YES | NO
EQPRepairJ22105948 | T6373-111 | CWC7J0672-01 | OverKill | TB1R8S74 | PA5137F-LT1 | P07NT.1 | 17 | TOO MANY O/S PM CHECK 143.5 OD11-A7024 | CPAUTO | Over Kill | O/S | 2022/08/08 21:51:00 | 2022/08/08 22:06:49 | 0.26 | OK | 15096 | 清針 二分法調針pass，range:8 | NO | NO | NO | YES | YES | NO
EQPRepairJ22105951 | T6373-86 | CWC7J0C21-01.009 | OverKill | JB1Q4S22 | PA6542C | J68S1 | 06 | too many o/s 158,4 pm checked od11-A6881 | CPAUTO | Over Kill | O/S | 2022/08/08 22:19:26 | 2022/08/08 22:43:17 | 0.40 | OK | 15096 | 重新對針PASS. | NO | NO | NO | YES | YES | NO
EQPRepairJ22105953 | T6373-111 | CWC7J0672-01 | OverKill | TB1R8S74 | PA5137F-LT1 | P07NT.1 | 17 | too many o/s 43,6 pm checked up/down od15-A6881 | CPAUTO | Over Kill | O/S | 2022/08/08 22:33:29 | 2022/08/09 00:56:34 | 2.38 | OK | 15096 | <低溫型號>-----下卡磨針，上機等待20分鐘，降溫setup pass. | NO | NO | NO | YES | YES | NO
EQPRepairJ22105959 | T6373-135 | CEC7J0106-01 | OverKill | JB1UJS13 | PA5129LS_12U80H-LT | P08WT.1 | 18 | EDA CONTINOUS SHOWS 2H DUE O/S-A6788 | CPAUTO | Over Kill | 2H | 2022/08/08 23:07:45 | 2022/08/08 23:50:54 | 0.72 | OK | 17674 | 1. 2H 更換備卡#18 PASS，溫針 RANGE Y=10 (81,0) | NO | NO | NO | YES | YES | NO
EQPRepairJ22105961 | T6373-16 | CWC7J0H53-01 | OverKill | JB1TAS21 | PA5208KF-LT | J67HK | 17 | o/s too many due p/c have particles-A6788 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/08 23:11:51 | 2022/08/09 00:45:35 | 1.56 | OK | 17674 | 1. 配件殘金 (金絲)更換備卡#26 PASS，溫針 RANGE Y=10 (115,1)#17 | NO | NO | NO | YES | YES | NO
EQPRepairJ22105964 | T6373-33 | CWC7J0A60-01 | OverKill | JB1TAS14 | PA5208KF2 | J68G8 | 11 | too many 8,5 pm up/down od11-A6881 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/08 23:16:32 | 2022/08/08 23:35:52 | 0.32 | OK | 11348 | 清針pass，Range:10um | NO | NO | NO | YES | YES | NO
EQPRepairJ22105983 | T6373-70 | CWC7J0M28-01 | OverKill | JB1Q4S16 | PA6542C | J683J | 06 | too many o/s (18.4) pm check done overkill process-A6881 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/09 00:37:39 | 2022/08/09 01:00:26 | 0.38 | OK | 17674 | 1. 清針Pass (12,4)，依 memo 單開 clean wafer | NO | NO | NO | YES | YES | NO
EQPRepairJ22105984 | T6373-135 | CEC7J0106-01 | OverKill | JB1UJS18 | PA5129LS_12U80H-LT | P08WT.1 | 18 | MBD 2H o/s-A6881 | CPAUTO | Over Kill | 2H | 2022/08/09 00:44:14 | 2022/08/09 00:47:48 | 0.06 | OK | 17674 | 1. 2H 剛已掛過已換過卡了 | NO | NO | NO | YES | YES | NO
EQPRepairJ22105991 | T6373-86 | CWC7J0C21-01.009 | OverKill | JB1Q4S22 | PA6542C | J68S1 | 07 | always O/S(OVERKILL DONE APPLY) | A6296 | Over Kill | O/S誤宰 | 2022/08/09 01:02:04 | 2022/08/09 02:54:57 | 1.88 | OK | 11348 | <低溫換卡>更換配件#15 上機等待20min 降溫setup pass | NO | NO | NO | YES | YES | NO
EQPRepairJ22105992 | T6373-26 | CWC8J0271-01 | OverKill | JB1LWS17 | ILISLQPEDQD-2CP | ILA13192.1 | 07 | BIN7 TOO MANY-19697 | CPAUTO | Over Kill | Function誤宰 | 2022/08/09 01:10:23 | 2022/08/09 02:01:06 | 0.85 | OK | 13597 | 清調針PASS，RANGE 10 UM | NO | NO | NO | YES | YES | NO
EQPRepairJ22106000 | T6373-27 | CWC8J0446-01.001 | Setup Fail | JB8Y9S02 | ILISLQPEDQD-2CP | ILA14003.1 | 13 | AUTO SETUP NG, PLEASE CHECK | A6963 | Setup Fail | tdr fail | 2022/08/09 01:42:24 | 2022/08/09 01:55:23 | 0.22 | OK | 13597 | 重mnt測試頭，重load程式ok | NO | NO | NO | YES | YES | NO
EQPRepairJ22106003 | T6373-152 | CWC8J0117-01 | OverKill | JB2P2S09 | R63455-A1S-EZ-CP | NAHM76.00-B11C | 05 | Continuous 2 wafers EDA shows "H",NO CORR WAFER | A6786 | Over Kill | Show H | 2022/08/09 02:04:35 | 2022/08/09 02:08:24 | 0.06 | OK | 17542 | Show H，[HW_BIN26>0.5%]7，不同批有相同問題(NAHM72.00)，請先行續測 | NO | NO | NO | YES | YES | NO
EQPRepairJ22106020 | T6373-26 | CWC8J0271-01 | OverKill | JB1LWS17 | ILISLQPEDQD-2CP | ILA13192.1 | 04 | too many os, bin7-A6963 | CPAUTO | Over Kill | Function誤宰 | 2022/08/09 03:07:22 | 2022/08/09 04:04:29 | 0.95 | OK | 13597 | 下卡磨針，range 10 um | NO | NO | NO | YES | YES | NO
EQPRepairJ22106022 | T6373-70 | CWC7J0M28-01 | OverKill | JB1Q4S16 | PA6542C | J683J | 07 | too many o/s-A6148 | CPAUTO | Over Kill | O/S | 2022/08/09 03:10:35 | 2022/08/09 03:56:28 | 0.76 | OK | 15096 | open 針位調整、clean wafer清針pass. | NO | NO | NO | YES | YES | NO
EQPRepairJ22106031 | T6373-47 | CWC8J0133-01.002 | OverKill | JB1HKS06 | NA01687G(COG48BA)-A-AT-N2-C | GGS416.1-BC22 | 17 | too many o/s-A6148 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/09 03:54:05 | 2022/08/09 04:29:46 | 0.59 | OK | 15895 | 機台清針，y:12 p/m(147.4)，先行還機待啟測觀察 | NO | NO | NO | YES | YES | NO
EQPRepairJ22106035 | T6373-70 |  | Prober Down | JB1Q4S16 | PA6542C | J683J |  | AlarmCode:7214 AlarmMessage:Soaktime correct Z-UP limit over !! | CPAUTO | Prober | Other | 2022/08/09 04:05:18 | 2022/08/09 07:12:55 | 3.13 | OK | 17674 | 1. Z軸易常，重對針Pass 溫針os 不穩 2. OS 針尖黑點，更換備卡#17 上機Bin55-O/S固定Pin，清調針無效 3. 更換備卡#14 上機pass，Range y=2 試打針位不佳針點大。 4. 原卡#16 磨針上機Bin55-O/S，亂跳Pin 5. 更換備卡#22 Pass，故將已上配件送修 6. #08 (86,2)溫針 range y=10 (46,0) | NO | NO | NO | YES | YES | NO
EQPRepairJ22106065 | T6373-47 | CWC8J0133-01.002 | OverKill | JB1HKS06 | NA01687G(COG48BA)-A-AT-N2-C | GGS416.1-BC22 | 18 | always open short-A6786 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/09 06:27:13 | 2022/08/09 08:16:27 | 1.82 | OK | 17229 | 1.CLEAN WAFER預溫重對針二分法RANGE:11um落點(164.1)5點pass. | NO | NO | NO | YES | YES | NO
EQPRepairJ22106067 | T6373-136 | CWC8J0118-01.001 | Setup Fail | JB6K8S02 | NC01682A(COG48AA)-A-AT-N3-C | GGR320.1-B1 | 23 | too large needle offset | A6788 | Setup Fail | Prober異常 prober abnormal | 2022/08/09 06:34:32 | 2022/08/09 08:15:30 | 1.68 | OK | 07764 | 自動對針異常 , 預溫後自動對針 OK | NO | NO | NO | YES | YES | NO
EQPRepairJ22106070 | T6373-90 | CRC7J0011-01 | OverKill | JB1GVS17 | NC01682A(COG48AA)-A-AT-N3-C | GCVP99.1-BC111121C | 01 | o/s too many-19697 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/09 06:42:50 | 2022/08/09 07:06:16 | 0.39 | OK | 17542 | 機台清針，重對針，range:16，試打(169,1)無異常，點測PASS穩定 | NO | NO | YES | YES | YES | NO
EQPRepairJ22106073 | T6373-12 | CWC8J0352-01.003 | Other | JB1PNS06 | NB01761A(COF94AA)-A-C | H2NC84.1-BC22 | 02 | 27,7 / 169.0 pm checked up/down od5 wait ae to confirm pm-A6881 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/09 06:44:35 | 2022/08/09 07:47:03 | 1.04 | OK | 10210 | 調整針位(197.7) | NO | NO | NO | YES | YES | NO
EQPRepairJ22106074 | T6373-28 | CWC8J0288-01.001 | OverKill | JB1LWS16 | ILISLQPEDQD-2CP | ILA13202.1 | 21 | O/S TOO MANY-19697 | CPAUTO | Over Kill | O/D不足 | 2022/08/09 06:45:54 | 2022/08/09 06:51:34 | 0.09 | OK | 13597 | 加針壓pass | NO | NO | NO | YES | YES | NO
EQPRepairJ22106079 | T6373-86 | CWC7J0C21-01.009 | OverKill | JB1Q4S15 | PA6542C | J68S1 | 08 | too many O/S(probemark too big already) | A6296 | Probe Card | #15 #23配件送修 | 2022/08/09 07:09:58 | 2022/08/09 10:34:54 | 3.42 | OK | 10210 | P/C DOWN--JB1Q4S15 JB1Q4S23送修 1.夜班交接#15 OPEN下卡磨針，接手上機Bin58 OPEN調針無效。 2.更換再庫卡#23 Bin51 Function。 3.故#15 #23配件送修，無備卡此機配件當。 | NO | NO | NO | YES | YES | NO
EQPRepairJ22106089 | T6373-13 | CWC8J0395-01 | OverKill | JB6H1S05 | NC01690G(COG48AA)-A-AT-N4-C | GGK811.1-B1C1 | 25 | #24 EDA I NO CORR-A6920 | CPAUTO | Over Kill | Show I | 2022/08/09 08:29:05 | 2022/08/09 08:40:53 | 0.20 | OK | 10210 | 單片Show I 請產線續測 | NO | NO | NO | YES | YES | NO
EQPRepairJ22106090 | T6373-90 | CRC7J0012-01 | OverKill | JB1GVS17 | NC01682A(COG48AA)-A-AT-N3-C | GCTG04.1-BC121C | 02 | o/s.-A6893 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/09 08:31:34 | 2022/08/09 10:37:37 | 2.10 | OK | 13581 | 1.清調針不過更換#26配件上機PASS(138.6) | NO | NO | NO | NO | YES | NO
EQPRepairJ22106092 | T6373-141 | CWC8J0248-01.002 | OverKill | JB1HKS07 | NA01687G(COG48BA)-A-AT-N2-C | GGS267.1-B11C22 | 05 | too many fail bin | A6858 | Over Kill | Function誤宰 | 2022/08/09 08:33:01 | 2022/08/09 09:46:05 | 1.22 | OK | 17229 | 1.更換配件#07>10預溫重對針二分法RANGE:10um落點172.3點測5點pass. | NO | NO | NO | YES | YES | NO
EQPRepairJ22106095 | T6373-47 | CWC8J0133-01.002 | OverKill | JB1HKS06 | NA01687G(COG48BA)-A-AT-N2-C | GGS416.1-BC22 | 18 | too many os overkill pass the os again-A6785 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/09 08:41:42 | 2022/08/09 08:57:02 | 0.26 | OK | 07764 | STDF 掛RUN 維修 | NO | NO | NO | YES | YES | NO
EQPRepairJ22106104 | T6373-147 | CWC7J0L14-01.002 | OverKill | TB1R8S85 | PA5137E1-LT | P097A.1 | 04 | #4 O/S FAIL | 09052 | Over Kill | O/S誤宰 | 2022/08/09 09:16:08 | 2022/08/09 10:29:37 | 1.22 | OK | 20230 | 1.O/S 2.機台清針,重新對針,二分法調針YRG=13,XRG=47,PM(43.6) 3.PASS | NO | NO | NO | YES | YES | NO
EQPRepairJ22106105 | T6373-47 | CWC8J0133-01.002 | OverKill | JB1HKS06 | NA01687G(COG48BA)-A-AT-N2-C | GGS416.1-BC22 | 18 | bin29 fail-07764 | CPAUTO | Over Kill | Function誤宰 | 2022/08/09 09:17:24 | 2022/08/09 09:40:55 | 0.39 | OK | 07764 | bin29 更換備卡 | NO | NO | NO | YES | YES | NO
EQPRepairJ22106111 | T6373-21 |  | Prober Down | JB1GKS08 | PA6543B | J6630 |  | AlarmCode:1057 AlarmMessage:MODEL NOT FOUND ( ROUGH ALIGNMENT ) !! | CPAUTO | Prober | Wafer Edge/Mode Not Found | 2022/08/09 09:52:49 | 2022/08/09 10:02:37 | 0.16 | OK | 17229 | 1Mode Not Found AND CHECK DIE OK檢查針痕173.5 OK | NO | NO | NO | YES | YES | NO
EQPRepairJ22106117 | T6373-136 | CWC8J0118-01.001 | OverKill | JB6K8S02 | NC01682A(COG48AA)-A-AT-N3-C | GGR320.1-B1C211 | 23 | TOO MANY O/S-A6143 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/09 10:07:47 | 2022/08/09 10:56:54 | 0.82 | OK | 20230 | 1.o/s 2.機台清針,重新對針,二分法調針rg=19,pm(46.6) 3.pass | NO | NO | NO | YES | YES | NO
EQPRepairJ22106122 | T6373-86 |  | Probe Card Down | JB1Q4S15 | PA6542C | J68S1 |  | P/C DOWN--JB1Q4S15 JB1Q4S23送修 1.夜班交接#15 OPEN下卡磨針，接手上機Bin58 OPEN調針無效。 2.更換再庫卡#23 Bin51 Function。 3.故#15 #23配件送修，無備卡此機配件當。 | 10210 | Probe Card | OPEN送修 | 2022/08/09 10:35:25 | 2022/08/09 15:31:49 | 4.94 | OK | 10210 | 1.此機#15 #23配件送修掛配件當 2.產線通知有#22上機多根Pin OPE 3.調針、重mnt、更換配件角度相同 4.拉ND4-184 #19上機Fail Pin相同 5.確認LCD#09 Fail，此機轉Tester down請常日協助 | NO | NO | NO | YES | YES | NO
EQPRepairJ22106159 | T6373-128 | CWC8J0088-01 | OverKill | JB1BYS11 | PA6587B_15U90H | J696C | 21 | EDA SHOW 2R O/S | A6785 | Over Kill | O/S誤宰 | 2022/08/09 13:14:07 | 2022/08/09 13:43:45 | 0.49 | OK | 07764 | 2r 調整針位/修改清針參數 | NO | NO | NO | YES | YES | NO
EQPRepairJ22106168 | T6373-136 | CWC8J0118-01.001 | OverKill | JB6K8S02 | NC01682A(COG48AA)-A-AT-N3-C | GGR320.1-B1C211 | 24 | eda show single I open/short-A6858 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/09 13:37:25 | 2022/08/09 13:56:02 | 0.31 | OK | 07764 | EDA I , 確認針痕/加壓#24(136.1) | NO | NO | NO | YES | YES | NO
EQPRepairJ22106169 | T6373-33 | CWC7J0A60-01 | Tester Down | JB1TAS14 | PA5208KF2 | J68G8 | 18 | AlarmCode:669 AlarmMessage:GP-IB COMMUNICATION TIMEOUT ERROR !! | CPAUTO | Tester | 離線(OCT/GPIB error) | 2022/08/09 13:38:53 | 2022/08/09 13:52:50 | 0.23 | OK | 10210 | 手動離線，起測OK | NO | NO | NO | YES | YES | NO
EQPRepairJ22106175 | T6373-152 | CWC8J0117-01 | OverKill | JB2P2S09 | R63455-A1S-EZ-CP | NAHM76.00-B11C | 09 | too many o/s overkill pass then o/s again-A6785 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/09 14:03:07 | 2022/08/09 14:47:33 | 0.74 | OK | 20352 | 1.清針,SOALKTIME重新對針 2.二分法 YRG=10 PM(74.0)(191.6) 3.五點水平 OD=10 | NO | NO | NO | YES | YES | NO
EQPRepairJ22106187 | T6373-34 | CWC7J0G36-01.001 | OverKill | JB1BYS10 | PA6587B_15U90H | J696A | 10 | TOO MANY O/S | A6131 | Over Kill | O/S誤宰 | 2022/08/09 14:45:13 | 2022/08/09 14:53:02 | 0.13 | OK | 17229 | 1.CLEAN UNIT ADJUST NEEDLE OD PASS. | NO | NO | NO | YES | YES | NO
EQPRepairJ22106196 | T6373-86 |  | Tester Down | JB1Q4S15 | PA6542C | J68S1 |  | 1.此機#15 #23配件送修掛配件當 2.產線通知有#22上機多根Pin OPE 3.調針、重mnt、更換配件角度相同 4.拉ND4-184 #19上機Fail Pin相同 5.確認LCD#09 Fail，此機轉Tester down請常日協助 | 10210 | Tester | LCD#09 | 2022/08/09 15:32:15 | 2022/08/09 19:18:56 | 3.78 | OK | 10210 | 1.常日SWAP LCD#09<>#05 2.上機Pass，重新降溫(159.4) | NO | NO | NO | YES | YES | NO
EQPRepairJ22106204 | T6373-13 | CWC8J0458-01 | OverKill | JB6H1S05 | NC01690G(COG48AA)-A-AT-N4-C | GGM575.1-B1C | 12 | #12 eda t44 show I NO CORR WAFER-A6132 | CPAUTO | Over Kill | Show I | 2022/08/09 16:06:07 | 2022/08/09 16:13:10 | 0.12 | OK | 10210 | 單片Show I 請產線續測 | NO | NO | NO | YES | YES | NO
EQPRepairJ22106217 | T6373-41 |  | Prober Down | JB8KZS16 | ND01657B(COF94AA)-A-C | S0V454.00-BC |  | AlarmCode:1057 AlarmMessage:MODEL NOT FOUND ( ROUGH ALIGNMENT ) !! | CPAUTO | Prober | Other | 2022/08/09 17:11:07 | 2022/08/09 17:44:14 | 0.55 | OK | 07764 | 水平重抓後 , AUTO OK #4(24.8)(248.7) 針壓有點重 , 退針壓15UM PASS | NO | NO | NO | YES | YES | NO
EQPRepairJ22106233 | T6373-147 | CWC7J0L14-01.002 | OverKill | TB1R8S85 | PA5137E1-LT | P097A.1 | 25 | #25 2H(CORR PASS)-09052 | CPAUTO | Over Kill | Function誤宰 | 2022/08/09 18:15:25 | 2022/08/09 18:33:45 | 0.31 | OK | 07764 | 2H 續測 | NO | NO | NO | YES | YES | NO
EQPRepairJ22106257 | T6373-15 | CWC8J0496-01 | OverKill | JB8Q6S04 | NA01761D(COF94AA)-A-Y-C | GGR932.1-B1C | 03 | #02 EDA SHOW I NO CORR WAFER-A6512 | CPAUTO | Over Kill | Show I | 2022/08/09 19:51:02 | 2022/08/09 19:53:14 | 0.04 | OK | 20277 | 輪工處理---- 1.單片Show I，[HW_P/DUTDIFF>4%] 非O/S造成，請產線Run Corr驗證，續測觀察 | NO | NO | NO | YES | YES | NO
EQPRepairJ22106258 | T6373-86 | CWC7J0C21-01.009 | OverKill | JB1Q4S19 | PA6542C | J68S1 | 09 | TOO MANY O/S..PM UP AND DOWN | A6673 | Over Kill | O/S誤宰 | 2022/08/09 19:55:32 | 2022/08/09 20:06:17 | 0.18 | OK | 09108 | 1.測試無異常 產線誤掛down 還機生產 | NO | NO | NO | YES | YES | NO
EQPRepairJ22106260 | T6373-12 | CWC8J0352-01.003 | OverKill | JB1PNS06 | NB01761A(COF94AA)-A-C | H2NC84.1-BC22 | 06 | TOO MANY O/S | A6941 | Over Kill | O/S誤宰 | 2022/08/09 20:05:55 | 2022/08/09 20:31:52 | 0.43 | OK | 09108 | 1.O/S 清針 調針PASS range:x28um y:18um (171.8) | NO | NO | NO | YES | YES | NO
EQPRepairJ22106261 | T6373-27 |  | Prober Down | JB8Y9S02 | ILISLQPEDQD-2CP | ILA14003.1 |  | AlarmCode:1060 AlarmMessage:CANNOT MEASURE INDEX SIZE !! | CPAUTO | Prober | Wafer Edge/Mode Not Found | 2022/08/09 20:08:58 | 2022/08/09 20:15:43 | 0.11 | OK | 16157 | 機台自動重新抓取OK，check die(42.1)OK | NO | NO | NO | YES | YES | NO
EQPRepairJ22106279 | T6373-129 | CWC7J0H70-01.004 | Setup Fail | TB1INS29 | NA01404H-A-AT-C | GGT500.1-B1 | 07 | ALREADY CHANGE P/C DUE TO FAILBIN 20(TB1IN28-->TB1IN29)ALREADY FOLLOW SOP.WAIT 20MINS SITLL WITH ICE | A6936 | Prober | Other | 2022/08/09 21:03:43 | 2022/08/09 22:27:36 | 1.40 | OK | 10897 | <<低溫換卡>>改機不過更換備卡結霜掛當 1.升高溫100度半小時，在降溫，重新對針查看bump無結霜還產線改機 | NO | NO | NO | NO | YES | NO
EQPRepairJ22106285 | T6373-47 | CWC7J0795-01 | OverKill | JB1HKS09 | NA01687G(COG48BA)-A-AT-N2-C | GGP605.1-BC11 | 10 | too many os-A6512 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/09 21:36:50 | 2022/08/09 21:48:48 | 0.20 | OK | 20277 | Remote---- 1.針位貼下 Range:X>20，Y:14 機台清針 點測五點PASS穩定，還機生產 | NO | NO | NO | YES | YES | NO
EQPRepairJ22106297 | T6373-141 | CWC8J0248-01.002 | OverKill | JB1HKS10 | NA01687G(COG48BA)-A-AT-N2-C | GGS267.1-B11C22 | 14 | o/s-07079 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/09 22:24:49 | 2022/08/09 22:54:51 | 0.50 | OK | 20277 | 輪工處理---- 1.調整針位、針壓 機台清針 Range:14 PASS | NO | NO | NO | YES | YES | NO
EQPRepairJ22106300 | T6373-25 | CWC8J0271-01.001 | OverKill | JB1LWS14 | ILISLQPEDQD-2CP | ILA13192.1 | 21 | o/s-08438 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/09 22:32:45 | 2022/08/09 22:46:13 | 0.22 | OK | 20277 | 輪工處理---- 1.調整針、位針壓 Range:14 起測一段PASS穩定，還機生產 | NO | NO | NO | YES | YES | NO
EQPRepairJ22106303 | T6373-33 | CWC7J0A60-01 | OverKill | JB1TAS14 | PA5208KF2 | J68G8 | 22 | too many O/S-A6857 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/09 22:40:14 | 2022/08/09 22:55:32 | 0.26 | OK | 09108 | 1.o/s 前蓋清針 機台清針pass (45.6) | NO | NO | NO | YES | YES | NO
EQPRepairJ22106306 | T6373-41 | CWC8J0484-01 | OverKill | JB8KZS16 | ND01657B(COF94AA)-A-C | S0V454.00-BC | 08 | #8 TOO MANY O/S (8SITE)-A6441 | CPAUTO | Over Kill | O/D不足 | 2022/08/09 23:09:19 | 2022/08/09 23:20:59 | 0.19 | OK | 20182 | OD-.10 PASS | NO | NO | NO | YES | YES | NO
EQPRepairJ22106309 | T6373-111 | CWC7J0A56-01 | OverKill | TB1R8S74 | PA5137F-LT1 | P094K.1 | 03 | EDA SHOW 2H | A6857 | Over Kill | Show H | 2022/08/09 23:19:03 | 2022/08/09 23:21:57 | 0.05 | OK | 09108 | 1.eda Show H 清針 續測 | NO | NO | NO | YES | YES | NO
EQPRepairJ22106312 | T6373-12 | CWC8J0352-01.003 | OverKill | JB1PNS06 | NB01761A(COF94AA)-A-C | H2NC84.1-BC22 | 07 | TOO MANY O/S..PM CHECK UP AND DOWN..OD 5-A6673 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/09 23:30:50 | 2022/08/09 23:48:31 | 0.29 | OK | 10897 | 機台清針，二分法。試打(53.8)STDF點測五點PASS | NO | NO | NO | YES | YES | NO
EQPRepairJ22106314 | T6373-38 | CWC7J0677-01 | OverKill | JB1GKS10 | PA6543B | J67H6 | 08 | EDA CONTINUE 2H | A7023 | Over Kill | 2H | 2022/08/09 23:33:22 | 2022/08/09 23:42:42 | 0.16 | OK | 15492 | 1.零星O/S，清、調針 PASS 2.P/M(158，6)OK 3.續測觀察 | NO | NO | NO | YES | YES | NO
EQPRepairJ22106320 | T6373-86 | CWC7J0C21-01.009 | OverKill | JB1Q4S19 | PA6542C | J68S1 | 10 | too many o/s-A6673 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/09 23:44:43 | 2022/08/10 00:30:32 | 0.76 | OK | 10897 | 機台清針，測試一段OK，可正常回收O/S | NO | NO | NO | NO | YES | NO
EQPRepairJ22106328 | T6373-15 | CWC8J0496-01 | OverKill | JB8Q6S04 | NA01761D(COF94AA)-A-Y-C | GGR932.1-B1C1 | 06 | #05 EDA SHOW I | 03491 | Over Kill | Show I | 2022/08/10 00:19:56 | 2022/08/10 00:28:21 | 0.14 | OK | 20277 | 輪工處理---- 1.#05 Show I [HW_P/DUTDIFF>4%]，前一片並無此問題，非O/S造成，請產線Run Corr驗證，續測觀察 | NO | NO | NO | YES | YES | NO
EQPRepairJ22106332 | T6373-129 | CWC7J0H70-01.004 | OverKill | TB1INS29 | NA01404H-A-AT-C | GGT500.1-B1C212 | 07 | too many os-A6446 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/10 00:30:36 | 2022/08/10 00:52:52 | 0.37 | OK | 09108 | 1.o/s 手動執行 soaktime 調針pass (140.6) | NO | NO | NO | YES | YES | NO
EQPRepairJ22106355 | T6373-141 | CWC8J0248-01.002 | Tester Down | JB1HKS10 | NA01687G(COG48BA)-A-AT-N2-C | GGS267.1-B11C22 | 14 | GP-IB-10851 | CPAUTO | Tester | 離線(OCT/GPIB error) | 2022/08/10 01:11:44 | 2022/08/10 01:25:23 | 0.23 | OK | 16157 | T44 Tester系統卡關，#14 wafer alignment請產線重新上貨 | NO | NO | NO | YES | YES | NO
EQPRepairJ22106362 | T6373-111 |  | Corr Fail | TB1R8S74 | PA5137F-LT1 | P094K.1 |  | Corr Fail | A6857 | Corr | Corr Fail | 2022/08/10 01:44:51 | 2022/08/10 01:50:26 | 0.09 | OK | 09108 | Corr Fail BIN127 產品問題 扣除BIN127 比對PASS | NO | NO | NO | YES | YES | NO
EQPRepairJ22106373 | T6373-25 | CWC8J0271-01.001 | OverKill | JB1LWS14 | ILISLQPEDQD-2CP | ILA13192.1 | 22 | TOO MANY O/S | A7023 | Over Kill | O/S誤宰 | 2022/08/10 02:35:33 | 2022/08/10 02:52:40 | 0.29 | OK | 15492 | 1.清、調針 PASS 2.P/M(234，10)OK | NO | NO | NO | YES | YES | NO
EQPRepairJ22106377 | T6373-147 | CWC7J0J51-01 | OverKill | TB1R8S85 | PA5137F-LT1 | P094H.1 | 11 | o/s-03491 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/10 02:49:21 | 2022/08/10 03:50:29 | 1.02 | OK | 20277 | 輪工處理---- 1.下卡磨針 上機PASS Range:10 | NO | NO | NO | YES | YES | NO
EQPRepairJ22106384 | T6373-38 | CWC7J0677-01 | OverKill | JB1GKS10 | PA6543B | J67H6 | 09 | O/S-20372 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/10 03:20:37 | 2022/08/10 03:36:05 | 0.26 | OK | 16157 | O/S 機台清針，調整針位pass | NO | NO | NO | YES | YES | NO
EQPRepairJ22106385 | T6373-86 | CWC7J0C21-01.009 | OverKill | JB1Q4S19 | PA6542C | J68S1 | 12 | too many o/s-A6941 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/10 03:28:59 | 2022/08/10 03:47:25 | 0.31 | OK | 09108 | 1.O/S 清針 調針PASS range:10um (129.5) | NO | NO | NO | YES | YES | NO
EQPRepairJ22106386 | T6373-35 | CWC8J0189-01.004 | OverKill | JB1GVS22 | NC01682A(COG48AA)-A-AT-N2-C | GGS284.1-B1C222 | 15 | too many o/s-10851 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/10 03:36:27 | 2022/08/10 03:45:33 | 0.15 | OK | 16157 | O/S 機台清針，二分法調針Range:12um pass | NO | NO | NO | YES | YES | NO
EQPRepairJ22106400 | T6373-21 | CMC8J0010-01 | OverKill | JB1GKS08 | PA6543B | J6630 | 12 | 2H-0358 | CPAUTO | Over Kill | 2H | 2022/08/10 04:54:02 | 2022/08/10 05:00:52 | 0.11 | OK | 15492 | 1.[HW_BIN60<98%](Bin 61 < 1109EA) 1052 2.CORR. WAFER PASSED，此為 CM LOT 3.待換批觀察 | NO | NO | NO | YES | YES | NO
EQPRepairJ22106426 | T6373-13 | CWC8J0503-01 | OverKill | JB6H1S05 | NC01690F(COG46BA)-A-AT-C | GGY201.1-B12C | 05 | eda show 2H no corr | A6446 | Over Kill | Show H | 2022/08/10 06:00:04 | 2022/08/10 06:21:09 | 0.35 | OK | 09108 | 1.bin26 eda Show H 無固定dut fail 續測 | NO | NO | NO | YES | YES | NO
EQPRepairJ22106433 | T6373-86 | CWC7J0C21-01.009 | OverKill | JB1Q4S19 | PA6542C | J68S1 | 13 | # 11-12 eda show 2R (O/S)-A6673 | CPAUTO | Over Kill | Show R | 2022/08/10 06:17:07 | 2022/08/10 06:58:30 | 0.69 | OK | 09108 | 1.EDA Show R 機台清針 重新對針 SETUP PASS range:10um (159.4) | NO | NO | NO | YES | YES | NO
EQPRepairJ22106446 | T6373-47 | CWC8J0348-01.002 | OverKill | JB1HKS09 | NA01687G(COG48BA)-A-AT-N2-C | GGR356.1-B11C12 | 14 | #10.11 (2r) (o/s)-03491 | CPAUTO | Over Kill | Other | 2022/08/10 06:54:46 | 2022/08/10 07:45:58 | 0.85 | OK | 07764 | Remote處理----2r bin22 與bin29 bin29 fail 同批於不同機台測試相同 bin22 open 確認針痕與清針 , 修改清針參數 , 續測觀察 | NO | NO | NO | YES | YES | NO
EQPRepairJ22106452 | T6373-37 | CWC8J0189-01.003 | OverKill | JB1GVS23 | NC01682A(COG48AA)-A-AT-N2-C | GGS284.1-B1C221 | 06 | EDA SHOW 2H NO CORR-10851 | CPAUTO | Over Kill | SHOW H | 2022/08/10 07:07:46 | 2022/08/10 07:32:41 | 0.42 | OK | 13581 | 1.EDA SHOW H PLEASE RUN CORR | NO | NO | NO | NO | YES | NO
EQPRepairJ22106467 | T6373-38 | CWC7J0677-01 | OverKill | JB1GKS10 | PA6543B | J67H6 | 13 | TOO MANY O/S-A6920 | CPAUTO | Over Kill | O/S | 2022/08/10 07:45:42 | 2022/08/10 08:17:45 | 0.53 | OK | 13222 | 機台清針OK | NO | NO | NO | YES | YES | NO
EQPRepairJ22106470 | T6373-86 | CWC7J0C21-01.009 | OverKill | JB1Q4S19 | PA6542C | J68S1 | 13 | #13 EDA CONTINUES SHOW 2R OS-A6132 | CPAUTO | Over Kill | 2R | 2022/08/10 07:58:37 | 2022/08/10 08:07:51 | 0.15 | OK | 15362 | 2R O/S,清調針,Y RANGE:10,P/M(173,3),續測觀察 | NO | NO | NO | YES | YES | NO
EQPRepairJ22106474 | T6373-111 | CWC7J0A56-01 | OverKill | TB1R8S74 | PA5137F-LT1 | P094K.1 | 11 | #11 eda continues show 2h no corr wafer-A6132 | CPAUTO | Over Kill | Show H | 2022/08/10 08:15:25 | 2022/08/10 08:22:34 | 0.12 | OK | 10210 | 2H 請產線續測 | NO | NO | NO | YES | YES | NO
EQPRepairJ22106500 | T6373-26 | CWC8J0345-01 | Prober Down | JB1LWS17 | ILISLQPEDQD-2CP | ILA14002.1 | 14 | CANNOT MEASURE INDEX SIZE-A6660 | CPAUTO | Prober | Wafer Edge/Mode Not Found | 2022/08/10 10:24:58 | 2022/08/10 10:51:08 | 0.44 | OK | 13222 | Wafer Edge/Mode Not Found 重抓OK | NO | NO | NO | YES | YES | NO
EQPRepairJ22106507 | T6373-34 | CWC8J0168-01.001 | OverKill | JB1BYS10 | PA6587B_15U90H | J6969 | 18 | TOO MANY O/S-A6564 | CPAUTO | Over Kill | 機台清針後Pass | 2022/08/10 10:47:55 | 2022/08/10 10:59:29 | 0.19 | OK | 13222 | 機台清針後Pass | NO | NO | NO | YES | YES | NO
EQPRepairJ22106508 | T6373-86 | CWC7J0C21-01.009 | OverKill | JB1Q4S19 | PA6542C | J68S1 | 17 | TOO MANY OS-A6564 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/10 10:51:44 | 2022/08/10 11:31:52 | 0.67 | OK | 20075 | 重新對針RANGE 12,MARK -1,3 | NO | NO | NO | YES | YES | NO
EQPRepairJ22106517 | T6373-33 | CWC7J0A87-01 | OverKill | JB1TAS14 | PA5208KF2 | J6875 | 04 | eda show 2 h-A6293 | CPAUTO | Over Kill | 2H | 2022/08/10 11:40:27 | 2022/08/10 11:49:52 | 0.16 | OK | 15362 | 2H兩片不同FUNCTION造成,清針,續測觀察 | NO | NO | NO | YES | YES | NO
EQPRepairJ22106527 | T6373-86 | CWC7J0C21-01.009 | OverKill | JB1Q4S19 | PA6542C | J68S1 | 20 | too many o/s-A6293 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/10 12:29:35 | 2022/08/10 12:38:35 | 0.15 | OK | 15362 | clean unit pass,p/m(159,4) | NO | NO | NO | YES | YES | NO
EQPRepairJ22106530 | T6373-34 | CWC8J0168-01.001 | OverKill | JB1BYS10 | PA6587B_15U90H | J6969 | 19 | EDA SHOW 2 CONTINUES R-A6564 | CPAUTO | Over Kill | eda 2r | 2022/08/10 12:58:42 | 2022/08/10 13:14:48 | 0.27 | OK | 20230 | 1.eda 2r 2.查看eda 零星o/s,機台清針,二分調針rg=13,pm(46.7) 3.pass | NO | NO | NO | YES | YES | NO
EQPRepairJ22106546 | T6373-15 | CWC8J0496-01 | OverKill | JB8Q6S04 | NA01761D(COF94AA)-A-Y-C | GGR932.1-B1C11 | 15 | #15 T44 EDA SHOW I ALREADY CORR PASSED-A6132 | CPAUTO | Over Kill | EDA I | 2022/08/10 13:59:56 | 2022/08/10 14:28:23 | 0.47 | OK | 20230 | 1.EDA I 2.機台清針,二分調針YRG=19,XRG=23,PM(165.8) 3.PASS | NO | NO | NO | YES | YES | NO
EQPRepairJ22106554 | T6373-39 | CWC7J0F95-01 | OverKill | JB1TAS21 | PA5208KF2 | J68G6 | 09 | TOO MANY O/S | A6919 | Over Kill | O/S誤宰 | 2022/08/10 14:43:31 | 2022/08/10 15:08:23 | 0.41 | OK | 20230 | 1.O/S 2.機台清針,二分法調針RG=14,PM(147.5) 3.PASS | NO | NO | NO | YES | YES | NO
EQPRepairJ22106566 | T6373-29 | CWC8J0134-01.001 | OverKill | JB1QNS04 | NA01687G(COG48BA)-A-AT-N2-C | GGS940.1-BC21 | 03 | TOO MANY FAIL BIN6-A6564 | CPAUTO | Over Kill | bin6 | 2022/08/10 15:44:44 | 2022/08/10 15:56:21 | 0.19 | OK | 13222 | bin6 機台清針ok | NO | NO | NO | YES | YES | NO
EQPRepairJ22106575 | T6373-90 | CWC5J0W76-01 | Corr Fail | JB1GVS26 | NC01682A(COG48AA)-A-AT-N3-C | GGG835.1-B | 16 | corr-08889 | CPAUTO | Corr | corr fail | 2022/08/10 16:30:32 | 2022/08/10 16:33:15 | 0.05 | OK | 20230 | 手動比對pass | NO | NO | NO | YES | YES | NO
EQPRepairJ22106579 | T6373-42 | CWC8J0291-01.002 | OverKill | JB1HKS06 | NA01687F(COG48BA)-A-AT-C | GGU014.1-BC22 | 03 | #3 PLS CHECK AND CONFIRM PN TOO MANY FAIL BIN 29-A6132 | CPAUTO | Over Kill | Function誤宰 | 2022/08/10 16:39:31 | 2022/08/10 18:04:18 | 1.41 | OK | 15362 | BIN29更換配件#07 PASS,Y RANGE:16,P/M(171,3) | NO | NO | NO | YES | YES | NO
EQPRepairJ22106582 | T6373-29 | CWC8J0134-01.001 | OverKill | JB1QNS04 | NA01687G(COG48BA)-A-AT-N2-C | GGS940.1-BC21 | 04 | #2.3 EDA 2H NO CORR | A6920 | Over Kill | eda 2h | 2022/08/10 17:07:50 | 2022/08/10 19:00:12 | 1.87 | OK | 20230 | 1.eda 2h 2.更換備卡s05,RG=16,PM(36.4) 3.PASS | NO | NO | NO | YES | YES | NO
EQPRepairJ22106584 | T6373-90 | CWC6J0996-01.001 | OverKill | JB1GVS26 | NC01682A(COG48AA)-A-AT-N3-C | GGG836.1-B1C211 | 02 | too many fail bin50-A6564 | CPAUTO | Over Kill | BIN50 | 2022/08/10 17:17:33 | 2022/08/10 17:28:11 | 0.18 | OK | 13222 | 機台清針、調針RANGE=8 (138.6 | NO | NO | NO | YES | YES | NO
EQPRepairJ22106590 | T6373-86 | CWC7J0C21-01 | OverKill | JB1Q4S19 | PA6542C | J68S1 | 01 | 01 ALREADY HANDLING OVERKILL PROCEDURE STILL NOT EFFECTIVE O/S TOO MANY-A6143 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/10 17:32:12 | 2022/08/10 20:52:09 | 3.33 | OK | 10897 | 日班交接O/S換備卡#16上機無法PASS 1.<<低溫換卡>>夜班接手更換配件#19等待20分與降溫時間後續上機2分法PASSS，試打(43.5) | NO | NO | NO | NO | YES | NO
EQPRepairJ22106596 | T6373-26 | CWC8J0345-01 | OverKill | JB1LWS17 | ILISLQPEDQD-2CP | ILA14002.1 | 17 | TOO MANY O/S-A6564 | CPAUTO | Over Kill | 針位調整 | 2022/08/10 17:57:18 | 2022/08/10 18:15:31 | 0.30 | OK | 13222 | 針位調整 RANGE=20 (178.13) | NO | NO | NO | YES | YES | NO
EQPRepairJ22106599 | T6373-12 | CWC8J0496-01.002 | Other | JB1PNS06 | NA01761D(COF94AA)-A-Y-C | GGR932.1-B1C12 | 03 | PM CHECK TOO BIG UP AND DOWN (26,1) PLEASE CONFIRM THANKS | A6785 | Other | 確認針痕 | 2022/08/10 18:06:03 | 2022/08/10 18:23:18 | 0.29 | OK | 15362 | #04 二分法X/Y RANGE:28/18,P/M(218,3) | NO | NO | NO | YES | YES | NO
EQPRepairJ22106600 | T6373-39 | CWC7J0F95-01 | OverKill | JB1TAS21 | PA5208KF2 | J68G6 | 10 | too many o/s | A6919 | Over Kill | Other | 2022/08/10 18:07:09 | 2022/08/10 18:40:57 | 0.56 | OK | 07764 | open 調針位pass #10(19.5) | NO | NO | NO | YES | YES | NO
EQPRepairJ22106609 | T6373-111 |  | Prober Down | TB1R8S74 | PA5137F-LT1 | P094K.1 |  | CPU 異常 | 18966 | Prober | 面板Down (Monitor Down) | 2022/08/10 18:41:36 | 2022/08/10 18:51:32 | 0.17 | OK | 18966 | 重開機正常 以開至第二段 | NO | NO | NO | NO | NO | NO
EQPRepairJ22106612 | T6373-147 | CWC7J0336-01.004 | OverKill | TB1R8S85 | PA5137E1-LT | P097C.1 | 12 | o/s | 07854 | Over Kill | O/S誤宰 | 2022/08/10 18:45:28 | 2022/08/10 18:52:59 | 0.13 | OK | 07764 | open 調整針位 | NO | NO | NO | YES | YES | NO
EQPRepairJ22106614 | T6373-42 | CWC8J0291-01.002 | OverKill | JB1HKS07 | NA01687F(COG48BA)-A-AT-C | GGU014.1-BC22 | 05 | EDA SHOWS 2 H O/S-A6919 | CPAUTO | Over Kill | H | 2022/08/10 18:52:32 | 2022/08/10 18:56:06 | 0.06 | OK | 17229 | 1.剛更換配件,先行還機待觀察當片. | NO | NO | NO | YES | YES | NO
EQPRepairJ22106616 | T6373-121 | CWC8J0133-01.002 | OverKill | JB1QNS01 | NA01687G(COG48BA)-A-AT-N2-C | GGS416.1-BC22 | 17 | EDA SHOW I O/S MODE TEST | A6785 | Over Kill | Show I | 2022/08/10 19:00:02 | 2022/08/10 19:25:57 | 0.43 | OK | 09108 | 1.eda Show I [!!! WARNING !!!OPEN AND SHORT MODE TEST] 續測 | NO | NO | NO | YES | YES | NO
EQPRepairJ22106617 | T6373-34 | CWC8J0168-01.001 | OverKill | JB1BYS10 | PA6587B_15U90H | J6969 | 21 | TOO MANY O/S-A6564 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/10 19:08:18 | 2022/08/10 20:00:05 | 0.86 | OK | 15492 | 1.下卡磨針，二分法(0，2)、RANGE:8 2.P/M(172，5)OK | NO | NO | NO | YES | YES | NO
EQPRepairJ22106618 | T6373-21 | CWC7J0K21-01 | OverKill | JB1TAS16 | PA5208KF-LT | J696N | 16 | too many o/s | A6660 | Over Kill | O/S誤宰 | 2022/08/10 19:16:32 | 2022/08/10 19:50:41 | 0.57 | OK | 15492 | 1.AUTO ONLINE INIT 2.O/S，清、調針 PASS 3.P/M(118,6)OK | NO | NO | NO | YES | YES | NO
EQPRepairJ22106628 | T6373-39 | CWC7J0F95-01 | OverKill | JB1TAS21 | PA5208KF2 | J68G6 | 11 | TOO MANY O/S..NO PARTICLE-A6516 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/10 20:11:43 | 2022/08/10 20:30:49 | 0.32 | OK | 20182 | 1.清針調針RG-8 PM(46.0) | NO | NO | NO | YES | YES | NO
EQPRepairJ22106644 | T6373-47 | CWC8J0348-01.002 | OverKill | JB1HKS09 | NA01687G(COG48BA)-A-AT-N2-C | GGR356.1-B11C12 | 24 | too many o/s pm up and down pls check | A6675 | Over Kill | O/S誤宰 | 2022/08/10 21:17:36 | 2022/08/10 21:38:48 | 0.35 | OK | 20277 | Remote---- 1.針位貼上 Range:14 機台清針 PASS | NO | NO | NO | YES | YES | NO
EQPRepairJ22106659 | T6373-39 | CWC7J0F95-01 | OverKill | JB1TAS21 | PA5208KF2 | J68G6 | 11 | TOO MANY OS | A6446 | Over Kill | O/S誤宰 | 2022/08/10 22:28:44 | 2022/08/10 22:43:12 | 0.24 | OK | 20277 | Remote---- 1.真偽貼上 Range:14 機台清針 起側一段PASS穩定，還機生產 | NO | NO | NO | YES | YES | NO
EQPRepairJ22106668 | T6373-111 |  | Other | TB1R8S74 | PA5137F-LT1 | P094K.1 |  | WAIT AE CHECK 3D | A6441 | Other | Other | 2022/08/10 22:51:18 | 2022/08/10 23:07:15 | 0.27 | OK | 09108 | 1.設定3D清針OK | NO | NO | NO | YES | YES | NO
EQPRepairJ22106681 | T6373-111 | CWC7J0762-01 | Other | TB1R8S74 | PA5137F-LT1 | P094K.1 | 11 | ice-10851 | CPAUTO | Prober | 結霜 | 2022/08/10 23:30:17 | 2022/08/11 01:32:13 | 2.03 | OK | 17184 | 1.升溫烘烤，ReS/U(46.0) | NO | NO | NO | NO | YES | NO
EQPRepairJ22106692 | T6373-34 | CWC8J0168-01.001 | OverKill | JB1BYS10 | PA6587B_15U90H | J6969 | 23 | # Already follow Overkilll handling procedures but still not effective. Too O/S | A6944 | Over Kill | O/S誤宰 | 2022/08/11 00:02:39 | 2022/08/11 00:34:18 | 0.53 | OK | 15492 | 1. 零星O/S，機台清針 PASS | NO | NO | NO | YES | YES | NO
EQPRepairJ22106732 | T6373-141 | CWC8J0391-01.001 | OverKill | JB1HKS10 | NA01687G(COG48BA)-A-AT-N2-C | GGS419.1-B11C2 | 07 | EDA Show 2R-A6318 | CPAUTO | Over Kill | 2R | 2022/08/11 02:05:40 | 2022/08/11 02:24:47 | 0.32 | OK | 15492 | 1.零星少數 O/S，清、調針 PASS，RANGE:12 2.P/M(146，4)OK | NO | NO | NO | YES | YES | NO
EQPRepairJ22106747 | T6373-11 | CWC8J0352-01 | OverKill | JB1PNS05 | NB01761A(COF94AA)-A-C | H2NC84.1-BC11 | 25 | PROBE MARK N.G PLEASE CHECK (158.7) | A6675 | Over Kill | Other | 2022/08/11 03:11:39 | 2022/08/11 04:56:18 | 1.74 | OK | 09108 | 1.產線掛down針位不佳 認針痕 針位貼下 2.因產線放置看BUMP畫面過久影像模糊 故重新對針 下針壓至+25無contact 無針痕 3.退片 Prober reset 重新SETUP PASS range: x>20 y:16um (52.8) #25重測 | NO | NO | NO | YES | YES | NO
EQPRepairJ22106754 | T6373-42 | CWC8J0291-01.002 | OverKill | JB1HKS07 | NA01687F(COG48BA)-A-AT-C | GGU014.1-BC22 | 14 | too many fail bin 29.-A7048 | CPAUTO | Over Kill | Function誤宰 | 2022/08/11 03:20:04 | 2022/08/11 03:51:21 | 0.52 | OK | 20277 | Remote---- 1.機台清針 Range:16 PASS | NO | NO | NO | YES | YES | NO
EQPRepairJ22106762 | T6373-86 | CWC7J0C21-01 | OverKill | JB1Q4S19 | PA6542C | J68S1 | 16 | TOO MANY O/S-A6938 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/11 03:57:04 | 2022/08/11 04:17:43 | 0.34 | OK | 17184 | 1.清針&調針(43.5) | NO | NO | NO | NO | YES | NO
EQPRepairJ22106764 | T6373-39 | CWC7J0F95-01 | OverKill | JB1TAS21 | PA5208KF2 | J68G6 | 13 | O/S | 03491 | Over Kill | O/S誤宰 | 2022/08/11 04:02:04 | 2022/08/11 04:18:43 | 0.28 | OK | 17091 | REMOTE----機台清針PASS | NO | NO | NO | YES | YES | NO
EQPRepairJ22106772 | T6373-21 | CWC7J0K21-01 | OverKill | JB1TAS16 | PA5208KF-LT | J696N | 19 | Already follow Overkilll handling procedures but still not effective. Too O/S | A6944 | Over Kill | O/S誤宰 | 2022/08/11 04:21:37 | 2022/08/11 04:34:27 | 0.21 | OK | 15492 | 1.零星O/S，清針 PASS 2.P/M(160,4)OK | NO | NO | NO | YES | YES | NO
EQPRepairJ22106777 | T6373-01 | CWC7J0K23-01 | OverKill | JB1TAS23 | PA5208KF-LT | J68QY | 18 | Already follow Overkilll handling procedures but still not effective. Too many o/s | A6944 | Over Kill | O/S誤宰 | 2022/08/11 04:43:58 | 2022/08/11 05:11:49 | 0.46 | OK | 15492 | 1.針位貼左，調針位 PASS 2.P/M(148，5)OK | NO | NO | NO | YES | YES | NO
EQPRepairJ22106779 | T6373-34 | CWC8J0168-01.001 | OverKill | JB1BYS10 | PA6587B_15U90H | J6969 | 25 | Already follow Overkilll handling procedures but still not effective. Too Many O/S | A6944 | Over Kill | O/S誤宰 | 2022/08/11 04:52:14 | 2022/08/11 05:28:31 | 0.60 | OK | 15492 | 1.清、調針 PASS 2.P/M(160,6)OK | NO | NO | NO | YES | YES | NO
EQPRepairJ22106780 | T6373-42 | CWC8J0291-01.002 | OverKill | JB1HKS07 | NA01687F(COG48BA)-A-AT-C | GGU014.1-BC22 | 15 | #12,14 (2H)(NO CORR) | 03491 | Over Kill | 2h | 2022/08/11 04:59:00 | 2022/08/11 06:22:10 | 1.39 | OK | 20182 | 2H Bin29過多更換備卡#S06，上機PASS 還機生產 觀察當片 | NO | NO | NO | YES | YES | NO
EQPRepairJ22106792 | T6373-35 |  | Corr Fail | JB1GVS22 | NC01682A(COG48AA)-A-AT-N3-C | GGP620.1-2B |  | Corr Fail | A6943 | Corr | corr. faIL | 2022/08/11 05:37:30 | 2022/08/11 05:41:01 | 0.06 | OK | 15492 | 1.比對 PASS | NO | NO | NO | YES | YES | NO
EQPRepairJ22106824 | T6373-33 |  | Other | JB1TAS14 | PA5208KF2 | J6875 |  | 3D 清針參數設定，待輪工確認 | 10851 | Other | Other | 2022/08/11 07:16:10 | 2022/08/11 07:24:02 | 0.13 | OK | 15362 | 設定3D清針 | NO | NO | NO | YES | YES | NO
EQPRepairJ22106826 | T6373-147 | CWC7J0672-01.002 | OverKill | TB1R8S85 | PA5137F-LT1 | P07NT.1 | 25 | O/S-07854 | CPAUTO | Over Kill | o/s | 2022/08/11 07:31:40 | 2022/08/11 08:05:42 | 0.57 | OK | 13222 | 1.配件#85 上貨測試VGL open bin2 2.更換備卡#82 上機 PASS 3.配件#85 送修 | NO | NO | NO | YES | YES | NO
EQPRepairJ22106828 | T6373-35 | CWC8J0118-01.001 | OverKill | JB1GVS22 | NC01682A(COG48AA)-A-AT-N3-C | GGR320.1-B1C211 | 24 | too many os-A6564 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/11 07:35:37 | 2022/08/11 07:47:51 | 0.20 | OK | 20230 | 1.O/S 2.機台清針,二分法調針RG=13,PM(47.6) 3.PASS | NO | NO | NO | YES | YES | NO
EQPRepairJ22106839 | T6373-41 |  | Prober Down | JB8KZS16 | ND01657B(COF94AA)-A-C | S0V380.00-21B11C |  | AlarmCode:1057 AlarmMessage:MODEL NOT FOUND ( ROUGH ALIGNMENT ) !! | CPAUTO | Prober | Wafer Edge/Mode Not Found | 2022/08/11 08:09:12 | 2022/08/11 08:25:08 | 0.27 | OK | 20075 | 抓mode ok | NO | NO | NO | YES | YES | NO
EQPRepairJ22106843 | T6373-86 | CWC7J0C21-01 | OverKill | JB1Q4S19 | PA6542C | J68S1 | 21 | TOO MNAY O/S PM CHECK UP AND DOWN | A6785 | Over Kill | O/S誤宰 | 2022/08/11 08:22:46 | 2022/08/11 09:19:21 | 0.94 | OK | 20230 | 1.O/S 2.機台清針,重新對針,二分法調針YRG=11,XRG=26,PM(44.5) 3.PASS | NO | NO | NO | YES | YES | NO
EQPRepairJ22106847 | T6373-38 | CWC6J0K13-01 | OverKill | JB1GKS10 | PA6543B | J6634 | 25 | TOO MANY OS-A6564 | CPAUTO | Over Kill | 機台清針後Pass | 2022/08/11 08:29:08 | 2022/08/11 09:14:07 | 0.75 | OK | 13222 | 機台清針後Pass | NO | NO | NO | YES | YES | NO
EQPRepairJ22106854 | T6373-21 | CWC7J0K21-01 | OverKill | JB1TAS16 | PA5208KF-LT | J696N | 20 | TOO MANY O/S-A6564 | CPAUTO | Over Kill | o/s | 2022/08/11 08:49:44 | 2022/08/11 09:14:51 | 0.42 | OK | 13222 | 機台清針後Pass | NO | NO | NO | YES | YES | NO
EQPRepairJ22106871 | T6373-141 | CWC8J0391-01.001 | OverKill | JB1HKS10 | NA01687G(COG48BA)-A-AT-N2-C | GGS419.1-B11C2 | 12 | too many os-A6564 | CPAUTO | Over Kill | 機台清針後Pass | 2022/08/11 10:10:36 | 2022/08/11 10:20:46 | 0.17 | OK | 13222 | 機台清針後Pass (9.1) | NO | NO | NO | YES | YES | NO
EQPRepairJ22106874 | T6373-121 | CWC8J0248-01.002 | OverKill | JB1QNS01 | NA01687G(COG48BA)-A-AT-N2-C | GGS267.1-B11C22 | 06 | EDA SHOW I WARNING O/S MODE TEST-A6131 | CPAUTO | Over Kill | Show I | 2022/08/11 10:18:29 | 2022/08/11 10:25:38 | 0.12 | OK | 10210 | eng mode check ok | NO | NO | NO | YES | YES | NO
EQPRepairJ22106882 | T6373-35 | CWC8J0135-01.001 | OverKill | JB1GVS22 | NC01682A(COG48AA)-A-AT-N3-C | GGP621.1-BC2111 | 01 | TOO MANY O/S-A6564 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/11 10:59:49 | 2022/08/11 11:29:35 | 0.50 | OK | 20230 | 1.o/s 2.機台清針,二分法調針rg=11,pm(47.6) 3.pass | NO | NO | NO | YES | YES | NO
EQPRepairJ22106888 | T6373-34 |  | Prober Down | JB1BYS10 | PA6587B_15U90H | J696A |  | AlarmCode:1060 AlarmMessage:CANNOT MEASURE INDEX SIZE !! | CPAUTO | Prober | cannot measure index size | 2022/08/11 11:22:15 | 2022/08/11 11:49:15 | 0.45 | OK | 20230 | cannot measure index size重抓ok check die ok 二分法調針 yrg=11,xrg=60,pm(132.7) | NO | NO | NO | YES | YES | NO
EQPRepairJ22106898 | T6373-16 | CWC7J0K81-01.001 | Other | JB1TAS26 | PA5208KF-LT | J69H8 | 17 | ae change PC-A6873 | CPAUTO | Other | Other | 2022/08/11 12:30:07 | 2022/08/11 12:40:53 | 0.18 | OK | 13222 | Other--- 1. 常日設備要求更換配件#28 上機PASS | NO | NO | NO | YES | YES | NO
EQPRepairJ22106902 | T6373-86 | CWC7J0C21-01 | OverKill | JB1Q4S19 | PA6542C | J68S1 | 23 | TOO MANY O/S | A6131 | Over Kill | O/S誤宰 | 2022/08/11 12:35:21 | 2022/08/11 13:01:08 | 0.43 | OK | 10210 | Unit清針 | NO | NO | NO | YES | YES | NO
EQPRepairJ22106912 | T6373-47 | CWC8J0133-01.003 | OverKill | JB1HKS09 | NA01687G(COG48BA)-A-AT-N2-C | GGS416.1-BC12 | 11 | TOO MANY O/S-A6919 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/11 13:09:28 | 2022/08/11 13:33:55 | 0.41 | OK | 14291 | clean wafer pass P/M(36.0) | NO | NO | NO | YES | YES | NO
EQPRepairJ22106923 | T6373-35 | CRC7J0011-01 | OverKill | JB1GVS22 | NC01682A(COG48AA)-A-AT-N3-C | GCVP99.1-BC111121C | 01 | TOO MANY O/S-A6564 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/11 13:36:01 | 2022/08/11 14:07:24 | 0.52 | OK | 13581 | 1.使用機台清針+調針位PASS(138.6) | NO | NO | NO | NO | YES | NO
EQPRepairJ22106926 | T6373-12 | CWC8J0496-01 | Other | JB1PNS06 | NA01761D(COF94AA)-A-Y-C | GGR932.1-B1C11 | 02 | PLEASE CONFIRM PROBEMARK. PM IS TOO BIG.-A6131 | CPAUTO | Other | 確認針痕 | 2022/08/11 13:58:50 | 2022/08/11 15:01:19 | 1.04 | OK | 15362 | #02 溫針,重新對針,x/y range:24/14,p/m(167,8) | NO | NO | NO | YES | YES | NO
EQPRepairJ22106929 | T6373-33 |  | Prober Down | JB1TAS14 | PA5208KF2 | J6875 |  | AlarmCode:1058 AlarmMessage:MODEL NOT FOUND ( FINE ALIGNMENT ) !! | CPAUTO | Prober | Wafer Edge/Mode Not Found | 2022/08/11 14:16:27 | 2022/08/11 14:21:05 | 0.08 | OK | 10210 | 重抓MODE OK(118.0) | NO | NO | NO | YES | YES | NO
EQPRepairJ22106934 | T6373-33 | CWC7J0A87-01 | OverKill | JB1TAS14 | PA5208KF2 | J6875 | 14 | open/short-A6858 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/11 14:37:05 | 2022/08/11 14:53:42 | 0.28 | OK | 20075 | 重新對針,range 10,mark 164,3 | NO | NO | NO | YES | YES | NO
EQPRepairJ22106937 | T6373-86 | CWC7J0C21-01 | OverKill | JB1Q4S19 | PA6542C | J68S1 | 25 | TOO MANY O/S-A6131 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/11 14:43:53 | 2022/08/11 17:19:02 | 2.59 | OK | 15362 | <低溫換卡>等待20分鐘降溫,更換配件#23,Y RANGE:12,P/M(12,4) ¦ | NO | NO | NO | YES | YES | NO
EQPRepairJ22106944 | T6373-07 | CWC7J0H15-01.002 | Prober Down | JB1TAS20 | PA5208K-LT | QT7WS | 09 | cannot measure index size-A6858 | CPAUTO | Prober | Wafer Edge/Mode Not Found | 2022/08/11 15:18:46 | 2022/08/11 15:29:52 | 0.19 | OK | 10210 | 重抓MODE(45.0) | NO | NO | NO | YES | YES | NO
EQPRepairJ22106948 | T6373-34 |  | Prober Down | JB1BYS10 | PA6587B_15U90H | J696A |  | AlarmCode:1060 AlarmMessage:CANNOT MEASURE INDEX SIZE !! | CPAUTO | Prober | Wafer Edge/Mode Not Found | 2022/08/11 15:33:04 | 2022/08/11 15:54:28 | 0.36 | OK | 13581 | 1.start機台重新抓\取OK | NO | NO | NO | NO | YES | NO
EQPRepairJ22106962 | T6373-39 | CWC7J0F95-01 | OverKill | JB1TAS21 | PA5208KF2 | J68G6 | 19 | TOO MANY O/S | A6873 | Over Kill | O/S誤宰 | 2022/08/11 16:44:57 | 2022/08/11 17:24:42 | 0.66 | OK | 07764 | OPEN 清針 | NO | NO | NO | YES | YES | NO
EQPRepairJ22106986 | T6373-12 | CWC8J0496-01 | OverKill | JB1PNS06 | NA01761D(COF94AA)-A-Y-C | GGR932.1-B1C111 | 05 | TOO MANY O/S. PROBEMARK IS TOO BIG. OD:5um | A6131 | Over Kill | O/S誤宰 | 2022/08/11 18:35:31 | 2022/08/11 18:58:14 | 0.38 | OK | 10210 | SOAKTIME Pass(170.0) | NO | NO | NO | YES | YES | NO
EQPRepairJ22106992 | T6373-25 | CWC8J0446-01.001 | OverKill | JB1LWS14 | ILISLQPEDQD-2CP | ILA14003.1 | 13 | TOO MANY O/S | A6513 | Over Kill | O/S誤宰 | 2022/08/11 18:57:23 | 2022/08/11 19:39:52 | 0.71 | OK | 17091 | 清針後確認針壓ok 二分法range x:-8~16 y:16~42 起測觀察一段穩定還機 | NO | NO | NO | YES | YES | NO
EQPRepairJ22106996 | T6373-147 | CWC7J0L15-01 | OverKill | TB1R8S82 | PA5137E1-LT | P097C.1 | 14 | o/s-07854 | CPAUTO | Over Kill | Show H | 2022/08/11 19:07:20 | 2022/08/11 19:22:37 | 0.25 | OK | 20277 | 輪工處理---- 1.產線誤掛O/S，2H 已更換過配件，請產線Run Corr驗證，續測觀察 | NO | NO | NO | YES | YES | NO
EQPRepairJ22106998 | T6373-38 |  | Corr Fail | JB1GKS10 | PA6543B | J6630 |  | Corr Fail | A6920 | Corr | CORR. FAIL | 2022/08/11 19:10:30 | 2022/08/11 19:25:10 | 0.24 | OK | 15492 | 1.比對 PASS | NO | NO | NO | YES | YES | NO
EQPRepairJ22107015 | T6373-136 | CWC8J0515-01 | OverKill | JB8Q6S05 | NA01761D(COF94AA)-A-Y-C | GGS331.1-BC | 01 | o/s pm not good wait ae check.-A7048 | CPAUTO | Over Kill | 針位調整 | 2022/08/11 20:12:29 | 2022/08/11 20:37:28 | 0.42 | OK | 20277 | Remote---- 1.Soaktime Range:X>20，Y:12 試打(218.5) 確認針位、針壓無異常，還機生產 | NO | NO | NO | YES | YES | NO
EQPRepairJ22107016 | T6373-15 | CWC8J0562-01 | OverKill | JB8Q6S04 | NA01761D(COF94AA)-A-Y-C | GGS330.1-BC1 | 10 | #09 EDA SHOW I (NO CORR WAFER)-A6512 | CPAUTO | Over Kill | Show I | 2022/08/11 20:19:50 | 2022/08/11 20:23:39 | 0.06 | OK | 20277 | 輪工處理---- 1.單片Show I [HW_P/DUTDIFF>4%]，前一片並無此問題 請產縣Run Corr驗證，續測觀察 | NO | NO | NO | YES | YES | NO
EQPRepairJ22107018 | T6373-27 | CWC8J0610-01 | OverKill | JB8Y9S02 | ILISLQPEDQD-2CP | ILA14006.1 | 06 | #6 Already follow the overkilling procedure but still OS | A6441 | Over Kill | O/S誤宰 | 2022/08/11 20:19:58 | 2022/08/11 20:39:00 | 0.32 | OK | 15492 | 1.針髒，機台清針 PASS | NO | NO | NO | YES | YES | NO
EQPRepairJ22107020 | T6373-151 | CWC8J0038-01 | OverKill | JB2P2S08 | R63455-A1S-EZ-CP | NAPY65.00-B1C | 03 | too many bin 23-A6516 | CPAUTO | Over Kill | Function誤宰 | 2022/08/11 20:21:39 | 2022/08/11 22:08:37 | 1.78 | OK | 20182 | 1.更換備卡#S10，上機一樣Bin23 2.查看EDA，有相同問題，續測 | NO | NO | NO | YES | YES | NO
EQPRepairJ22107021 | T6373-34 | CWC7J0G36-01.001 | OverKill | JB1BYS10 | PA6587B_15U90H | J696A | 07 | #7 Already follow the overkilling procedure but still OS | A6896 | Over Kill | O/S誤宰 | 2022/08/11 20:34:09 | 2022/08/11 21:39:15 | 1.09 | OK | 15492 | 1.下卡磨針，上機PASS，RANGE:8 | NO | NO | NO | YES | YES | NO
EQPRepairJ22107022 | T6373-07 |  | Prober Down | JB1TAS20 | PA5208KF-LT | J682N |  | AlarmCode:1060 AlarmMessage:CANNOT MEASURE INDEX SIZE !! | CPAUTO | Prober | Wafer Edge/Mode Not Found | 2022/08/11 20:35:02 | 2022/08/11 20:49:44 | 0.24 | OK | 09108 | Remote處理 ----- 機台再次自動重抓OK (147.1) | NO | NO | NO | YES | YES | NO
EQPRepairJ22107023 | T6373-147 |  | Prober Down | TB1R8S82 | PA5137E1-LT | P097C.1 |  | AlarmCode:1057 AlarmMessage:MODEL NOT FOUND ( ROUGH ALIGNMENT ) !! | CPAUTO | Prober | Wafer Edge/Mode Not Found | 2022/08/11 20:38:29 | 2022/08/11 20:45:37 | 0.12 | OK | 20277 | Remote---- 1.抓水平OK | NO | NO | NO | YES | YES | NO
EQPRepairJ22107032 | T6373-42 | CWC8J0291-01.003 | OverKill | JB1HKS06 | NA01687F(COG48BA)-A-AT-C | GGU014.1-BC12 | 16 | EDA SHOW 2H..NO CORR WAFER-A6516 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/11 21:07:16 | 2022/08/11 22:01:30 | 0.90 | OK | 20277 | Remote---- 1.下卡磨針 上機PASS Range:16 | NO | NO | NO | YES | YES | NO
EQPRepairJ22107036 | T6373-39 | CWC7J0F95-01 | OverKill | JB1TAS21 | PA5208KF2 | J68G6 | 20 | too many o/s..no particle-A6516 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/11 21:16:34 | 2022/08/11 21:38:36 | 0.37 | OK | 20277 | Remote---- 1.Range:X>20，Y:14 機台清針 回收PASS穩定，還機生產 | NO | NO | NO | YES | YES | NO
EQPRepairJ22107040 | T6373-128 |  | Setup Fail | VB1BVS01 | PA5270D-LT | U38F8.00 |  | 21:13 system error. | A7048 | Prober | 溫度異常 (Temperature Error) | 2022/08/11 21:25:26 | 2022/08/11 22:53:53 | 1.47 | OK | 20277 | Prober Down---- 1.Setup Fail溫控異常，重開後卡在第一段進不去 2.重開多次相同，改掛Prober Down，待處理 | NO | NO | NO | YES | YES | NO
EQPRepairJ22107043 | T6373-136 | CWC8J0515-01 | OverKill | JB8Q6S05 | NA01761D(COF94AA)-A-Y-C | GGS331.1-BC | 02 | #01 EDA SHOW I (NO CORR WAFER)-A6512 | CPAUTO | Over Kill | Show I | 2022/08/11 21:37:36 | 2022/08/11 21:41:09 | 0.06 | OK | 20277 | 輪工處理---- 1.單片Show I [HW_P/DUTDIFF>4%]，前一次維修已在回收 8Site無法多次回收造成，續測觀察 | NO | NO | NO | YES | YES | NO
EQPRepairJ22107044 | T6373-70 |  | Setup Fail | JB1Q4S16 | PA6542C | J68S1 |  | soaktime correct Z up limit over-A7048 | CPAUTO | Setup Fail | Prober異常 prober abnormal | 2022/08/11 21:39:19 | 2022/08/11 22:15:28 | 0.60 | OK | 20277 | Remote---- 1.重新對針 重抓Model OK，Change setup again | NO | NO | NO | YES | YES | NO
EQPRepairJ22107050 | T6373-21 | CMC8J0010-01 | Prober Down | JB1GKS08 | PA6543B | J6630 | 01 | F AXIS SOFTWARE ERROR (DRIVE RANGE OVER) | A6440 | Prober | SOAKTIME ERROR | 2022/08/11 21:58:09 | 2022/08/11 22:06:18 | 0.14 | OK | 15492 | REMOTE--- 1.RESOAK TIME OK | NO | NO | NO | YES | YES | NO
EQPRepairJ22107052 | T6373-27 | CWC8J0610-01 | OverKill | JB8Y9S02 | ILISLQPEDQD-2CP | ILA14006.1 | 06 | #06 too many o/s (8site) | A6441 | Over Kill | 針位調整 | 2022/08/11 22:11:07 | 2022/08/11 22:25:34 | 0.24 | OK | 17091 | 針位偏移 重新二分法調針 r:-2~18 起測回收os ok | NO | NO | NO | YES | YES | NO
EQPRepairJ22107055 | T6373-07 | CWC7J0H17-01 | Prober Down | JB1TAS20 | PA5208KF-LT | J682N | 05 | AlarmCode:1503 AlarmMessage:Coolant decrease. | CPAUTO | Prober | 溫度異常 (Temperature Error) | 2022/08/11 22:37:20 | 2022/08/12 00:41:21 | 2.07 | OK | 17184 | 1.Prober&Chiller ShutDown 2.ReS/U(46.0) | NO | NO | NO | NO | YES | NO
EQPRepairJ22107058 | T6373-128 |  | Prober Down | VB1BVS01 | PA5270D-LT | U38F8.00 |  | Prober Down---- 1.Setup Fail溫控異常，重開後卡在第一段進不去 2.重開多次相同，改掛Prober Down，待處理 | 20277 | Prober | 硬碟 | 2022/08/11 22:54:19 | 2022/08/12 14:44:50 | 15.84 | OK | 19771 | 常日協助更換硬碟後OK 改掛setup | NO | NO | NO | YES | YES | NO
EQPRepairJ22107064 | T6373-21 |  | Prober Down | JB1GKS08 | PA6543B | J6630 |  | AlarmCode:4526 AlarmMessage:LoadPort1 safety error! | CPAUTO | Prober | Loader Down | 2022/08/11 23:15:33 | 2022/08/11 23:49:03 | 0.56 | OK | 15492 | 1.退片不完全，推回 WAFER ，INITIAL OK | NO | NO | NO | YES | YES | NO
EQPRepairJ22107067 | T6373-11 | CWC8J0496-01.004 | OverKill | JB1PNS05 | NA01761D(COF94AA)-A-Y-C | GGR932.1-B1C112 | 19 | too many o/s pm up/down od10-A6943 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/11 23:24:47 | 2022/08/11 23:44:56 | 0.34 | OK | 17184 | 1.調針(198.7) | NO | NO | NO | NO | YES | NO
EQPRepairJ22107068 | T6373-136 | CWC8J0515-01 | Corr Fail | JB8Q6S05 | NA01761D(COF94AA)-A-Y-C | GGS331.1-BC | 03 | CORR FAIL-03491 | CPAUTO | Corr | CORR FAIL | 2022/08/11 23:26:44 | 2022/08/11 23:57:11 | 0.51 | OK | 20182 | 手動比對PASS | NO | NO | NO | YES | YES | NO
EQPRepairJ22107087 | T6373-38 | CWC6J0815-01.001 | OverKill | JB1GKS10 | PA6543B | J6630 | 23 | ALREADY CORR PASSED BUT STILL CONTINUOUS 2H | A6944 | Over Kill | 2h | 2022/08/12 00:40:47 | 2022/08/12 00:45:37 | 0.08 | OK | 15492 | 1.[HW_BIN0<94%](Bin 1 < 1064EA) 953(Hold Wafer) 2.續測觀察 | NO | NO | NO | YES | YES | NO
EQPRepairJ22107093 | T6373-07 |  | Prober Down | JB1TAS20 | PA5208KF-LT | J682N |  | CANNOT MEASURE INDEX SIZE | A6898 | Other | Other | 2022/08/12 01:06:13 | 2022/08/12 01:14:15 | 0.13 | OK | 20182 | 機台重抓OK | NO | NO | NO | NO | YES | NO
EQPRepairJ22107109 | T6373-21 | CMC8J0010-01 | OverKill | JB1GKS08 | PA6543B | J6630 | 03 | #3 Wafer 1.2 eda show H (No Corr./Test time exceeded) | A6896 | Over Kill | 2h | 2022/08/12 02:19:22 | 2022/08/12 02:35:59 | 0.28 | OK | 17091 | 2h [HW_BIN17>1%](Bin 18 > 11EA) 19(Hold Wafer)，[HW_BIN21>1%](Bin 22 > 11EA) 21(Hold Wafer)[HW_BIN60<95.5%](Bin 61 < 1081EA) 1056(Hold Wafer)，工程批非os造成，請產線換批驗證 | NO | NO | NO | YES | YES | NO
EQPRepairJ22107116 | T6373-70 | CWC8J0216-01 | OverKill | JB1Q4S17 | PA6542C | J683M | 02 | TOO MANY OS-A6512 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/12 02:42:55 | 2022/08/12 03:05:53 | 0.38 | OK | 20277 | Remote---- 1.Range:X>20，Y:14 機台清針 PASS | NO | NO | NO | YES | YES | NO
EQPRepairJ22107129 | T6373-29 | CWC8J0291-01.002 | OverKill | JB1QNS05 | NA01687F(COG48BA)-A-AT-C | GGU014.1-BC221 | 05 | #5 Already follow the overkilling procedure but still OS | A6896 | Over Kill | O/S誤宰 | 2022/08/12 03:35:11 | 2022/08/12 04:01:49 | 0.44 | OK | 17091 | 清針後確認針ok 二分法r:-10~8 點測穩定 | NO | NO | NO | YES | YES | NO
EQPRepairJ22107146 | T6373-21 |  | Prober Down | JB1GKS08 | PA6543B | J6630 |  | AlarmCode:4526 AlarmMessage:LoadPort1 safety error! | CPAUTO | Prober | Loader Down | 2022/08/12 04:45:36 | 2022/08/12 05:16:37 | 0.52 | OK | 15492 | 1.退片不完全，INITIAL OK | NO | NO | NO | YES | YES | NO
EQPRepairJ22107149 | T6373-43 | CWC7J0L43-01 | Setup Fail | VB1BVS02 | PA5270D-LT | U38F5.00 | 14 | 03:51 #14 id ok, setup failed due to function fail bin 34 no other available p/c polish at 6th floor | A6053 | Setup Fail | Setup Fail | 2022/08/12 04:47:52 | 2022/08/12 06:45:01 | 1.95 | OK | 20182 | 1.重新Setup ok | NO | NO | NO | NO | YES | NO
EQPRepairJ22107150 | T6373-86 | CWC7J0J58-01.003 | OverKill | JB1Q4S23 | PA6542C | J683L | 20 | too many open/short | A6599 | Over Kill | O/S誤宰 | 2022/08/12 04:58:29 | 2022/08/12 05:30:17 | 0.53 | OK | 17184 | 1.清針&調針(132.0) | NO | NO | NO | NO | YES | NO
EQPRepairJ22107171 | T6373-70 | CWC8J0216-01 | OverKill | JB1Q4S17 | PA6542C | J683M | 04 | TOO MANY OS-A6512 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/12 06:17:58 | 2022/08/12 06:44:38 | 0.44 | OK | 17091 | REM OTE----機台清針二分法調針PASS 測試一段穩定還機 | NO | NO | NO | YES | YES | NO
EQPRepairJ22107173 | T6373-12 |  | Tester Down | JB1PNS06 | NA01761D(COF94AA)-A-Y-C | GGR932.1-B1C111 |  | eng fail | 10515 | Tester | Linearity Init / Cal Fail | 2022/08/12 06:41:03 | 2022/08/12 09:41:25 | 3.01 | OK | 13535 | 夜班交接:init fail 兩次, 處理後init pass, run cal , diag ,交接日班 日班: diag pass 還機 | NO | NO | NO | NO | YES | NO
EQPRepairJ22107177 | T6373-15 |  | Prober Down | JB8Q6S04 | NA01761D(COF94AA)-A-Y-C | GGS330.1-BC1 |  | AlarmCode:4526 AlarmMessage:LoadPort1 safety error! | CPAUTO | Prober | Loader Down | 2022/08/12 06:57:44 | 2022/08/12 07:02:19 | 0.08 | OK | 20182 | LOADER INIT OK | NO | NO | NO | YES | YES | NO
EQPRepairJ22107187 | T6373-21 |  | Prober Down | JB1GKS08 | PA6543B | J6630 |  | AlarmCode:4526 AlarmMessage:LoadPort1 safety error! | CPAUTO | Prober | Loader Down | 2022/08/12 07:39:38 | 2022/08/12 08:03:34 | 0.40 | OK | 16631 | 1.#04退片凸片，手動推回init ok | NO | NO | NO | YES | YES | NO
EQPRepairJ22107190 | T6373-42 | CWC8J0391-01 | OverKill | JB1HKS06 | NA01687G(COG48BA)-A-AT-N2-C | GGS419.1-B11C11 | 06 | too many failbin 29, pm check up/down (145,4)-A6986 | CPAUTO | Over Kill | Function誤宰 | 2022/08/12 07:50:01 | 2022/08/12 09:21:07 | 1.52 | OK | 19771 | #06起測全面Bin 29 重load PGN OK | NO | NO | NO | YES | YES | NO
EQPRepairJ22107191 | T6373-43 |  | Setup Fail | VB1BVS02 | PA5270D-LT | U38F5.00 |  | setup04.contact BIN34-A6668 | CPAUTO | Setup Fail | Function Fail | 2022/08/12 07:50:46 | 2022/08/12 09:13:12 | 1.37 | OK | 19771 | 1.RUN CORR 全面bin34 重開測試機無效 2.load #25 點測pass 3.再次拉回#14點測OK | NO | NO | NO | YES | YES | NO
EQPRepairJ22107207 | T6373-29 | CWC8J0291-01.002 | OverKill | JB1QNS05 | NA01687F(COG48BA)-A-AT-C | GGU014.1-BC221 | 11 | fail bin 13-20022 | CPAUTO | Over Kill | Function誤宰 | 2022/08/12 08:50:59 | 2022/08/12 09:11:14 | 0.34 | OK | 11385 | 重新對針調整針位，點測pass穩定 | NO | NO | NO | YES | YES | NO
EQPRepairJ22107222 | T6373-70 | CWC8J0216-01 | OverKill | JB1Q4S17 | PA6542C | J683M | 06 | O/S-A6921 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/12 09:57:02 | 2022/08/12 10:47:07 | 0.83 | OK | 08774 | Remote 處理---- 1.執行soaktime 重新對針 機台清針 43.5 | NO | NO | NO | YES | YES | NO
EQPRepairJ22107225 | T6373-42 | CWC8J0391-01 | Tester Down | JB1HKS06 | NA01687G(COG48BA)-A-AT-N2-C | GGS419.1-B11C11 | 06 | RUN TIME ERROR-A6986 | CPAUTO | Product Issue | RTE | 2022/08/12 10:01:45 | 2022/08/12 10:18:24 | 0.28 | OK | 08774 | 無異常訊息，STDF 退片重新上貨 | NO | NO | NO | YES | YES | NO
EQPRepairJ22107227 | T6373-21 | CMC8J0010-01 | Prober Down | JB1GKS08 | PA6543B | J6630 | 05 | AlarmCode:4526 AlarmMessage:LoadPort1 safety error! | CPAUTO | Prober | Loader Down | 2022/08/12 10:05:40 | 2022/08/12 10:34:27 | 0.48 | OK | 16631 | 1.init ok，請產線更換cassette再觀察 | NO | NO | NO | YES | YES | NO
EQPRepairJ22107228 | T6373-121 | CWC8J0134-01.001 | Probe Card Down | JB1QNS01 | NA01687G(COG48BA)-A-AT-N2-C | GGS267.1-B11C2211 | 02 | Probe Card PM-05671 | CPAUTO | Probe Card | 定期PM | 2022/08/12 10:07:37 | 2022/08/12 12:33:15 | 2.43 | OK | 17712 | 低溫換卡，等待20分鐘，S01 PM，更換備卡S04上機PASS。 | NO | NO | NO | YES | YES | NO
EQPRepairJ22107230 | T6373-11 | CWC8J0496-01.004 | OverKill | JB1PNS05 | NA01761D(COF94AA)-A-Y-C | GGR932.1-B1C112 | 23 | too many o/s-A6894 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/12 10:09:52 | 2022/08/12 10:41:58 | 0.54 | OK | 17680 | remote----換片o/s,OD+2 & 溫針點測穩定pass | NO | NO | NO | YES | YES | NO
EQPRepairJ22107233 | T6373-147 |  | Prober Down | TB1R8S82 | PA5137E1-LT | P097C.1 |  | AlarmCode:1057 AlarmMessage:MODEL NOT FOUND ( ROUGH ALIGNMENT ) !! | CPAUTO | Prober | Wafer Edge/Mode Not Found | 2022/08/12 10:15:05 | 2022/08/12 10:34:21 | 0.32 | OK | 08774 | Remote 處理----重抓 45.0 | NO | NO | NO | YES | YES | NO
EQPRepairJ22107241 | T6373-42 | CWC8J0391-01 | OverKill | JB1HKS06 | NA01687G(COG48BA)-A-AT-N2-C | GGS419.1-B11C11 | 06 | STILL TOO MANY FAILBIN 29, PM CHECK (144,4) UP/DOWN-A6986 | CPAUTO | Over Kill | Function誤宰 | 2022/08/12 10:51:12 | 2022/08/12 11:43:06 | 0.87 | OK | 19771 | 重開測試機OK | NO | NO | NO | YES | YES | NO
EQPRepairJ22107245 | T6373-39 | CWC7J0F27-01 | OverKill | JB1TAS21 | PA5208KF2 | J6874 | 01 | P/M NG,.too many o/s-A6671 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/12 11:04:41 | 2022/08/12 12:30:41 | 1.43 | OK | 19771 | 更換備卡ok #1(147.5) | NO | NO | NO | YES | YES | NO
EQPRepairJ22107256 | T6373-42 | CWC8J0391-01 | Tester Down | JB1HKS06 | NA01687G(COG48BA)-A-AT-N2-C | GGS419.1-B11C11 | 06 | AlarmCode:665 AlarmMessage:GP-IB STOP COMMAND RECEIVED !! | CPAUTO | Other | Other | 2022/08/12 12:06:12 | 2022/08/12 12:07:29 | 0.02 | OK | 19771 | TA 誤掛 | NO | NO | NO | YES | YES | NO
EQPRepairJ22107265 | T6373-103 | CWC8J0612-01 | OverKill | JB6K8S02 | NC01682A(COG48AA)-A-AT-N3-C | GGS288.1-BC | 01 | too many o/s-A6052 | CPAUTO | Over Kill | O/S | 2022/08/12 12:35:24 | 2022/08/12 13:07:38 | 0.54 | OK | 11385 | 溫針重新對針調整針位 (167.5) | NO | NO | NO | YES | YES | NO
EQPRepairJ22107270 | T6373-151 | CWC8J0038-01 | OverKill | JB2P2S08 | R63455-A1S-EZ-CP | NAPY65.00-B1C | 05 | BIN 23-A6916 | CPAUTO | Over Kill | Function誤宰 | 2022/08/12 12:50:54 | 2022/08/12 13:26:59 | 0.60 | OK | 16631 | 1.#05 大量BIN23，先前已有換卡無效，通知產品負責人是否關閉ALARM，告知先換機測試 | NO | NO | NO | YES | YES | NO
EQPRepairJ22107274 | T6373-42 | CWC8J0391-01 | OverKill | JB1HKS06 | NA01687G(COG48BA)-A-AT-N2-C | GGS419.1-B11C11 | 06 | BIN 29-A6916 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/12 13:02:12 | 2022/08/12 14:06:32 | 1.07 | OK | 08774 | BIN29 下卡磨針 | NO | NO | NO | YES | YES | NO
EQPRepairJ22107276 | T6373-136 | CWC8J0515-01 | OverKill | JB8Q6S05 | NA01761D(COF94AA)-A-Y-C | GGS331.1-BC | 09 | # EDA SHOW I, ALREADY CORR PASSED-A6986 | CPAUTO | Over Kill | show I | 2022/08/12 13:06:56 | 2022/08/12 13:09:59 | 0.05 | OK | 17680 | single I ,keep test | NO | NO | NO | YES | YES | NO
EQPRepairJ22107280 | T6373-15 | CWC8J0562-01 | OverKill | JB8Q6S04 | NA01761D(COF94AA)-A-Y-C | GGS330.1-BC11 | 19 | Continuous 2 wafers EDA shows "I" corr passed | A6921 | Over Kill | 2I | 2022/08/12 13:21:10 | 2022/08/12 15:14:32 | 1.89 | OK | 19771 | 2I 配件磨針 #19(49.8) | NO | NO | NO | YES | YES | NO
EQPRepairJ22107286 | T6373-43 | CWC7J0L43-01 | Corr Fail | VB1BVS02 | PA5270D-LT | U38F5.00 | 15 | t32 Corr Fail-A6986 | CPAUTO | Corr | Other | 2022/08/12 13:45:52 | 2022/08/12 13:51:26 | 0.09 | OK | 08774 | T32先RUN CORR 後比對 1.查看S.T RUN CORR FAIL較多，量產YIELD較高，確認針壓針位 續測觀察 | NO | NO | NO | YES | YES | NO
EQPRepairJ22107293 | T6373-70 | CWC8J0216-01 | OverKill | JB1Q4S17 | PA6542C | J683M | 08 | too many o/s, pm check up/down (127,5)-A6986 | CPAUTO | Over Kill | O/S | 2022/08/12 14:04:10 | 2022/08/12 14:37:46 | 0.56 | OK | 11385 | 溫針重新對針調整針位清針 | NO | NO | NO | YES | YES | NO
EQPRepairJ22107294 | T6373-90 | CRC7J0011-01 | Prober Down | JB1GVS15 | NC01682A(COG48AA)-A-AT-N3-C | GCVP99.1-BC111121C | 01 | UNIT2 HEIGHT OUT OF TOLERANT AREA ..-A6916 | CPAUTO | Prober | Other | 2022/08/12 14:07:14 | 2022/08/12 14:23:21 | 0.27 | OK | 16631 | 1.電話詢問常日啟家，告知先關閉Touch Sensor待處理 | NO | NO | NO | YES | YES | NO
EQPRepairJ22107308 | T6373-11 | CWC8J0496-01.004 | OverKill | JB1PNS05 | NA01761D(COF94AA)-A-Y-C | GGR932.1-B1C112 | 24 | #24 too many o/s-19617 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/12 14:43:50 | 2022/08/12 15:06:39 | 0.38 | OK | 08774 | Remote 處理---- 1.執行soaktime 重新對針 | NO | NO | NO | YES | YES | NO
EQPRepairJ22107314 | T6373-34 | CWC8J0168-01 | OverKill | JB1BYS10 | PA6587B_15U90H | J6969 | 02 | o/s-20022 | CPAUTO | Over Kill | 機台清針後Pass | 2022/08/12 15:17:29 | 2022/08/12 15:31:33 | 0.23 | OK | 16631 | 1.機台清針後Pass | NO | NO | NO | YES | YES | NO
EQPRepairJ22107319 | T6373-39 | CWC7J0F27-01 | Other | JB1TAS26 | PA5208KF2 | J6874 | 02 | AE CHANGE PC-A6516 | CPAUTO | Probe Card | 配件Open更換P/C | 2022/08/12 15:52:33 | 2022/08/12 16:24:08 | 0.53 | OK | 08774 | 配件室要求換卡 46.6 | NO | NO | NO | YES | YES | NO
EQPRepairJ22107320 | T6373-90 | CRC7J0011-01 | OverKill | JB1GVS15 | NC01682A(COG48AA)-A-AT-N3-C | GCVP99.1-BC111121C | 01 | O/S TOO MANY-A6916 | CPAUTO | Over Kill | O/D不足 | 2022/08/12 15:55:13 | 2022/08/12 16:08:42 | 0.22 | OK | 16631 | 1.重新對針 181.4 | NO | NO | NO | YES | YES | NO
EQPRepairJ22107323 | T6373-21 | CMC8J0010-01 | OverKill | JB1GKS08 | PA6543B | J6630 | 08 | O/S TOO MANY-A6916 | CPAUTO | Over Kill | O/S | 2022/08/12 16:00:50 | 2022/08/12 16:39:15 | 0.64 | OK | 11385 | 溫針重新對針調整針位 (130.0) | NO | NO | NO | YES | YES | NO
EQPRepairJ22107328 | T6373-152 | CWC8J0166-01 | Setup Fail | JB2P2S09 | R63455-A1S-EZ-CP | NAHM66.00-B1 | 16 | Setup Fail 有針點沒contact 輪工確認中 | 14212 | Setup Fail | O/S Fail | 2022/08/12 16:10:17 | 2022/08/12 16:26:44 | 0.27 | OK | 08774 | 有針點沒contact 確認tester UTILITU PS燈號不亮 1.此台為AUTO CARD CHANGE機台 2.退片執行AUTO CARD CHANGE，抬頭 3.蓋頭，CARD CHANGE後燈號正常，續掛S/U | NO | NO | NO | YES | YES | NO
EQPRepairJ22107329 | T6373-43 | CWC7J0L43-01 | OverKill | VB1BVS02 | PA5270D-LT | U38F5.00 | 15 | #15 bin34-13081 | CPAUTO | Over Kill | Function誤宰 | 2022/08/12 16:21:19 | 2022/08/12 17:17:55 | 0.94 | OK | 19771 | 調整後可回收,換片起測一段無異常,還機續測觀察 | NO | NO | NO | YES | YES | NO
EQPRepairJ22107333 | T6373-111 | CWC7J0761-01.004 | OverKill | TB1R8S74 | PA5137F-LT1 | P07HJ.1 | 13 | EDA SHOW 2H(NO CORR WAFER)-17436 | CPAUTO | Product Issue | 2H | 2022/08/12 16:40:43 | 2022/08/12 16:52:02 | 0.19 | OK | 13749 | Remote--- 1.eda Slot #01-#02 show 2H [HW_BIN0<95.5%](Bin 1 < 814EA) 808(Hold Wafer) . 2.原因為前流程Fail Bin造成，非當下CP4流程測試產生 Product Issue續測. | NO | NO | NO | YES | YES | NO
EQPRepairJ22107336 | T6373-12 | CWC8J0496-01 | OverKill | JB1PNS06 | NA01761D(COF94AA)-A-Y-C | GGR932.1-B1C111 | 12 | O/S-17436 | CPAUTO | Over Kill | 針位調整 | 2022/08/12 16:51:55 | 2022/08/12 17:09:30 | 0.29 | OK | 13749 | 1.Slot #12 測試中大量O/S，且針痕過大 貼下擠壓BUMP Check NG PM(65.8). 2.對針前 測試OD:5um Contact 高度:13679，對針完成後 2分法PASS 測試OD:5um 高度為:13666 低於原先13um 即可pass. 3.2分法 Range=21um PM(197.1) 針位X=0 Y=4 | NO | NO | NO | YES | YES | NO
EQPRepairJ22107344 | T6373-29 | CWC8J0291-01.002 | OverKill | JB1QNS05 | NA01687F(COG48BA)-A-AT-C | GGU014.1-BC221 | 22 | BIN 13-A6916 | CPAUTO | Probe Card | BIN13 | 2022/08/12 17:15:23 | 2022/08/12 17:53:16 | 0.63 | OK | 11385 | P/C down----#05配件測試Bin13 Function過多，清針調整針位無效，配件送修 | NO | NO | NO | YES | YES | NO
EQPRepairJ22107352 | T6373-29 |  | Probe Card Down | JB1QNS05 | NA01687F(COG48BA)-A-AT-C | GGU014.1-BC221 |  | P/C down----#05配件測試Bin13 Function過多，清針調整針位無效，配件送修 | 11385 | Probe Card | bin13 | 2022/08/12 17:53:48 | 2022/08/12 19:58:49 | 2.08 | OK | 13597 | P/C down----#05配件測試Bin13 Function過多，清針調整針位無效，配件送修 1.日班交接更換配件#01，上機bin13不穩 2.更換配件#06，降溫交接夜班 3.接手上機pass，range 8 um 4.點測bin13 pass，續測觀察 | NO | NO | NO | YES | YES | NO
EQPRepairJ22107357 | T6373-103 | CWC8J0612-01 | OverKill | JB6K8S02 | NC01682A(COG48AA)-A-AT-N3-C | GGS288.1-BC | 07 | too many o/s-A6052 | CPAUTO | Over Kill | 機台清針後Pass | 2022/08/12 18:03:55 | 2022/08/12 18:28:49 | 0.42 | OK | 16631 | 1.機台CLEAN WAFER清針後Pass | NO | NO | NO | YES | YES | NO
EQPRepairJ22107367 | T6373-47 | CWC8J0348-01.004 | Corr Fail | JB1HKS09 | NA01687G(COG48BA)-A-AT-N2-C | GGR761.1-BC2 | 08 | 第二次良率比較好 | 08341 | Corr | Other | 2022/08/12 19:01:38 | 2022/08/12 19:02:57 | 0.02 | OK | 08774 | Remote 處理----Corr Fail---- Corr 良率比 Corr Wafer 好，前流程 Fail Bin 回收成 Pass，手動比對 Corr Pass | NO | NO | NO | YES | YES | NO
EQPRepairJ22107375 | T6373-34 | CWC8J0168-01 | OverKill | JB1BYS10 | PA6587B_15U90H | J6969 | 04 | #2,3 EDA SHOW 2 H #2 DO NOT NEED TO RUNCORR-A6497 | CPAUTO | Over Kill | Show H | 2022/08/12 19:45:33 | 2022/08/12 20:02:43 | 0.29 | OK | 17542 | Show H，#02 [HW_BIN60<95.5%](Bin 61 < 1090EA) 1084(Hold Wafer(Yield>90% donot neet to run corr,and keep test))，#03 [HW_BIN20>0.3%](Bin 21 > 3EA) 5，不同Fail Bin造成Show H，請產線先行續測待後續觀察 | NO | NO | NO | YES | YES | NO
EQPRepairJ22107386 | T6373-111 |  | Corr Fail | TB1R8S74 | PA5137F-LT1 | P07HJ.1 |  | Corr Fail yield不足 | 19697 | Corr | 產品問題 | 2022/08/12 20:10:16 | 2022/08/12 20:29:41 | 0.32 | OK | 11348 | CORR FAIL----邊緣BIN127 續測 | YES | NO | NO | YES | YES | NO
EQPRepairJ22107387 | T6373-43 | CWC7J0L43-01 | OverKill | VB1BVS02 | PA5270D-LT | U38F5.00 | 17 | BIN34 TOO MANY OD15-A6788 | CPAUTO | Over Kill | Function誤宰 | 2022/08/12 20:12:16 | 2022/08/12 21:54:30 | 1.70 | OK | 13510 | 1.更換配件#01後上機一樣bin34 2.重load程式後ok | NO | NO | NO | YES | YES | NO
EQPRepairJ22107392 | T6373-47 | CWC8J0348-01 | OverKill | JB1HKS09 | NA01687G(COG48BA)-A-AT-N2-C | GGR356.1-B11C111111 | 13 | too many open short-A6786 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/12 20:29:36 | 2022/08/12 20:54:13 | 0.41 | OK | 13510 | 1.清針後重新對針pass | NO | NO | NO | YES | YES | NO
EQPRepairJ22107403 | T6373-141 | CWC8J0291-01.003 | OverKill | JB1HKS10 | NA01687F(COG48BA)-A-AT-C | GGU014.1-BC121 | 02 | FAILBIN TO O/S | A6404 | Over Kill | O/S誤宰 | 2022/08/12 21:16:37 | 2022/08/12 23:38:23 | 2.36 | OK | 17542 | 下卡磨針，上機點測發生online init，溫針重對針，range:12，試打(146,0)無異常，點測PASS穩定 | NO | NO | YES | YES | YES | NO
EQPRepairJ22107419 | T6373-70 | CWC8J0216-01 | OverKill | JB1Q4S17 | PA6542C | J683M | 13 | #11,12 eda 1st map too many o/s,pls check by leader-A7033 | CPAUTO | Over Kill | O/D不足 | 2022/08/12 22:03:24 | 2022/08/12 23:21:04 | 1.29 | OK | 17674 | 1. OD 8->15 PASS (12,4)#13 | NO | NO | NO | YES | YES | NO
EQPRepairJ22107420 | T6373-135 | CWC8J0184-01 | OverKill | VB4BKS02 | PA5129LS_12U80H-LT | P07CL.1 | 10 | #3.4 EDA 2H-A6925 | CPAUTO | Over Kill | SHOW H | 2022/08/12 22:03:47 | 2022/08/12 22:21:41 | 0.30 | OK | 13597 | 後續測試SHOW N，請產線RUN CORR驗證 | NO | NO | NO | YES | YES | NO
EQPRepairJ22107421 | T6373-25 | CWC8J0610-01 | OverKill | JB1LWS14 | ILISLQPEDQD-2CP | ILA14006.1 | 02 | TOO MANY O/S | A6404 | Over Kill | O/D不足 | 2022/08/12 22:05:05 | 2022/08/12 22:24:24 | 0.32 | OK | 20276 | 增加OD 8um>10um range 13 P/M(125.-1) | NO | NO | NO | YES | YES | NO
EQPRepairJ22107425 | T6373-90 | CWC8J0119-01 | Corr Fail | JB1GVS15 | NC01682A(COG48AA)-A-AT-N3-C | GGP620.1-2B | 21 | Corr Fail | A5929 | Corr | Other | 2022/08/12 22:27:25 | 2022/08/12 22:47:43 | 0.34 | OK | 13597 | 比對PASS | NO | NO | NO | YES | YES | NO
EQPRepairJ22107439 | T6373-15 | CWC8J0562-01 | Prober Down | JB8Q6S04 | NA01761D(COF94AA)-A-Y-C | GGS330.1-BC11 | 23 | PROBER NG-A6925 | CPAUTO | Prober | Other | 2022/08/12 23:27:54 | 2022/08/13 02:08:16 | 2.67 | OK | 17674 | 1. 測試中Prober 面板當，R/C 與實際 wafer 不符，R/C 有備註 拆批 及hold TPD 片號，但#01片 無備註，請產線確認OK 2. 機台INIT 退片、檢查EDA 正常、C-LOT S/U PASS，上機機台預溫20、下針壓溫針退針壓、Data in後 機台又溫針 3. #23 (112,5)溫針 range y=15 x=20 (212,6)，重新上貨 | NO | NO | NO | YES | YES | NO
EQPRepairJ22107441 | T6373-90 | CRC7J0009-01 | OverKill | JB1GVS15 | NC01682A(COG48AA)-A-AT-N3-C | GFK688.1-BC12111C | 25 | too many failbin probe mark NG shadow on lower side-A6497 | CPAUTO | Over Kill | Function誤宰 | 2022/08/12 23:47:07 | 2022/08/13 00:27:13 | 0.67 | OK | 17542 | 重對針，range:12，試打(140,0)無異常，點測PASS穩定 | NO | NO | YES | YES | YES | NO
EQPRepairJ22107443 | T6373-38 | CWC8J0615-01 | OverKill | JB1GKS10 | PA6543B | J6AC7 | 02 | too many o/s already overkill-A6497 | CPAUTO | Over Kill | 配件卡毛 | 2022/08/12 23:54:37 | 2022/08/13 01:36:23 | 1.70 | OK | 20276 | 配件卡毛 下卡清針 range 10 P/M(131.7) | NO | NO | NO | YES | YES | NO
EQPRepairJ22107445 | T6373-01 | CWC7J0K82-01 | OverKill | JB1TAS27 | PA5208KF-LT | J682K | 02 | Too many o/s-A6053 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/13 00:18:55 | 2022/08/13 01:17:51 | 0.98 | OK | 13597 | 更換配件#23，RANGE 10 UM | NO | NO | NO | YES | YES | NO
EQPRepairJ22107452 | T6373-135 | CWC8J0184-01 | OverKill | VB4BKS02 | PA5129LS_12U80H-LT | P07CL.1 | 16 | EDA SHOW 2H (O/S) | A6296 | Over Kill | 2H | 2022/08/13 00:53:04 | 2022/08/13 01:06:02 | 0.22 | OK | 17674 | 1. 2H (Bin 73 > 2EA) 4 (hold wafer) ，bin73-short 良率99% | NO | NO | NO | YES | YES | NO
EQPRepairJ22107455 | T6373-70 | CWC8J0216-01 | OverKill | JB1Q4S17 | PA6542C | J683M | 14 | too many o/s-A7033 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/13 01:02:04 | 2022/08/13 02:46:46 | 1.75 | OK | 17542 | 1.測試中O/S不穩，配件嚴重黑點 2.更換在庫配件JB1Q4S16，上機對針ok，溫針重對針，點測固定PIN OPEN，清調針 & 重MNT & 更換POGO PIN無效，故將配件送修 3.原卡磨針，上機對針ok，溫針ok，Data-In發生Soaktime Error，重對針，點測PASS Data-In後執行Soaktime，再次重對針，range:8，試打(43,5)無異常，起測一段Pass穩定 | NO | NO | NO | YES | YES | NO
EQPRepairJ22107457 | T6373-16 | CWC7J0K22-01 | OverKill | JB1TAS28 | PA5208KF-LT | J67LH | 01 | too many open short-A6786 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/13 01:12:26 | 2022/08/13 01:59:14 | 0.78 | OK | 15171 | Remote處理 --- 重新對針，RANGE:8UM，(45.6) | NO | NO | NO | YES | YES | NO
EQPRepairJ22107458 | T6373-43 | CWC7J0L43-01 | OverKill | VB1BVS01 | PA5270D-LT | U38F5.00 | 18 | eda show 2"H" PASS BIN(NO CORR WAFER) | A6296 | Over Kill | SHOW H | 2022/08/13 01:14:49 | 2022/08/13 01:36:17 | 0.36 | OK | 15171 | Remote處理 --- 針壓針位OK續測 | NO | NO | NO | YES | YES | NO
EQPRepairJ22107467 | T6373-43 | CWC7J0L43-01 | OverKill | VB1BVS01 | PA5270D-LT | U38F5.00 | 18 | TOO MANY FAIL BIN-A6786 | CPAUTO | Over Kill | Function誤宰 | 2022/08/13 02:03:47 | 2022/08/13 02:16:44 | 0.22 | OK | 13510 | 1.重load程式ok | NO | NO | NO | YES | YES | NO
EQPRepairJ22107474 | T6373-25 | CWC8J0610-01 | OverKill | JB1LWS14 | ILISLQPEDQD-2CP | ILA14006.1 | 04 | #2,#3 SHOW 2H(NO CORR)-19697 | CPAUTO | Over Kill | Show H | 2022/08/13 02:34:38 | 2022/08/13 02:50:34 | 0.27 | OK | 17542 | Show H，不同Fail Bin造成Show H，不同機卡有相同問題，續測 | NO | NO | NO | YES | YES | NO
EQPRepairJ22107478 | T6373-39 | CWC7J0F27-01 | OverKill | JB1TAS25 | PA5208KF2 | J6874 | 08 | CONTINOUS EDA SHOWS 2R DUE TO OS-A6788 | CPAUTO | Over Kill | SHOW R | 2022/08/13 02:58:44 | 2022/08/13 04:19:37 | 1.35 | OK | 13597 | 機台清針PASS，測試一段穩定，續測觀察 | NO | NO | NO | YES | YES | NO
EQPRepairJ22107483 | T6373-111 | CWC7J0761-01.004 | OverKill | TB1R8S74 | PA5137F-LT1 | P07HJ.1 | 21 | TOO MANY O/S(OD15um )overkill done apply. | A6296 | Over Kill | O/S | 2022/08/13 03:20:08 | 2022/08/13 04:02:22 | 0.70 | OK | 15096 | open 清、調針pass. | NO | NO | NO | YES | YES | NO
EQPRepairJ22107486 | T6373-43 | CWC7J0L43-01 | OverKill | VB1BVS01 | PA5270D-LT | U38F5.00 | 18 | (111.3)(79.3)p-f (73.5)f-f too many o/s and bin34 od15-A6788 | CPAUTO | Over Kill | Function誤宰 | 2022/08/13 03:33:26 | 2022/08/13 05:11:18 | 1.63 | OK | 17674 | 1. Bin34 今夜已有換過卡無效，有重load 程式pass紀錄，故將 Tester 重開起測穩定pass，還機觀察 | NO | NO | NO | YES | YES | NO
EQPRepairJ22107509 | T6373-15 | CWC8J0562-01 | OverKill | JB8Q6S04 | NA01761D(COF94AA)-A-Y-C | GGS330.1-BC11 | 25 | #23 EDA show i.. no corr wafer | A6925 | Over Kill | 2I | 2022/08/13 05:35:11 | 2022/08/13 05:44:45 | 0.16 | OK | 17674 | 1. 2I 下一片SHOW N | NO | NO | NO | YES | YES | NO
EQPRepairJ22107511 | T6373-32 | CWC8J0348-01.013 | Prober Down | JB1HKS07 | NA01687G(COG48BA)-A-AT-N2-C | GGR356.1-B11C1122 | 23 | LOADPORT1 SAFETY ERROR-A6497 | CPAUTO | Prober | Loader Down | 2022/08/13 05:47:38 | 2022/08/13 06:02:36 | 0.25 | OK | 13597 | init ok | NO | NO | NO | YES | YES | NO
EQPRepairJ22107512 | T6373-37 | CWC8J0012-01 | OverKill | TB2U7S09 | R69005A0FNCPJDQV-P | AYN442000-B1C | 02 | too many o/s-A6497 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/13 05:52:45 | 2022/08/13 06:13:53 | 0.35 | OK | 17542 | 執行saoktime，重對針，range:18，試打(213,5)無異常，起測一段pass穩定 | NO | NO | NO | YES | YES | NO
EQPRepairJ22107514 | T6373-43 | CWC7J0L43-01 | OverKill | VB1BVS01 | PA5270D-LT | U38F5.00 | 18 | TOO MANY FAIL BIN | A6786 | Over Kill | Function誤宰 | 2022/08/13 05:54:01 | 2022/08/13 06:15:47 | 0.36 | OK | 17674 | 1. Function誤宰 34 重LOAD PGM PASS， 待日班給產品確認 | NO | NO | NO | YES | YES | NO
EQPRepairJ22107521 | T6373-39 | CWC7J0F27-01 | OverKill | JB1TAS25 | PA5208KF2 | J6874 | 09 | EDA SHOW 2R | A6786 | Over Kill | 2R | 2022/08/13 06:09:30 | 2022/08/13 06:17:34 | 0.13 | OK | 17674 | 1. 2R 補掛 下一片SHOW N，機台清針還機觀察 | NO | NO | NO | YES | YES | NO
EQPRepairJ22107522 | T6373-70 | CWC8J0216-01 | Prober Down | JB1Q4S17 | PA6542C | J683M | 17 | AlarmCode:7214 AlarmMessage:Soaktime correct Z-UP limit over !! | CPAUTO | Prober | Other | 2022/08/13 06:16:01 | 2022/08/13 06:42:54 | 0.45 | OK | 17674 | 1. 機台自動抓正常 | NO | NO | NO | YES | YES | NO
EQPRepairJ22107527 | T6373-16 | CWC7J0K22-01 | OverKill | JB1TAS28 | PA5208KF-LT | J67LH | 03 | too many open short-A6786 | CPAUTO | Over Kill | 機台清針後Pass | 2022/08/13 06:35:47 | 2022/08/13 06:55:19 | 0.33 | OK | 17674 | 機台清針後Pass | NO | NO | NO | YES | YES | NO
EQPRepairJ22107530 | T6373-103 | CWC8J0612-01 | OverKill | JB6K8S02 | NC01682A(COG48AA)-A-AT-N3-C | GGS288.1-BC | 22 | os too many 6sites-A6899 | CPAUTO | Over Kill | 配件殘金 | 2022/08/13 06:43:34 | 2022/08/13 08:19:21 | 1.60 | OK | 11385 | 下卡磨針 (168.5) | NO | NO | NO | YES | YES | NO
EQPRepairJ22107536 | T6373-37 | CWC8J0012-01 | OverKill | TB2U7S09 | R69005A0FNCPJDQV-P | AYN442000-B1C | 02 | TOO MANY O/S(RE-PROBING NO PASS)PROBEMARK TOO BIG ALREADY | A6296 | Over Kill | O/S | 2022/08/13 07:04:28 | 2022/08/13 08:02:37 | 0.97 | OK | 11385 | 重新對針調整針位清針 | NO | NO | NO | YES | YES | NO
EQPRepairJ22107544 | T6373-111 | CWC7J0562-01.002 | OverKill | TB1R8S74 | PA5137F-LT1 | P07HH.1 | 11 | #11 too many o/s-19617 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/13 07:52:59 | 2022/08/13 08:19:03 | 0.43 | OK | 17680 | remote----調整針位range:6um,position(-1,8),mark(117,6),OD-2um,起測穩定 | NO | NO | NO | YES | YES | NO
EQPRepairJ22107552 | T6373-43 | CWC7J0L43-01 | OverKill | VB1BVS01 | PA5270D-LT | U38F5.00 | 19 | #19 已依誤宰處理流程處置無效 (Bin34 過多) | 19617 | Over Kill | Function誤宰 | 2022/08/13 08:31:58 | 2022/08/13 09:03:10 | 0.52 | OK | 19771 | 清調針OK | NO | NO | NO | YES | YES | NO
EQPRepairJ22107557 | T6373-70 | CWC8J0216-01 | OverKill | JB1Q4S17 | PA6542C | J683M | 18 | always alarm, too many o/s-A6986 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/13 09:03:04 | 2022/08/13 09:36:47 | 0.56 | OK | 17680 | REMOTE----針壓不足重新對針RANGE:13UM,POSITION(0,1),MARK(131,0)，起測穩定 | NO | NO | NO | YES | YES | NO
EQPRepairJ22107559 | T6373-152 | CWC8J0448-01.004 | OverKill | JB252S01 | TD7850-A0S-AZ-CP | SC1GT-B111C | 02 | o/s | A6571 | Over Kill | 針位調整 | 2022/08/13 09:09:40 | 2022/08/13 09:28:19 | 0.31 | OK | 19771 | 針位調整ok #2(147.7) | NO | NO | NO | YES | YES | NO
EQPRepairJ22107579 | T6373-21 | CWC7J0L61-01 | OverKill | VB405S07 | PA5208K-LT | QT7CH | 17 | o/s-a6571 | CPAUTO | Over Kill | 針位調整 | 2022/08/13 10:21:12 | 2022/08/13 10:41:54 | 0.35 | OK | 16631 | 1.針位貼下，針位調整PASS，3D清針 | NO | NO | NO | YES | YES | NO
EQPRepairJ22107581 | T6373-27 | CWC8J0610-01.001 | OverKill | JB1LWS16 | ILISLQPEDQD-2CP | ILA14006.1 | 24 | CONTINOUS EDA 2 SHOW H ALREADY CORR PASSED-A6049 | CPAUTO | Over Kill | 2H | 2022/08/13 10:33:55 | 2022/08/13 10:41:03 | 0.12 | OK | 08819 | 1.2H，同批不同機卡有相同問題，Corr Passed續測。 | NO | NO | NO | YES | YES | NO
EQPRepairJ22107584 | T6373-29 | CWC8J0291-01.003 | OverKill | JB1QNS06 | NA01687F(COG48BA)-A-AT-C | GGU014.1-BC121 | 08 | #2.6 EDA SHOW 2H, NO CORR WAFER-A6986 | CPAUTO | Over Kill | 2H | 2022/08/13 10:58:12 | 2022/08/13 11:01:53 | 0.06 | OK | 16631 | 1.2H，請產線續測 | NO | NO | NO | YES | YES | NO
EQPRepairJ22107586 | T6373-152 | CWC8J0448-01.004 | OverKill | JB252S01 | TD7850-A0S-AZ-CP | SC1GT-B111C | 02 | TOO MANY O/S | A6513 | Over Kill | O/D不足 | 2022/08/13 11:13:17 | 2022/08/13 11:55:23 | 0.70 | OK | 19771 | 重新對針ok(53.7) | NO | NO | NO | YES | YES | NO
EQPRepairJ22107599 | T6373-70 |  | Prober Down | JB1Q4S17 | PA6542C | J683M |  | AlarmCode:7214 AlarmMessage:Soaktime correct Z-UP limit over !! | CPAUTO | Prober | soak time z limit | 2022/08/13 12:25:44 | 2022/08/13 12:54:37 | 0.48 | OK | 19771 | 重新對針ok #20(13.4) | NO | NO | NO | YES | YES | NO
EQPRepairJ22107602 | T6373-35 | CRC7J0024-01 | Setup Fail | JB1QNS05 | NA01687G(COG48BA)-A-AT-N2-C | GGN561.1-B1C12111 | 11 | setup Bin56 調針清針換卡#01→#05無效 | 05671 | Tester | bin56 | 2022/08/13 12:34:22 | 2022/08/13 14:55:00 | 2.34 | OK | 11385 | Tester down----1.改測NA01687G(COG48BA)-A-AT-N2-C / CP6 All Bin56 Function，更換配件一樣 2.Full call PASS上機一樣Bin56，待後續維修 | NO | NO | NO | YES | YES | NO
EQPRepairJ22107603 | T6373-33 | CWC7J0A88-01 | OverKill | JB1TAS14 | PA5208KF2 | J68R0 | 04 | Already follow Overkilll handling procedures but still not effective too many o/s | A6897 | Over Kill | O/S誤宰 | 2022/08/13 12:38:16 | 2022/08/13 13:02:39 | 0.41 | OK | 17680 | remote-----零星o/s，調整針位position(8,3)，mark(46,0),設定memo單3D clean參數(未設定)，起測回收穩定 | NO | NO | NO | YES | YES | NO
EQPRepairJ22107609 | T6373-42 | CWC8J0347-01.001 | OverKill | JB1HKS06 | NA01687F(COG48BA)-A-AT-C | GGU954.1-BC2 | 10 | bin 29-a6571 | CPAUTO | Over Kill | Function誤宰 | 2022/08/13 12:55:03 | 2022/08/13 13:51:25 | 0.94 | OK | 19771 | 配件磨針PASS #10(35.4) | NO | NO | NO | YES | YES | NO
EQPRepairJ22107610 | T6373-43 | CWC7J0L43-01 | OverKill | VB1BVS01 | PA5270D-LT | U38F5.00 | 20 | BIN 34 | A6921 | Over Kill | Function誤宰 | 2022/08/13 12:57:17 | 2022/08/13 13:37:04 | 0.66 | OK | 08774 | Remote 處理---確認後pass 續測 | NO | NO | NO | YES | YES | NO
EQPRepairJ22107614 | T6373-129 | CWC7J0197-01.008 | Prober Down | TB1INS29 | NA01404H-A-AT-C | GGR689.1-BC21112 | 05 | AlarmCode:4827 AlarmMessage:ARM @@VACUUM ERROR !! | CPAUTO | Prober | 手臂異常(Arm Error) | 2022/08/13 13:12:36 | 2022/08/13 13:25:38 | 0.22 | OK | 17680 | Prober Down--- 1.機台自動執行Clean Wafer 清針時，發生Arm1 Vacuum Error. 2.查看機台有開啟Delay Time 仍發生異常. 3.此為低溫型號測試，開機至第一段 先行升溫待確認. | NO | NO | NO | YES | YES | NO
EQPRepairJ22107616 | T6373-128 | CWC7J0M27-01.001 | OverKill | JB1BYS11 | PA6587B_15U90H | J696G | 25 | o/s-13081 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/13 13:21:32 | 2022/08/13 13:54:55 | 0.56 | OK | 08774 | Remote 處理---重新對針 機台清針 46.7 | NO | NO | NO | YES | YES | NO
EQPRepairJ22107618 | T6373-129 |  | Prober Down | TB1INS29 | NA01404H-A-AT-C | GGR689.1-BC21112 |  | Prober Down--- 1.機台自動執行Clean Wafer 清針時，發生Arm1 Vacuum Error. 2.查看機台有開啟Delay Time 仍發生異常. 3.此為低溫型號測試，開機至第一段 先行升溫待確認. | 13749 | Prober | 手臂異常(Arm Error) | 2022/08/13 13:26:13 | 2022/08/13 14:23:24 | 0.95 | OK | 13749 | 1.Slot #05首片測試發生Arm1 Vacuum Error. 2.Prober 重開 C- Resetup #15 PASS OK. (25.5) 3.Slot #05 Hold To TPD T44 STDF. | NO | NO | NO | YES | YES | NO
EQPRepairJ22107622 | T6373-33 | CWC7J0A88-01 | OverKill | JB1TAS14 | PA5208KF2 | J68R0 | 05 | #03,04 2H / Corr wafer overdue-19617 | CPAUTO | Over Kill | 2H | 2022/08/13 13:33:21 | 2022/08/13 13:35:44 | 0.04 | OK | 17680 | remote----兩片皆為不同bin,keep test | NO | NO | NO | YES | YES | NO
EQPRepairJ22107632 | T6373-111 | CWC7J0562-01.002 | OverKill | TB1R8S74 | PA5137F-LT1 | P07HH.1 | 16 | always O/S alarm-A6954 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/13 13:56:34 | 2022/08/13 14:09:40 | 0.22 | OK | 20231 | 1.清針起測一段穩定pass，PM(143,1) | NO | NO | NO | YES | YES | NO
EQPRepairJ22107636 | T6373-43 | CWC7J0L43-01 | OverKill | VB1BVS01 | PA5270D-LT | U38F5.00 | 20 | TOO MANY BIN34 | A6668 | Over Kill | Function誤宰 | 2022/08/13 14:08:46 | 2022/08/13 14:14:00 | 0.09 | OK | 19771 | 確認後pass還機續測 | NO | NO | NO | YES | YES | NO
EQPRepairJ22107645 | T6373-27 | CWC8J0610-01.001 | OverKill | JB1LWS16 | ILISLQPEDQD-2CP | ILA14006.1 | 25 | #25 too many o/s-19617 | CPAUTO | Over Kill | O/S | 2022/08/13 14:39:27 | 2022/08/13 15:04:52 | 0.42 | OK | 11385 | 溫針重新對針調整針位 (177.0) | NO | NO | NO | YES | YES | NO
EQPRepairJ22107648 | T6373-07 |  | Prober Down | JB1TAS20 | PA5208KF-LT | J682N |  | AlarmCode:1060 AlarmMessage:CANNOT MEASURE INDEX SIZE !! | CPAUTO | Prober | Wafer Edge/Mode Not Found | 2022/08/13 14:44:28 | 2022/08/13 14:47:37 | 0.05 | OK | 13749 | 重抓 CHECK DIE OK | NO | NO | NO | YES | YES | NO
EQPRepairJ22107650 | T6373-35 |  | Tester Down | JB1GVS19 | NC01682A(COG48AA)-A-AT-N3-C | GFP985.1-BC12211C |  | Tester down----1.改測NA01687G(COG48BA)-A-AT-N2-C / CP6 All Bin56 Function，更換配件一樣 2.Full call PASS上機一樣Bin56，待後續維修 | 11385 | Tester | Open / Short / Function Fail | 2022/08/13 14:56:50 | 2022/08/13 15:52:58 | 0.94 | OK | 16631 | 1.排程mail告知IDLE | NO | NO | NO | YES | YES | NO
EQPRepairJ22107652 | T6373-112 |  | Tester Down | JB1GVS25 | NC01682B(COG48AA)-A-AT-N-C | GGR939.1-BC2 |  | ENG FAIL:Diag Fail 105450. | 17001 | Tester | Diag Fail | 2022/08/13 15:05:10 | 2022/08/13 20:12:03 | 5.11 | OK | 16532 | DIAG 105450 185810 FAIL SWAP LCD#1<>#3 HVDC#1<>#2 更換備品PASS | NO | NO | NO | NO | YES | YES
EQPRepairJ22107665 | T6373-43 | CWC7J0L43-01 | OverKill | VB1BVS01 | PA5270D-LT | U38F5.00 | 20 | bin 34-13081 | CPAUTO | Over Kill | Function誤宰 | 2022/08/13 16:07:23 | 2022/08/13 16:33:49 | 0.44 | OK | 08774 | 安裝風管 續測觀察 | NO | NO | NO | YES | YES | NO
EQPRepairJ22107666 | T6373-32 | CWC8J0348-01.017 | OverKill | JB1HKS07 | NA01687G(COG48BA)-A-AT-N2-C | GGR356.1-B11C12122 | 18 | fail bin 50-20022 | CPAUTO | Over Kill | bin50 | 2022/08/13 16:07:51 | 2022/08/13 17:21:55 | 1.23 | OK | 11385 | 下卡磨針(36.0) | NO | NO | NO | YES | YES | NO
EQPRepairJ22107688 | T6373-29 | CWC8J0457-01 | Tester Down | JB1QNS06 | NA01687G(COG48BA)-A-AT-N2-C | GGR761.1-BC1 | 15 | work mode eng-A6049 | CPAUTO | Product Issue | 自動化異常 | 2022/08/13 17:51:51 | 2022/08/13 18:00:26 | 0.14 | OK | 16631 | 1.疑似自動化未自動move in，造成stdf卡關，退片請產線重新上貨並手動切機況 | NO | NO | NO | YES | YES | NO
EQPRepairJ22107701 | T6373-103 | CWC8J0622-01 | OverKill | JB6K8S02 | NC01682A(COG48AA)-A-AT-N2-C | GGT630.1-BC | 05 | EDA SHOW I NO CORR WAFER-A6893 | CPAUTO | Over Kill | SHOW I | 2022/08/13 18:37:23 | 2022/08/13 18:42:32 | 0.09 | OK | 16631 | 1.#01 SHOW I，#02 03 04 SHOW N，續測 | NO | NO | NO | YES | YES | NO
EQPRepairJ22107712 | T6373-136 | CWC8J0514-01 | OverKill | JB8Q6S05 | NA01761D(COF94AA)-A-Y-C | GGS422.1-BC | 02 | #1 T44 1I-20022 | CPAUTO | Over Kill | 單片I | 2022/08/13 19:14:59 | 2022/08/13 19:32:02 | 0.28 | OK | 17674 | 1. 單片I [HW_Bin30<=90.00％(1448EA)]1405(IF HW_Bin30<=90.00％, Please HOLD)，Short 、function過多 | NO | NO | NO | YES | YES | NO
EQPRepairJ22107739 | T6373-70 | CWC8J0440-01 | Other | JB1Q4S17 | PA6542C | J6C0T | 01 | plaese check probe mark NG too large-A6497 | CPAUTO | Prober | 針位異常(Probe Mark Shift) | 2022/08/13 21:08:37 | 2022/08/13 21:37:07 | 0.48 | OK | 17674 | check P/M ------ 1. 換批上貨掛確認針位，#01 查看未測 die已有針痕針點偏大貼上，故判斷為前流程造成 cp1 2. 重對針pass，#01 (88,2)溫針 range y=8 (40,0) | NO | NO | NO | YES | YES | NO
EQPRepairJ22107740 | T6373-43 | CWC7J0L43-01 | OverKill | VB1BVS01 | PA5270D-LT | U38F5.00 | 22 | too many fail bin-A6924 | CPAUTO | Over Kill | Function誤宰 | 2022/08/13 21:09:01 | 2022/08/13 21:59:23 | 0.84 | OK | 17674 | 1. Function誤宰 34 重load pgm pass | NO | NO | NO | YES | YES | NO
EQPRepairJ22107742 | T6373-136 | CWC8J0514-01 | OverKill | JB8Q6S05 | NA01761D(COF94AA)-A-Y-C | GGS422.1-BC | 03 | EDA SHOW 2 CONT I-A7033 | CPAUTO | Over Kill | Show I | 2022/08/13 21:14:42 | 2022/08/13 21:23:22 | 0.14 | OK | 17542 | Remote處理 ----- Show I，[HW_P/DUTDIFF>4%]，不同批有相同問題(GGS330.1)，請產線先行續測待後續Run Corr驗證 | NO | NO | NO | YES | YES | NO
EQPRepairJ22107754 | T6373-135 | CWC8J0227-01 | OverKill | VB4BKS02 | PA5129LS_12U80H-LT | P07CJ.1 | 09 | #1,2 eda shows contionous 2R DUE TO O/S LEADERS COMMAND MBD-A6788 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/13 22:12:38 | 2022/08/13 22:36:03 | 0.39 | OK | 03426 | 使用Clean Wafer 清針 | NO | NO | NO | YES | YES | NO
EQPRepairJ22107762 | T6373-136 | CWC8J0514-01 | Prober Down | JB8Q6S05 | NA01761D(COF94AA)-A-Y-C | GGS422.1-BC | 03 | IT FAILED IN LOCK/UNLOCK OF THE MANUAL RING LELVER-A6497 | CPAUTO | Prober | Lock / Unlock Error | 2022/08/13 22:54:38 | 2022/08/13 23:17:59 | 0.39 | OK | 17542 | Lock / Unlock Error --- 調整後嘗試多次無異常，還機待觀察 | NO | NO | NO | YES | YES | NO
EQPRepairJ22107763 | T6373-43 | CWC7J0L43-01 | OverKill | VB1BVS01 | PA5270D-LT | U38F5.00 | 23 | too many fail bin-A6786 | CPAUTO | Over Kill | Function誤宰 | 2022/08/13 23:01:19 | 2022/08/13 23:19:05 | 0.30 | OK | 15171 | Remote處理 --- BIN34，LOG電流值過高，AL ON PASS | NO | NO | NO | YES | YES | NO
EQPRepairJ22107772 | T6373-29 | CWC8J0248-01.003 | OverKill | JB1QNS06 | NA01687G(COG48BA)-A-AT-N2-C | GGS267.1-B11C12 | 04 | ALREADY OVERKILL NOT EFFECIVE,TOO MANY FAILBIN | A6404 | Over Kill | Function誤宰 | 2022/08/13 23:23:22 | 2022/08/14 01:34:57 | 2.19 | OK | 17542 | 低溫型號，下卡磨針上機pass，range:16，試打(170,3)無異常，點測pass，還機待觀察 | NO | NO | YES | YES | YES | NO
EQPRepairJ22107774 | T6373-111 | CWC7J0J51-01 | OverKill | TB1R8S74 | PA5137F-LT1 | P094H.1 | 11 | too many o/s-A6924 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/13 23:41:46 | 2022/08/13 23:54:25 | 0.21 | OK | 20276 | 清針後 pass range 10 P/M(118.6) | NO | NO | NO | YES | YES | NO
EQPRepairJ22107789 | T6373-135 | CWC8J0227-01 | OverKill | VB4BKS02 | PA5129LS_12U80H-LT | P07CJ.1 | 15 | #13,14 CONTINUOUS 2H O/S-A6497 | CPAUTO | Over Kill | Show H | 2022/08/14 00:53:46 | 2022/08/14 00:59:24 | 0.09 | OK | 17542 | Remote處理 ----- Show H，[HW_BIN72>0.3%](Bin 73 > 2EA) 3，Short Bin73過多(零星Short)，不同批有相同問題(P060P.1)，續測 | NO | NO | NO | YES | YES | NO
EQPRepairJ22107791 | T6373-21 | CWC8J0615-01.003 | Prober Down | JB1GKS09 | PA6543B | J67H7 | 01 | cannot measure index-A6404 | CPAUTO | Prober | Wafer Edge/Mode Not Found | 2022/08/14 01:03:37 | 2022/08/14 01:10:22 | 0.11 | OK | 17542 | Re-start ok，check die ok，P/M(162,1)ok | NO | NO | NO | YES | YES | NO
EQPRepairJ22107802 | T6373-43 | CWC7J0L43-01 | OverKill | VB1BVS01 | PA5270D-LT | U38F5.00 | 24 | TOO MANY FAILBIN-A6497 | CPAUTO | Over Kill | Function誤宰 | 2022/08/14 01:42:42 | 2022/08/14 04:38:00 | 2.92 | OK | 17674 | 1. Function誤宰 21 更換在庫卡#02 烘烤上機，Tester swap MDC #02 <>03 ，Cal 機 fail ，轉掛Tester down ，休息 | NO | NO | NO | YES | YES | NO
EQPRepairJ22107803 | T6373-34 | CWC8J0088-01.001 | OverKill | JB1BYS10 | PA6587B_15U90H | J696C | 05 | 連續2片EDA秀"H" Continuous 2 wafers EDA shows "H" os | A5929 | Over Kill | O/S誤宰 | 2022/08/14 01:44:18 | 2022/08/14 01:51:22 | 0.12 | OK | 03426 | 針位調整 機台清針PASS | NO | NO | NO | YES | YES | NO
EQPRepairJ22107818 | T6373-15 | CWC8J0508-01 | OverKill | JB8Q6S04 | NA01761D(COF94AA)-A-Y-C | GGS421.1-BC | 11 | #10 EDA /I (CORR PASSED)-A6925 | CPAUTO | Over Kill | 單片I | 2022/08/14 02:25:19 | 2022/08/14 02:57:38 | 0.54 | OK | 17674 | 1. 單片I [HW_P/DUTDIFF>4%]、CORR PASS | NO | NO | NO | YES | YES | NO
EQPRepairJ22107834 | T6373-136 | CWC8J0514-01 | OverKill | JB8Q6S05 | NA01761D(COF94AA)-A-Y-C | GGS422.1-BC | 05 | STILL EDA SHOW I CORR PASSED-A6497 | CPAUTO | Over Kill | 2I | 2022/08/14 03:49:24 | 2022/08/14 04:42:14 | 0.88 | OK | 17674 | 1.2I Short、不同function 過多，CORR PASS | NO | NO | NO | YES | YES | NO
EQPRepairJ22107835 | T6373-103 | CWC8J0622-01 | OverKill | JB6K8S02 | NC01682A(COG48AA)-A-AT-N2-C | GGT630.1-BC | 12 | Continuous wafer EDA shows " I " No corr. wafer | A6525 | Over Kill | Show I | 2022/08/14 03:56:23 | 2022/08/14 04:04:24 | 0.13 | OK | 17542 | Show I，[HW_P/DUTDIFF>5%]，不同Dut Low Yield，故續測 | NO | NO | NO | YES | YES | NO
EQPRepairJ22107846 | T6373-43 |  | Tester Down | VB1BVS01 | PA5270D-LT | U38F5.00 |  | Tester Down----- 1. Function誤宰 34 更換在庫卡#02 烘烤上機，Tester swap MDC #02 <>03 ，Cal 機 fail ，轉掛Tester down | 17674 | Tester | Other | 2022/08/14 04:39:01 | 2022/08/14 06:42:34 | 2.06 | OK | 17674 | Tester down ---- 1. Function誤宰 34 更換在庫卡#02 烘烤上機，Tester swap MDC #02<>03 ，Cal 機 fail 2. Swap HVDC #09<>#08 Cal 機pass 3. 低溫換卡、機台重新上機pass，RANGE Y=10 (172,2)#24 | NO | NO | NO | YES | YES | NO
EQPRepairJ22107854 | T6373-121 | CWC7J0197-01.002 | OverKill | TB1INS25 | NA01404H-A-AT-C | GGR689.1-BC121 | 18 | Too many o/s-A6053 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/14 04:56:51 | 2022/08/14 05:19:26 | 0.38 | OK | 11348 | 重新對針PASS | NO | NO | NO | YES | YES | NO
EQPRepairJ22107855 | T6373-37 | CWC8J0012-01 | Corr Fail | TB2U7S09 | R69005A0FNCPJDQV-P | AYN442000-B1 | 10 | setup-07658 | CPAUTO | Corr | Corr Fail | 2022/08/14 05:04:39 | 2022/08/14 05:14:30 | 0.16 | OK | 17542 | Corr Fail---- Corr 良率比 Corr Wafer 好，前流程 Fail Bin 回收成 Pass，手動比對 Corr Pass | NO | NO | NO | YES | YES | NO
EQPRepairJ22107867 | T6373-37 | CWC8J0011-01 | OverKill | TB2U7S09 | R69005A0FNCPJDQV-P | AYN442000-B2C | 13 | too many o/s,probecard have particles | A6404 | Over Kill | O/S誤宰 | 2022/08/14 05:59:44 | 2022/08/14 07:39:38 | 1.67 | OK | 20231 | 1.夜班交皆低溫型號卡毛下卡，上機降溫中，日班接手重對針、調針位pass，rg:18，PM(177,1)，中央od5可pass，邊緣需od14才可pass | NO | NO | NO | YES | YES | NO
EQPRepairJ22107872 | T6373-121 | CWC7J0197-01.002 | OverKill | TB1INS25 | NA01404H-A-AT-C | GGR689.1-BC121 | 19 | too many O/S(overkill done apply) | A6296 | Over Kill | O/S誤宰 | 2022/08/14 06:17:29 | 2022/08/14 06:27:04 | 0.16 | OK | 11348 | 二分法調針Pass，Range:10um | NO | NO | NO | YES | YES | NO
EQPRepairJ22107879 | T6373-21 |  | Prober Down | JB1GKS09 | PA6543B | J6AC7 |  | AlarmCode:1058 AlarmMessage:MODEL NOT FOUND ( FINE ALIGNMENT ) !! | CPAUTO | Prober | Wafer Edge/Mode Not Found | 2022/08/14 06:41:56 | 2022/08/14 06:46:04 | 0.07 | OK | 03426 | 重抓Module 查檢Check each ok | NO | NO | NO | YES | YES | NO
EQPRepairJ22107889 | T6373-111 | CWC7J0L14-01.001 | OverKill | TB1R8S74 | PA5137E1-LT | P097A.1 | 03 | too many o/s | A6897 | Over Kill | O/S誤宰 | 2022/08/14 07:33:11 | 2022/08/14 09:14:28 | 1.69 | OK | 13749 | 1.#03 配件金絲沾黏 造成 o/s ，更換備卡#S74->#S82 OK. 2.上機等待20分鐘後開始降溫 Setup pass (45.0) | NO | NO | NO | YES | YES | NO
EQPRepairJ22107896 | T6373-37 | CWC8J0011-01 | OverKill | TB2U7S09 | R69005A0FNCPJDQV-P | AYN442000-B2C | 13 | too many o/s-A6919 | CPAUTO | Over Kill | O/D不足 | 2022/08/14 08:02:02 | 2022/08/14 08:17:49 | 0.26 | OK | 20231 | 1.重新對針pass，PM(178,1) | NO | NO | NO | YES | YES | NO
EQPRepairJ22107906 | T6373-121 | CWC7J0H70-01 | OverKill | TB1INS25 | NA01404H-A-AT-C | GGT500.1-B1C1 | 01 | too many o/s | A6897 | Over Kill | O/D不足 | 2022/08/14 08:30:23 | 2022/08/14 08:52:37 | 0.37 | OK | 17712 | 二分法調針，OD+5um，#1(189.4)RG10 | NO | NO | NO | YES | YES | NO
EQPRepairJ22107912 | T6373-43 | CWC7J0L43-01 | OverKill | VB1BVS02 | PA5270D-LT | U38F5.00 | 24 | bin34 | A6921 | Over Kill | Function誤宰 | 2022/08/14 08:47:36 | 2022/08/14 08:57:17 | 0.16 | OK | 19771 | 調整針位OK | NO | NO | NO | YES | YES | NO
EQPRepairJ22107918 | T6373-121 | CWC7J0H70-01 | Tester Down | TB1INS25 | NA01404H-A-AT-C | GGT500.1-B1C1 | 01 | work mode eng-A6049 | CPAUTO | Other | Other | 2022/08/14 08:57:49 | 2022/08/14 09:07:12 | 0.16 | OK | 13749 | 1.#01 機台狀態未掛Run； TA Prober起測造成系統卡關. 2.依Run Card指示退片重新上貨測試. | NO | NO | NO | YES | YES | NO
EQPRepairJ22107920 | T6373-29 | CWC8J0347-01.003 | OverKill | JB1QNS06 | NA01687F(COG48BA)-A-AT-C | GGU954.1-BC222 | 01 | TOO MANY FAILED BIN | A6919 | Tester | Bin13 | 2022/08/14 09:13:41 | 2022/08/14 11:26:23 | 2.21 | OK | 11385 | Tester down---1.NA01687F(COG48BA)-A-AT-C / CP3 測試Bin13 Function過多，更換JB1QNS01.04.06一樣Bin13問題2.cal init待後續維修 | NO | NO | NO | YES | YES | NO
EQPRepairJ22107921 | T6373-43 | CWC7J0L43-01 | OverKill | VB1BVS02 | PA5270D-LT | U38F5.00 | 24 | o/s, p/c has particle-a6921 | CPAUTO | Over Kill | 機台清針後Pass | 2022/08/14 09:14:57 | 2022/08/14 09:21:04 | 0.10 | OK | 16631 | 1.機台清針後Pass | NO | NO | NO | YES | YES | NO
EQPRepairJ22107932 | T6373-33 | CWC7J0A88-01 | OverKill | JB1TAS14 | PA5208KF2 | J68R0 | 13 | p/c has particle too many o/s | A6897 | Over Kill | 配件卡毛 | 2022/08/14 09:49:13 | 2022/08/14 10:05:40 | 0.27 | OK | 17680 | REMOTE----前蓋清針，起測穩定 | NO | NO | NO | YES | YES | NO
EQPRepairJ22107942 | T6373-43 | CWC7J0L43-01 | OverKill | VB1BVS02 | PA5270D-LT | U38F5.00 | 24 | bin34 too many-12254 | CPAUTO | Over Kill | Function誤宰 | 2022/08/14 10:26:49 | 2022/08/14 10:47:53 | 0.35 | OK | 16631 | 1.溫針PASS | NO | NO | NO | YES | YES | NO
EQPRepairJ22107959 | T6373-43 | CWC7J0L43-01 | OverKill | VB1BVS02 | PA5270D-LT | U38F5.00 | 24 | too many fail bin-A6671 | CPAUTO | Over Kill | Function誤宰 | 2022/08/14 11:05:16 | 2022/08/14 11:10:11 | 0.08 | OK | 19771 | 調整針位OK | NO | NO | NO | YES | YES | NO
EQPRepairJ22107962 | T6373-33 | CWC7J0A88-01 | OverKill | JB1TAS14 | PA5208KF2 | J68R0 | 14 | CONTINUOUS EDA SHOW 2H,,CORR PASSEDD ALREDY-A6671 | CPAUTO | Over Kill | 2h | 2022/08/14 11:10:50 | 2022/08/14 11:18:22 | 0.13 | OK | 20231 | [HW_BIN0<94%](Bin 1 < 820EA) 815(Hold Wafer) [HW_BIN20>0.5%](Bin 21 > 4EA) 5(Hold Wafer)----- 1.查看eda資料，不同貨批(J6874-T)不同機卡有相同問題，續測在觀察 | NO | NO | NO | YES | YES | NO
EQPRepairJ22107963 | T6373-29 |  | Tester Down | JB1QNS06 | NA01687F(COG48BA)-A-AT-C | GGU954.1-BC222 |  | Tester down---1.NA01687F(COG48BA)-A-AT-C / CP3 測試Bin13 Function過多，更換JB1QNS01.04.06一樣Bin13問題2.cal init待後續維修 | 11385 | Tester | bin13 | 2022/08/14 11:26:56 | 2022/08/14 13:50:59 | 2.40 | OK | 11385 | Tester down---1.NA01687F(COG48BA)-A-AT-C / CP3 測試Bin13 Function過多，更換JB1QNS01.04.06一樣Bin13問題2.cal init PASS，上機點測皆PASS穩定，STDF退片重新上貨 | NO | NO | NO | YES | YES | NO
EQPRepairJ22107964 | T6373-43 | CWC7J0L43-01 | Prober Down | VB1BVS02 | PA5270D-LT | U38F5.00 | 25 | AlarmCode:1182 AlarmMessage:NEEDLE ALIGNMENT ERROR (# OF NEEDLE POSITIONS FOUND LESS THAN EXPECTED) !! | CPAUTO | Prober | Needle Alighment Error | 2022/08/14 11:28:14 | 2022/08/14 11:44:52 | 0.28 | OK | 19771 | 重新對針 #25(14.7) | NO | NO | NO | YES | YES | NO
EQPRepairJ22107965 | T6373-38 |  | Other | JB1GKS10 | PA6543B | J67H5 |  | please set 3D clean mode | A6668 | Over Kill | Other | 2022/08/14 11:29:39 | 2022/08/14 11:33:30 | 0.06 | OK | 11385 | 設定ok | NO | NO | NO | YES | YES | NO
EQPRepairJ22107968 | T6373-12 | CWC8J0562-01 | OverKill | JB1PNS06 | NA01761D(COF94AA)-A-Y-C | GGS330.1-BC1111 | 04 | #14 check P/M o/s | 20022 | Over Kill | O/S | 2022/08/14 11:44:31 | 2022/08/14 12:37:11 | 0.88 | OK | 11385 | 溫針重新對針調整針位 (197.1) | NO | NO | NO | YES | YES | NO
EQPRepairJ22107970 | T6373-128 | CWC7J0M27-01.004 | OverKill | JB1BYS11 | PA6587B_15U90H | J696G | 13 | #11.12 EDA SHOW 2R , O/S-A6986 | CPAUTO | Over Kill | 2R | 2022/08/14 11:48:07 | 2022/08/14 13:57:07 | 2.15 | OK | 19771 | 無備卡 配件磨針 #13(132.7) | NO | NO | NO | YES | YES | NO
EQPRepairJ22107974 | T6373-136 | CWC8J0514-01 | OverKill | JB8Q6S05 | NA01761D(COF94AA)-A-Y-C | GGS422.1-BC | 10 | #9 EDA shows "I", ALREADY CORR PASSED | A6986 | Over Kill | Show I | 2022/08/14 12:11:38 | 2022/08/14 12:19:17 | 0.13 | OK | 08819 | Remote處理 --- 1.#09 Show I，不同機卡不同Dut皆有相同問題。 2.Corr Passed，續測。 | NO | NO | NO | YES | YES | NO
EQPRepairJ22107993 | T6373-13 | CWC8J0612-01.002 | OverKill | JB1GVS19 | NC01682A(COG48AA)-A-AT-N3-C | GGS288.1-BC22 | 05 | o/s-17436 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/14 14:35:22 | 2022/08/14 14:56:18 | 0.35 | OK | 17712 | 機台清針，確認針痕針壓，#5(139.6)rg8 | NO | NO | NO | YES | YES | NO
EQPRepairJ22108001 | T6373-33 | CWC7J0A88-01 | OverKill | JB1TAS14 | PA5208KF2 | J68R0 | 15 | O/S-17436 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/14 15:16:40 | 2022/08/14 16:13:18 | 0.94 | OK | 17680 | REMOTE-----REPLACE #24,POSITION(0,5),MARK(46,0) | NO | NO | NO | YES | YES | NO
EQPRepairJ22108002 | T6373-128 | CWC7J0M27-01.004 | OverKill | JB1BYS11 | PA6587B_15U90H | J696G | 13 | BIN 12 | A6921 | Over Kill | Function誤宰 | 2022/08/14 15:21:58 | 2022/08/14 16:18:57 | 0.95 | OK | 16631 | 1.測試中BIN12 FUNCTION，下卡磨針上機相同，查看先前有多次相同送修紀錄，故配件送修，無備卡。 | NO | NO | NO | YES | YES | NO
EQPRepairJ22108007 | T6373-38 | CMC8J0010-01 | OverKill | JB1GKS10 | PA6543B | J6630 | 01 | o/s-a6571 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/14 15:44:57 | 2022/08/14 16:05:28 | 0.34 | OK | 20231 | 1.清調針pass，二分法座標(98,1)，RG:12，PM(162,1) | NO | NO | NO | YES | YES | NO
EQPRepairJ22108013 | T6373-128 |  | Probe Card Down | JB1BYS11 | PA6587B_15U90H | J696G |  | Probe Card Down---- 1.測試中BIN12 FUNCTION，下卡磨針上機相同，查看先前有多次相同送修紀錄，故配件送修，無備卡。 | 16631 | Probe Card | Other | 2022/08/14 16:19:31 | 2022/08/15 01:01:48 | 8.70 | OK | 17674 | P/C DOWN ----- 1. 日班交接配件當改測，改極低溫型號，約 18:50開始降溫，依低溫改機SOP 改機上機Bin42 Function 下卡磨針 2. 接手，23:20 磨針上機 、23:40 降溫、00:45 溫度-40、上機點測 bin41、42、45 不穩，此批為特需單請產線正常量產批驗證 3. range y=10 (133,0)#09 | NO | NO | NO | YES | YES | NO
EQPRepairJ22108016 | T6373-38 | CMC8J0010-01 | OverKill | JB1GKS10 | PA6543B | J6630 | 01 | too many o/s-A6919 | CPAUTO | Over Kill | O/S | 2022/08/14 16:36:43 | 2022/08/14 17:09:26 | 0.55 | OK | 11385 | 重新對針調整針位清針 (131.0) | NO | NO | NO | YES | YES | NO
EQPRepairJ22108020 | T6373-35 |  | Prober Down | JB1GVS24 | NC01682B(COG48AA)-A-AT-N-C | GGR939.1-BC222 |  | AlarmCode:4827 AlarmMessage:ARM @@VACUUM ERROR !! | CPAUTO | Prober | 氣體供應異常(Air/Vacuum Supply Error) | 2022/08/14 17:08:46 | 2022/08/14 19:37:40 | 2.48 | OK | 03426 | 早班交接VACUUM ERROR 重新開機降溫 接手Set up PASS 請產線重新上貨 | NO | NO | NO | YES | YES | NO
EQPRepairJ22108026 | T6373-29 | CWC8J0347-01.003 | OverKill | JB1QNS04 | NA01687F(COG48BA)-A-AT-C | GGU954.1-BC222 | 04 | #1.2 2H no corr wafer-14212 | CPAUTO | Over Kill | show H | 2022/08/14 17:18:52 | 2022/08/14 18:48:19 | 1.49 | OK | 11385 | EDA H 更換JB1QNS06 (35.0) | NO | NO | NO | YES | YES | NO
EQPRepairJ22108032 | T6373-42 |  | Prober Down | JB1PNS07 | NA01761D(COF94AA)-A-Y-C | GGS330.1-BC1112 |  | AlarmCode:1057 AlarmMessage:MODEL NOT FOUND ( ROUGH ALIGNMENT ) !! | CPAUTO | Prober | Wafer Edge/Mode Not Found | 2022/08/14 17:51:22 | 2022/08/14 18:02:15 | 0.18 | OK | 16631 | 1.重抓位置/check die/針痕 | NO | NO | NO | YES | YES | NO
EQPRepairJ22108052 | T6373-39 | CWC8J0186-01 | OverKill | TB4C1S03 | PA1638K-10E-B-LT | RDKAA0000 | 12 | #11 1I O/S-20022 | CPAUTO | Over Kill | Show I | 2022/08/14 19:04:31 | 2022/08/14 19:34:15 | 0.50 | OK | 17542 | Show I，[!!! WARNING !!!OPEN AND SHORT MODE TEST]，T32A型號 單片Show I，請續測 | NO | NO | NO | YES | YES | NO
EQPRepairJ22108057 | T6373-43 | CWC7J0M33-01.002 | OverKill | VB1BVS02 | PA5270D-LT | U38F8.00 | 12 | CONTINOUS EDA SHOWS 2H NO CORR-A6788 | CPAUTO | Over Kill | Show H | 2022/08/14 19:28:04 | 2022/08/14 19:40:28 | 0.21 | OK | 17542 | Show H，[HW_BIN0<100%](Bin 1 < 798EA) 716，請產線先行續測待觀察 | NO | NO | NO | YES | YES | NO
EQPRepairJ22108064 | T6373-86 | CWC7J0M28-01 | OverKill | JB1Q4S23 | PA6542C | J683J | 12 | too many o/s | A6926 | Over Kill | O/S誤宰 | 2022/08/14 19:49:31 | 2022/08/14 20:07:08 | 0.29 | OK | 11348 | Remote 處理----重新對針，Range:10um | NO | NO | NO | YES | YES | NO
EQPRepairJ22108067 | T6373-38 | CMC8J0010-01 | OverKill | JB1GKS10 | PA6543B | J6630 | 03 | CONTINUOUS EDA SHOW 2H,NO CORR WAFER-A6881 | CPAUTO | Over Kill | Show H | 2022/08/14 20:00:50 | 2022/08/14 20:06:39 | 0.10 | OK | 03426 | Show H-[CUSTOMER_HOLD_CRITERIA] [HW_BIN0<94%](Bin 1 < 1064EA) 1006(Hold Wafer) -----請產線續測 | NO | NO | NO | YES | YES | NO
EQPRepairJ22108068 | T6373-39 | CWC8J0186-01 | OverKill | TB4C1S03 | PA1638K-10E-B-LT | RDKAA0000 | 15 | TOO MANY OPEN SHORT-A6786 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/14 20:01:13 | 2022/08/14 20:20:37 | 0.32 | OK | 11348 | Remote 處理----重新對針pass | NO | NO | NO | YES | YES | NO
EQPRepairJ22108073 | T6373-32 |  | Prober Down | KB1PNS01 | NA01761D(COF94AA)-A-Y-C | GGS331.1-BC2 |  | AlarmCode:1057 AlarmMessage:MODEL NOT FOUND ( ROUGH ALIGNMENT ) !! | CPAUTO | Prober | Wafer Edge/Mode Not Found | 2022/08/14 20:15:50 | 2022/08/14 20:22:13 | 0.11 | OK | 03426 | 重抓Module 查檢Check each ok 218.4 | NO | NO | NO | YES | YES | NO
EQPRepairJ22108076 | T6373-35 | CWC8J0558-01.005 | OverKill | JB1GVS24 | NC01682B(COG48AA)-A-AT-N-C | GGR939.1-BC222 | 04 | BIN18 TOO MANY-19697 | CPAUTO | Over Kill | Function誤宰 | 2022/08/14 20:18:54 | 2022/08/14 20:33:25 | 0.24 | OK | 03426 | Bin18 Leakage 接風管Pass | NO | NO | NO | YES | YES | NO
EQPRepairJ22108086 | T6373-35 | CWC8J0558-01.005 | OverKill | JB1GVS24 | NC01682B(COG48AA)-A-AT-N-C | GGR939.1-BC222 | 04 | TOO MANY O/S-A6053 | CPAUTO | Over Kill | O/D不足 | 2022/08/14 20:51:40 | 2022/08/14 21:09:16 | 0.29 | OK | 03426 | 重新對針PASS | NO | NO | NO | YES | YES | NO
EQPRepairJ22108090 | T6373-33 | CWC7J0A88-01 | OverKill | JB1TAS24 | PA5208KF2 | J68R0 | 18 | #16,#17 show 2h-19697 | CPAUTO | Over Kill | Show H | 2022/08/14 20:55:38 | 2022/08/14 21:16:47 | 0.35 | OK | 11348 | Show H 續測 | NO | NO | NO | YES | YES | NO
EQPRepairJ22108103 | T6373-39 |  | Prober Down | TB4C1S03 | PA1638K-10E-B-LT | RDKAA0000 |  | AlarmCode:1060 AlarmMessage:CANNOT MEASURE INDEX SIZE !! | CPAUTO | Prober | Model Not Found | 2022/08/14 21:42:52 | 2022/08/14 22:08:15 | 0.42 | OK | 17542 | Model Not Found --- Re-start OK，Check Die OK，P/M(169,4)OK | NO | NO | NO | YES | YES | NO
EQPRepairJ22108115 | T6373-39 |  | Prober Down | TB4C1S03 | PA1638K-10E-B-LT | RDKAA0000 |  | AlarmCode:1058 AlarmMessage:MODEL NOT FOUND ( FINE ALIGNMENT ) !! | CPAUTO | Prober | Wafer Edge/Mode Not Found | 2022/08/14 22:30:08 | 2022/08/14 22:41:19 | 0.19 | OK | 03426 | 重抓Module 查檢Check each ok | NO | NO | NO | YES | YES | NO
EQPRepairJ22108121 | T6373-43 | CWC7J0M33-01.002 | OverKill | VB1BVS02 | PA5270D-LT | U38F8.00 | 13 | TOO MANY FAIL BIN34-A6786 | CPAUTO | Over Kill | Function誤宰 | 2022/08/14 23:01:59 | 2022/08/14 23:41:02 | 0.65 | OK | 17674 | 1. Function誤宰 34 重Load PGM PASS | NO | NO | NO | YES | YES | NO
EQPRepairJ22108124 | T6373-135 | CWC8J0620-01 | OverKill | VB4BKS02 | PA5129LS_12U80H-LT | P07CK.1 | 16 | CONTINUOUS EDA SHOW 2H,O/S-A6881 | CPAUTO | Over Kill | 2H | 2022/08/14 23:05:18 | 2022/08/14 23:17:37 | 0.21 | OK | 17542 | 1. 2H (Bin 73 > 2EA) 3(hold wafer) ，SHORT過多 | NO | NO | NO | YES | YES | NO
EQPRepairJ22108125 | T6373-33 | CWC7J0A88-01 | OverKill | JB1TAS24 | PA5208KF2 | J68R0 | 18 | O/S-13225 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/14 23:09:47 | 2022/08/14 23:45:18 | 0.59 | OK | 11348 | 更換配件#09 pass Range:10um | NO | NO | NO | YES | YES | NO
EQPRepairJ22108128 | T6373-39 | CWC8J0186-01 | OverKill | TB4C1S03 | PA1638K-10E-B-LT | RDKAA0000 | 21 | CONTINUOUS EDA SHOW 2I-A6881 | CPAUTO | Over Kill | Show I | 2022/08/14 23:24:49 | 2022/08/14 23:45:48 | 0.35 | OK | 17542 | Show I，[!!! WARNING !!!OPEN AND SHORT MODE TEST]，零星Short過多，不同機台測試有相同問題，請續測 | NO | NO | NO | YES | YES | NO
EQPRepairJ22108133 | T6373-07 |  | Prober Down | JB1TAS20 | PA5208KF-LT | J68QR |  | AlarmCode:1060 AlarmMessage:CANNOT MEASURE INDEX SIZE !! | CPAUTO | Prober | Wafer Edge/Mode Not Found | 2022/08/14 23:46:26 | 2022/08/15 00:11:00 | 0.41 | OK | 11348 | 重抓ok | NO | NO | NO | YES | YES | NO
EQPRepairJ22108134 | T6373-21 | CWC8J0621-01 | OverKill | JB1GKS09 | PA6543B | J6AC8 | 03 | EDA 2 show H o/s-A6053 | CPAUTO | Over Kill | Show H | 2022/08/14 23:47:43 | 2022/08/15 00:42:52 | 0.92 | OK | 17542 | Show H，O/S不穩，更換配件#08，上機PASS，溫針重對針，range:12，試打(161,1)無異常，起測PASS | NO | NO | NO | YES | YES | NO
EQPRepairJ22108145 | T6373-29 | CWC8J0347-01.003 | OverKill | JB1QNS06 | NA01687F(COG48BA)-A-AT-C | GGU954.1-BC222 | 08 | #6.#7 SHOW 2H (NO CORR)-19697 | CPAUTO | Over Kill | Show H | 2022/08/15 00:49:53 | 2022/08/15 01:10:54 | 0.35 | OK | 17542 | Show H，Bin13 Function Fail過多，調整機台後方風管位置(放置一段時間)，點測pass穩定，請產線先行續測待觀察 | NO | NO | YES | YES | YES | NO
EQPRepairJ22108150 | T6373-43 | CWC7J0M33-01.002 | OverKill | VB1BVS02 | PA5270D-LT | U38F8.00 | 14 | TOO MANY FAILED BIN (BIN34) NO PARTICLE-A6925 | CPAUTO | Over Kill | Function誤宰 | 2022/08/15 01:27:43 | 2022/08/15 01:43:27 | 0.26 | OK | 17542 | 換片起測全面Bin34，機台清針 & 重Load程式，非燒Code型號 wafer alignment重測，起測pass穩定，還機待觀察 | NO | NO | NO | YES | YES | NO
EQPRepairJ22108152 | T6373-128 | CWC7J0K25-01.001 | OverKill | JB1BYS12 | PA6587B_15U90H | J696F | 01 | too many fail bin-A6786 | CPAUTO | Over Kill | Function誤宰 | 2022/08/15 01:38:55 | 2022/08/15 02:14:32 | 0.59 | OK | 17674 | 1. 改機全面性BIN42 Function Fail，下卡磨針上機Bin41、42、45不穩，清調針、換批皆無效，故將配件送修無備卡 | NO | NO | NO | YES | YES | NO
EQPRepairJ22108159 | T6373-128 |  | Probe Card Down | JB1BYS12 | PA6587B_15U90H | J696F |  | Probe Card Down---- 1. 改機全面性BIN42 Function Fail，下卡磨針上機Bin41、42、45不穩，清調針、換批皆無效，故將配件送修無備卡 | 17674 | Over Kill | O/S誤宰 | 2022/08/15 02:16:04 | 2022/08/15 06:15:14 | 3.99 | OK | 17674 | P/C DOWN ---- 0. 配件送修 改測高溫 1. TB1R8S67改機Pass， Range y=2，針位上貼下、下貼上 內縮 2. 更換備卡#77，上機Pass， Range y=4，針位內縮 3. 更換備卡#74，上機 Pass，Range y=4，針位上下貼 4. 更換備卡#84 上機Pass， Range y=6，故將以上配件送修 5. #24 (81,3)溫針，試打(44,0)，掛 S/U 04 | NO | NO | NO | YES | YES | NO
EQPRepairJ22108162 | T6373-152 |  | Setup Fail | JB252S01 | TD7850-A0S-AZ-CP | SA2PK-B2 |  | setup fail due to bin10 o/s already unmount the tester and change p/c JB252S01JB252S02 but sill same p/c polish-A6937 | CPAUTO | Setup Fail | O/S Fail | 2022/08/15 02:39:55 | 2022/08/15 05:09:06 | 2.49 | OK | 17674 | S/U FAIL ----- 1. 改機open 一大串不過 log無contact 、查看針位有針痕 2. 下卡時發現沒有雲母片，抬頭檢查雲母片在Pogopin 上面，重新擦拭雲母片，重新上機pass 掛s/u 02，休息 | NO | NO | NO | YES | YES | NO
EQPRepairJ22108164 | T6373-47 | CWC8J0390-01.001 | OverKill | JB1HKS06 | NA01687F(COG48BA)-A-AT-C | GGS941.1-1BC2 | 04 | too many fail bin and o/s,(36.4)pm ,od5um,check,pm up/ down close to the edge of the bump,no particle-A5717 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/15 02:42:18 | 2022/08/15 04:21:25 | 1.65 | OK | 17674 | 1. 更換備卡#07PASS ，#04 (92,2) 溫針 RANGE y=10 (163,1) | NO | NO | NO | YES | YES | NO
EQPRepairJ22108174 | T6373-141 | CWC8J0347-01.004 | OverKill | JB1HKS10 | NA01687F(COG48BA)-A-AT-C | GGU954.1-BC121 | 25 | Too many O/S Bin22 pm checked(87.0)-A6525 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/15 03:30:51 | 2022/08/15 03:43:00 | 0.20 | OK | 03426 | 針位調整PASS | NO | NO | NO | YES | YES | NO
EQPRepairJ22108175 | T6373-129 | CWC8J0457-01.001 | Prober Down | JB1QNS01 | NA01687G(COG48BA)-A-AT-N2-C | GGR761.1-BC2 | 22 | AlarmCode:1002 AlarmMessage:WAFER EDGE NOT FOUND | CPAUTO | Prober | Wafer Edge/Mode Not Found | 2022/08/15 03:32:58 | 2022/08/15 03:47:39 | 0.24 | OK | 13597 | 重抓OK | NO | NO | NO | YES | YES | NO
EQPRepairJ22108178 | T6373-33 | CWC7J0A88-01 | OverKill | JB1TAS19 | PA5208KF2 | J68R0 | 20 | TOO MANY O/S | A6926 | Over Kill | O/S誤宰 | 2022/08/15 03:43:49 | 2022/08/15 03:59:57 | 0.27 | OK | 03426 | 針位調整 機台清針PASS | NO | NO | NO | YES | YES | NO
EQPRepairJ22108186 | T6373-47 |  | Prober Down | JB1HKS07 | NA01687F(COG48BA)-A-AT-C | GGS941.1-1BC2 |  | AlarmCode:1057 AlarmMessage:MODEL NOT FOUND ( ROUGH ALIGNMENT ) !! | CPAUTO | Prober | Wafer Edge/Mode Not Found | 2022/08/15 04:45:50 | 2022/08/15 04:55:19 | 0.16 | OK | 15171 | Remote處理 --- START OK，CHECK DIE OK | NO | NO | NO | YES | YES | NO
EQPRepairJ22108190 | T6373-103 |  | Prober Down | JB6K8S02 | NC01682A(COG48AA)-A-AT-N2-C | GGS284.1-B1C22111 |  | AlarmCode:1060 AlarmMessage:CANNOT MEASURE INDEX SIZE !! | CPAUTO | Prober | Model Not Found | 2022/08/15 04:55:59 | 2022/08/15 05:22:56 | 0.45 | OK | 17542 | Model Not Found --- Re-start ok，Check Die ok，6-Site型號，range:14，P/M(16,5)ok | NO | NO | YES | YES | YES | NO
EQPRepairJ22108191 | T6373-33 | CWC7J0A88-01 | OverKill | JB1TAS19 | PA5208KF2 | J68R0 | 21 | Too many o/s-A6053 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/15 05:00:06 | 2022/08/15 05:33:52 | 0.56 | OK | 11348 | 更換配件#25 PASS Range:10um | NO | NO | NO | YES | YES | NO
EQPRepairJ22108195 | T6373-103 | CWC8J0189-01.003 | OverKill | JB6K8S02 | NC01682A(COG48AA)-A-AT-N2-C | GGS284.1-B1C22111 | 03 | DN check, corr. fail-A6722 | CPAUTO | Corr | Other | 2022/08/15 05:33:28 | 2022/08/15 05:36:33 | 0.05 | OK | 03426 | T44 產品無 Corr Wafer Run Corr 比對 Fail | NO | NO | NO | YES | YES | NO
EQPRepairJ22108196 | T6373-135 | CWC8J0704-01 | OverKill | VB4BKS02 | PA5129LS_12U80H-LT | P07GY.1 | 06 | #4,5 CONTINUES EDA SHOW 2H PM CHECK 136.4 OD12-A7024 | CPAUTO | Over Kill | SHOW H | 2022/08/15 05:33:53 | 2022/08/15 05:43:56 | 0.17 | OK | 11348 | Remote 處理----SHOW H 二分法調針Pass，Range:12um 續測 | NO | NO | NO | YES | YES | NO
EQPRepairJ22108200 | T6373-136 | CWC8J0514-01 | OverKill | JB8Q6S05 | NA01761D(COF94AA)-A-Y-C | GGS422.1-BC | 22 | #20.21 CORR PASSED BUT STILL EDA SHOW 2H-A5717 | CPAUTO | Over Kill | Show I | 2022/08/15 05:44:53 | 2022/08/15 06:02:22 | 0.29 | OK | 11348 | Remote 處理----Show I 無固定dut corr pass 續測 | NO | NO | NO | YES | YES | NO
EQPRepairJ22108203 | T6373-12 |  | Other | JB1PNS06 | NA01761D(COF94AA)-A-Y-C | GGS330.1-BC111111 |  | check p/m..( 169.8 )-A6148 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/15 06:04:05 | 2022/08/15 07:35:17 | 1.52 | OK | 10210 | 1.夜班交接更換配件#08 2.確認針位(53.0) | NO | NO | NO | YES | YES | NO
EQPRepairJ22108204 | T6373-152 | CWC7J0A18-01.003 | Corr Fail | JB252S02 | TD7850-A0S-AZ-CP | SA2PK-B2 | 16 | corr fail | A6788 | Corr | CORR FAIL | 2022/08/15 06:04:08 | 2022/08/15 06:17:51 | 0.23 | OK | 17674 | 1. CORR PASS | NO | NO | NO | YES | YES | NO
EQPRepairJ22108210 | T6373-39 | CWC8J0186-01 | OverKill | JB1FDS08 | PA1638K-10E-B-LT | RDKAA0000 | 15 | EDA SHOW 2H-A6786 | CPAUTO | Over Kill | Show H | 2022/08/15 06:28:59 | 2022/08/15 06:40:41 | 0.20 | OK | 17542 | Remote處理 ----- Show H，[HW_BIN20>1.0%](Bin 21 > 9EA) 10，不同機卡有相同問題(RDAXG1000-1-S)，機台清針，OD+3um，T32A型號No Corr，Setup時Run Corr首片比對PASS，故請產線先行續測待後續觀察 | NO | NO | NO | YES | YES | NO
EQPRepairJ22108212 | T6373-111 | CWC7J0L14-01.001 | OverKill | TB1R8S82 | PA5137E1-LT | P097A.1 | 16 | TOO MANY O/S RUN CORR | A6926 | Over Kill | O/S誤宰 | 2022/08/15 06:36:03 | 2022/08/15 06:58:33 | 0.38 | OK | 13597 | 清調針pass，range 8 um | NO | NO | NO | YES | YES | NO
EQPRepairJ22108220 | T6373-103 |  | Prober Down | JB6K8S02 | NC01682A(COG48AA)-A-AT-N2-C | GGS284.1-B1C22111 |  | AlarmCode:1060 AlarmMessage:CANNOT MEASURE INDEX SIZE !! | CPAUTO | Prober | Wafer Edge/Mode Not Found | 2022/08/15 07:08:56 | 2022/08/15 07:24:42 | 0.26 | OK | 13222 | Wafer Edge/Mode Not Found ok | NO | NO | NO | YES | YES | NO
EQPRepairJ22108223 | T6373-37 | CWC8J0203-01 | OverKill | TB2U7S09 | R69005A0FNCPJDQV-P | AYP684000-B1C | 12 | too many o/s | A6785 | Over Kill | O/D不足 | 2022/08/15 07:50:11 | 2022/08/15 08:04:33 | 0.24 | OK | 15362 | 重新對針,RANGE:16,P/M(176,8) | NO | NO | NO | YES | YES | NO
EQPRepairJ22108226 | T6373-86 | CWC8J0217-01 | OverKill | JB1Q4S15 | PA6542C | J68S2 | 03 | #3 PLS CHECK AND CONFIRM PM TOO MANY O/S-A6132 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/15 08:09:28 | 2022/08/15 08:31:01 | 0.36 | OK | 07764 | open 針痕貼下 , 調整針位pass #3(13.4) | NO | NO | NO | YES | YES | NO
EQPRepairJ22108227 | T6373-135 | CWC8J0704-01 | OverKill | VB4BKS02 | PA5129LS_12U80H-LT | P07GY.1 | 11 | #09.10 EDA 2H CORR PASS-05131 | CPAUTO | Over Kill | 2H | 2022/08/15 08:29:55 | 2022/08/15 08:31:23 | 0.02 | OK | 10210 | 2H 請產線續測 | NO | NO | NO | YES | YES | NO
EQPRepairJ22108228 | T6373-33 | CWC7J0A88-01 | OverKill | JB1TAS25 | PA5208KF2 | J68R0 | 22 | #22 PLS CHECK AND CONFIRM PM TOO MANY O/S-A6132 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/15 08:32:42 | 2022/08/15 08:57:27 | 0.41 | OK | 07764 | open清針 | NO | NO | NO | YES | YES | NO
EQPRepairJ22108232 | T6373-21 | CWC8J0621-01 | OverKill | JB1GKS08 | PA6543B | J6AC8 | 07 | too many o/s-A6942 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/15 09:13:50 | 2022/08/15 09:30:39 | 0.28 | OK | 15362 | 清調針PASS,Y RANGE:10,P/M(158,6) | NO | NO | NO | YES | YES | NO
EQPRepairJ22108234 | T6373-34 | CWC7J0G36-01.002 | OverKill | JB1BYS10 | PA6587B_15U90H | J696A | 19 | too many o/s | A6942 | Over Kill | o/s | 2022/08/15 09:25:50 | 2022/08/15 09:37:18 | 0.19 | OK | 13222 | 機台清針、調針ok (44.0)range=12 | NO | NO | NO | YES | YES | NO
EQPRepairJ22108248 | T6373-34 |  | Prober Down | JB1BYS10 | PA6587B_15U90H | J696A |  | AlarmCode:1060 AlarmMessage:CANNOT MEASURE INDEX SIZE !! | CPAUTO | Prober | Wafer Edge/Mode Not Found | 2022/08/15 10:07:42 | 2022/08/15 10:43:48 | 0.60 | OK | 13222 | Wafer Edge/Mode Not Found | NO | NO | NO | YES | YES | NO
EQPRepairJ22108249 | T6373-38 | CMC8J0010-01 | OverKill | JB1GKS10 | PA6543B | J6630 | 13 | TOO MANY O/S | A6785 | Over Kill | O/S | 2022/08/15 10:09:33 | 2022/08/15 10:44:22 | 0.58 | OK | 13222 | 機台清針OK | NO | NO | NO | YES | YES | NO
EQPRepairJ22108252 | T6373-33 | CWC7J0A88-01 | OverKill | JB1TAS25 | PA5208KF2 | J68R0 | 23 | TOO MANY O/S AND HAVE PARTICLE | A6131 | Over Kill | O/S誤宰 | 2022/08/15 10:18:55 | 2022/08/15 11:29:08 | 1.17 | OK | 07764 | open 卡毛 , 前蓋清針 | NO | NO | NO | YES | YES | NO
EQPRepairJ22108259 | T6373-32 | CWC8J0508-01.003 | Other | KB1PNS01 | NA01761D(COF94AA)-A-Y-C | GGS421.1-BC12 | 18 | pm check too big up and down risky device due to pm please check thanks | A6785 | Other | 確認針痕 | 2022/08/15 11:01:29 | 2022/08/15 11:39:32 | 0.63 | OK | 15362 | 二分法 x/y range:22/20,p/m(169,8) | NO | NO | NO | YES | YES | NO
EQPRepairJ22108265 | T6373-38 | CMC8J0010-01 | OverKill | JB1GKS10 | PA6543B | J6630 | 13 | TOO MANY O/S | A6873 | Over Kill | O/S誤宰 | 2022/08/15 11:30:31 | 2022/08/15 11:55:57 | 0.42 | OK | 20075 | 重新對針rnage 10 ,mark -1,3 | NO | NO | NO | YES | YES | NO
EQPRepairJ22108268 | T6373-29 | CWC8J0347-01.005 | OverKill | JB1QNS06 | NA01687F(COG48BA)-A-AT-C | GGU954.1-BC122 | 17 | eda 2H NO CORR WAFER BIN 13.15-A6873 | CPAUTO | Over Kill | 2H BIN13 | 2022/08/15 11:38:04 | 2022/08/15 12:08:32 | 0.51 | OK | 20075 | BIN13 多次處理無效,5點點測PASS,#17未有BIN13,續測觀察 | NO | NO | NO | YES | YES | NO
EQPRepairJ22108283 | T6373-43 | CWC7J0M33-01 | Corr Fail | VB1BVS02 | PA5270D-LT | U38F8.00 | 04 | t32 比對Corr Fail | 11550 | Corr | Other | 2022/08/15 12:36:56 | 2022/08/15 12:40:31 | 0.06 | OK | 10210 | 手動比對Pass | NO | NO | NO | YES | YES | NO
EQPRepairJ22108289 | T6373-12 | CWC8J0508-01.002 | OverKill | JB1PNS08 | NA01761D(COF94AA)-A-Y-C | GGS421.1-BC22 | 07 | TOO MANY O/S | A6942 | Over Kill | O/S誤宰 | 2022/08/15 13:34:02 | 2022/08/15 14:22:09 | 0.80 | OK | 14291 | clean wafer pass,od5>od8 P/M(54.0)RG:10 | NO | NO | NO | YES | YES | NO
EQPRepairJ22108296 | T6373-39 |  | Corr Fail | JB1FDS08 | PA1638K-10E-B-LT | RDKAA0000 |  | Corr Fail | 11550 | Corr | Prober Card問題-換Prober Card | 2022/08/15 13:51:35 | 2022/08/15 14:34:46 | 0.72 | OK | 20075 | 換卡#s08,RANGE 10,MARK -1,3 | NO | NO | NO | YES | YES | NO
EQPRepairJ22108298 | T6373-141 | CWC8J0753-01.001 | OverKill | JB1HKS10 | NA01687F(COG48BA)-A-AT-C | GGV504.1-BC2 | 05 | EDA 2H O/S | A6873 | Over Kill | 2H | 2022/08/15 13:54:51 | 2022/08/15 14:22:23 | 0.46 | OK | 13222 | 2H 機台清針OK | NO | NO | NO | YES | YES | NO
EQPRepairJ22108315 | T6373-33 | CWC7J0A88-01 | OverKill | JB1TAS25 | PA5208KF2 | J68R0 | 24 | TOO MANY O/S-A6942 | CPAUTO | Over Kill | 機台清針後Pass | 2022/08/15 15:22:05 | 2022/08/15 16:33:46 | 1.19 | OK | 20075 | 機台清針後Pass,164,3 | NO | NO | NO | YES | YES | NO
EQPRepairJ22108328 | T6373-129 |  | Setup Fail | JB1FMS08 | NC01267C(COG2)-A-AT-C | FASA73.1-B |  | it failed in lock/unlock manual ing lever-A6942 | CPAUTO | Setup Fail | Prober異常 prober abnormal | 2022/08/15 16:49:59 | 2022/08/15 17:10:13 | 0.34 | OK | 14291 | 1.改機 lock unlock error 2.手動調整驗證ok 3.請產線setup again | NO | NO | NO | YES | YES | NO
EQPRepairJ22108330 | T6373-21 | CWC8J0621-01 | OverKill | JB1GKS08 | PA6543B | J6AC8 | 11 | too many O/S-A6873 | CPAUTO | Over Kill | O/D不足 | 2022/08/15 16:53:31 | 2022/08/15 17:23:18 | 0.50 | OK | 14291 | Remote處理----重新對針pass P/M(131.0)RG:10 | NO | NO | NO | YES | YES | NO
EQPRepairJ22108338 | T6373-33 | CWC7J0A88-01 | OverKill | JB1TAS25 | PA5208KF2 | J68R0 | 25 | o/s | A6921 | Over Kill | O/S誤宰 | 2022/08/15 17:22:41 | 2022/08/15 18:01:29 | 0.65 | OK | 14291 | 換卡s26 pass P/M(45.0)RG:10 | NO | NO | NO | YES | YES | NO
EQPRepairJ22108341 | T6373-111 | CWC7J0761-01.007 | Prober Down | TB1R8S82 | PA5137F-LT1 | P07HJ.1 | 18 | AlarmCode:8732 AlarmMessage:It failed in lock/unlock of the manual ring lever. | CPAUTO | Prober | LOCK/UNLOCK | 2022/08/15 17:38:18 | 2022/08/15 17:53:59 | 0.26 | OK | 13222 | LOCK/UNLOCK 確認OK | NO | NO | NO | YES | YES | NO
EQPRepairJ22108343 | T6373-135 | CWC8J0746-01 | OverKill | VB4BKS02 | PA5129LS_12U80H-LT | P07HA.1 | 06 | EDA 2H, O/S | A6668 | Over Kill | 2H | 2022/08/15 17:41:18 | 2022/08/15 17:47:41 | 0.11 | OK | 10210 | 2H 請產線續測 | NO | NO | NO | YES | YES | NO
EQPRepairJ22108348 | T6373-152 | CWC8J0448-01.004 | OverKill | JB252S02 | TD7850-A0S-AZ-CP | SC1GT-B111C | 09 | o/s OD:15 | 08889 | Over Kill | O/S誤宰 | 2022/08/15 18:10:16 | 2022/08/15 18:17:28 | 0.12 | OK | 10210 | 調整針位(52.7) | NO | NO | NO | YES | YES | NO
EQPRepairJ22108349 | T6373-86 | CWC8J0217-01 | OverKill | JB1Q4S15 | PA6542C | J68S2 | 07 | O/S | A6921 | Over Kill | 機台清針後Pass | 2022/08/15 18:12:27 | 2022/08/15 18:25:37 | 0.22 | OK | 07764 | Remote處理----起測open , 檢查針痕ok , 清針pass | NO | NO | NO | YES | YES | NO
EQPRepairJ22108357 | T6373-112 | CWC8J0622-01 | OverKill | JB1GVS25 | NC01682A(COG48AA)-A-AT-N2-C | GGT630.1-BC1 | 10 | 已依誤宰處理流程處置無效 (O/S 過多) | 11550 | Over Kill | O/S誤宰 | 2022/08/15 18:24:54 | 2022/08/15 18:52:20 | 0.46 | OK | 20075 | 清調針RANGE 10,MARK 182,3 | NO | NO | NO | YES | YES | NO
EQPRepairJ22108360 | T6373-11 | CWC8J0515-01 | Prober Down | JB1PNS05 | NA01761D(COF94AA)-A-Y-C | GGS331.1-BC11 | 22 | AlarmCode:240 AlarmMessage:PREALIGNMENT ERROR !! | CPAUTO | Prober | Wafer Edge/Mode Not Found | 2022/08/15 18:34:33 | 2022/08/15 18:44:10 | 0.16 | OK | 07764 | Remote處理----stdf 掛run 維修 | NO | NO | NO | YES | YES | NO
EQPRepairJ22108362 | T6373-111 | CWC7J0761-01.007 | Prober Down | TB1R8S82 | PA5137F-LT1 | P07HJ.1 | 18 | AlarmCode:8732 AlarmMessage:It failed in lock/unlock of the manual ring lever. | CPAUTO | Prober | lock/unlock of the manual ring lever | 2022/08/15 18:49:49 | 2022/08/15 20:48:40 | 1.98 | OK | 10897 | 微調LOCK桿無效，後續確認氣壓筏，發現螺絲脫落，重新固定正常，低溫型號等待20分鐘，與降溫重下針壓PASS，試打(115.0) | NO | NO | NO | NO | YES | NO
EQPRepairJ22108377 | T6373-128 | CWC7J0A84-01 | OverKill | TB1R8S84 | PA5137F-LT1 | P094G.1 | 06 | TOO MANY OS | A6512 | Over Kill | O/S誤宰 | 2022/08/15 20:09:36 | 2022/08/15 20:24:16 | 0.24 | OK | 20182 | 清針調針PASS RG-8 PM(44,6) | NO | NO | NO | YES | YES | NO
EQPRepairJ22108388 | T6373-34 | CWC7J0G36-01.002 | OverKill | JB1BYS10 | PA6587B_15U90H | J696A | 25 | P/C HAS A PARTICLE | A6857 | Over Kill | 配件卡毛 | 2022/08/15 20:42:02 | 2022/08/15 20:53:27 | 0.19 | OK | 17091 | 不影響測試 二分法調針r12~24 pm(18，6) 起測觀察一段穩定還機 | NO | NO | NO | YES | YES | NO
EQPRepairJ22108392 | T6373-21 | CWC8J0621-01 | OverKill | JB1GKS08 | PA6543B | J6AC8 | 13 | TOO MANY O/S PC HAS A PARTICLE | A6857 | Over Kill | O/S誤宰 | 2022/08/15 20:59:47 | 2022/08/15 21:12:38 | 0.21 | OK | 15492 | 1.零星 OI/S，清針 PASS | NO | NO | NO | YES | YES | NO
EQPRepairJ22108396 | T6373-38 | CWC8J0752-01 | OverKill | JB1GKS10 | PA6543B | J6AK6 | 20 | too many O/S-A6857 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/15 21:21:55 | 2022/08/15 21:51:32 | 0.49 | OK | 15492 | 1.換片起測 O/S，針位貼下，二分法調針位(4，16)、RANGE:12 2.P/M(158,6)OK | NO | NO | NO | YES | YES | NO
EQPRepairJ22108398 | T6373-129 | CWC2J0Y28-01 | Setup Fail | JB1LBS03 | NB01682F(COG48BA)-A-AT-Z-C | GFSW63.1-B1 | 20 | set up fail due to prober down, AE said change to MBD | A6943 | Prober | Other | 2022/08/15 21:30:58 | 2022/08/15 21:51:10 | 0.34 | OK | 09108 | 1.setup 無法抬頭 Prober reset 無效 2. 調整 sensor ok 還機setup | NO | NO | NO | YES | YES | NO
EQPRepairJ22108401 | T6373-99 | CWC7J0L61-01.001 | OverKill | JB1TAS16 | PA5208K-LT | QT7CH | 19 | o/s-03491 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/15 21:41:01 | 2022/08/15 22:02:00 | 0.35 | OK | 20277 | Remote---- 1.Range:14 機台清針 PASS | NO | NO | NO | YES | YES | NO
EQPRepairJ22108419 | T6373-34 | CWC7J0G36-01.002 | OverKill | JB1BYS10 | PA6587B_15U90H | J696A | 25 | p/c has a particle-A6516 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/15 22:23:28 | 2022/08/15 23:19:52 | 0.94 | OK | 17091 | 下卡磨針上機PAss r:10um | NO | NO | NO | YES | YES | NO
EQPRepairJ22108420 | T6373-33 | CWC7J0416-01.002 | OverKill | JB1TAS26 | PA5208KF2 | J68GA | 06 | TOO MANY O/S | A6675 | Over Kill | O/S誤宰 | 2022/08/15 22:33:35 | 2022/08/15 23:54:13 | 1.34 | OK | 17184 | 1.磨針(148.5) | NO | NO | NO | NO | YES | NO
EQPRepairJ22108431 | T6373-47 | CWC8J0753-01.002 | OverKill | JB1HKS07 | NA01687F(COG48BA)-A-AT-C | GGV504.1-BC12 | 18 | already do the overkill process but still too many o/s-A6937 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/15 23:27:59 | 2022/08/15 23:40:31 | 0.21 | OK | 10110 | Remote---clean wafer清針調針後pass | NO | NO | NO | YES | YES | NO
EQPRepairJ22108436 | T6373-34 |  | Tester Down | JB1BYS10 | PA6587B_15U90H | J696A |  | GPIB-A6857 | CPAUTO | Over Kill | LOAD PGM FAIL | 2022/08/16 00:00:59 | 2022/08/16 00:16:40 | 0.26 | OK | 15492 | 1.TDR FAIL,重 MNT PASS 2.RELOAD PGM OK | NO | NO | NO | YES | YES | NO
EQPRepairJ22108442 | T6373-151 | CWC8J0038-01 | OverKill | JB2P2S08 | R63455-A1S-EZ-CP | NAPY65.00-B1C | 11 | TOO MANY BIN 23-A6512 | CPAUTO | Over Kill | Function誤宰 | 2022/08/16 00:25:14 | 2022/08/16 00:40:09 | 0.25 | OK | 20277 | Remote---- 1.Bin23 Pass-->Pass，Fail-->Fail穩定，Bin23不同機卡也無法回收，產品問題，續測 | NO | NO | YES | YES | YES | NO
EQPRepairJ22108444 | T6373-111 | CWC8J0440-01 | Prober Down | JB1Q4S23 | PA6542C | J6C0T | 01 | cannot measure index size | A6675 | Prober | Wafer Edge/Mode Not Found | 2022/08/16 00:30:29 | 2022/08/16 00:47:14 | 0.28 | OK | 09108 | Remote處理 ----- 重抓Model OK (128.5) | NO | NO | NO | YES | YES | NO
EQPRepairJ22108447 | T6373-129 | CWC2J0Y28-01 | Tester Down | JB1LBS03 | NB01682F(COG48BA)-A-AT-Z-C | GFSW63.1-B1C1 | 20 | GPIB | A6675 | Tester | 離線(OCT/GPIB error) | 2022/08/16 00:38:11 | 2022/08/16 01:01:27 | 0.39 | OK | 10897 | 1.測試中Tester跳電，重開機後ok 2.此為STDF型號無法接測，請協助將#20 Hold To TPD. 3.已通知DN &產線協助後續作業. | NO | NO | NO | YES | YES | NO
EQPRepairJ22108452 | T6373-70 | CWC7J0M28-01.003 | Setup Fail | JB1Q4S17 | PA6542C | J683J | 01 | 0030 setup fail due to soktime correct z up limit over wait ae check (3d not ok) | A6441 | Setup Fail | O/S Fail | 2022/08/16 01:07:10 | 2022/08/16 01:39:44 | 0.54 | OK | 17091 | remote-----重新對針 二分法r:-4~6 試打針點Ok 請產線掛setup04 | NO | NO | NO | YES | YES | NO
EQPRepairJ22108459 | T6373-32 | CWC8J0562-01.001 | Other | KB1PNS01 | NA01761D(COF94AA)-A-Y-C | GGS330.1-BC2 | 07 | PM NG WAIT AE CHECK-A7048 | CPAUTO | Other | CHECK P/M | 2022/08/16 01:34:49 | 2022/08/16 01:54:54 | 0.33 | OK | 15492 | 1.二分法調針位(2，1)、RANGE:18 2.P/M(213，6)OK | NO | NO | NO | YES | YES | NO
EQPRepairJ22108469 | T6373-21 |  | Prober Down | JB1GKS08 | PA6543B | J6AC7 |  | AlarmCode:4526 AlarmMessage:LoadPort1 safety error! | CPAUTO | Prober | Loader Down | 2022/08/16 02:05:08 | 2022/08/16 02:26:26 | 0.36 | OK | 17091 | 退片不完全 手推回後INIT OK 請產線重LOAD#13 | NO | NO | NO | YES | YES | NO
EQPRepairJ22108470 | T6373-129 | CWC2J1067-01.006 | OverKill | JB1LBS03 | NB01682F(COG48BA)-A-C | GFTW28.1-BC21121 | 22 | TOO MANY OS | A6446 | Over Kill | O/S誤宰 | 2022/08/16 02:07:05 | 2022/08/16 02:32:01 | 0.42 | OK | 10897 | 針壓不足，重新對針，Range:9um，試打(139.6) | NO | NO | NO | YES | YES | NO
EQPRepairJ22108472 | T6373-90 | CWC7J0G36-01.001 | Setup Fail | JB1BYS11 | PA6587B_15U90H | J696A | 01 | (down temp) 2327 #1 set up passed but during run corr too many o/s, AE said change to MBD, id ok (72.4) dicho (-3,-8) range x(-32,26) y(2,-13), no corr wafer, change P/C JB1BYS12--11 due to fail bin 42 | A6943 | Setup Fail | BIN42 | 2022/08/16 02:08:34 | 2022/08/16 04:01:19 | 1.88 | OK | 17091 | 1.配件#12改機Fail全面Bin42 Function下卡磨針 2.上機相同，更換在庫卡#11上機Bin32 Open，清調針無效 3.配件#11、12皆送修無備卡配件當 | NO | NO | NO | YES | YES | NO
EQPRepairJ22108475 | T6373-29 | CWC8J0390-01.002 | OverKill | JB1QNS06 | NA01687F(COG48BA)-A-AT-C | GGS941.1-1BC22 | 01 | many fail bin.-A7048 | CPAUTO | Over Kill | Function誤宰 | 2022/08/16 02:27:36 | 2022/08/16 04:33:20 | 2.10 | OK | 17091 | 測試中大量bin13 function清調針無效點測不穩，更換配件04 上機全面bin13，拉下cal ABS轉掛Tester down | NO | NO | NO | YES | YES | NO
EQPRepairJ22108482 | T6373-32 | CWC8J0562-01.001 | OverKill | KB1PNS01 | NA01761D(COF94AA)-A-Y-C | GGS330.1-BC2 | 07 | always o/s pm not good.-A7048 | CPAUTO | Over Kill | O/D不足 | 2022/08/16 02:45:21 | 2022/08/16 03:02:46 | 0.29 | OK | 15492 | 1.溫針、重新對針，二分法(2，-5)、RANGE:20 2.P/M(198，7)OK | NO | NO | NO | YES | YES | NO
EQPRepairJ22108501 | T6373-90 |  | Probe Card Down | JB1BYS11 | PA6587B_15U90H | J696A |  | 1.配件#12改機Fail全面Bin42 Function下卡磨針 2.上機相同，更換在庫卡#11上機Bin32 Open，清調針無效 3.配件#11、12皆送修無備卡配件當 | 17091 | Probe Card | Function Fail | 2022/08/16 04:01:44 | 2022/08/16 06:34:57 | 2.55 | OK | 15492 | 1.配件#12送修回來，上機一樣全面Bin42 Function，無備卡配件當 | NO | NO | NO | YES | YES | NO
EQPRepairJ22108504 | T6373-121 | CWC7J0197-01.010 | OverKill | TB1INS25 | NA01404H-A-AT-C | GGR689.1-BC1212 | 22 | # 19-20 EDA SHOW 2H (NO CORR WAFER) | A6673 | Over Kill | SHOW H | 2022/08/16 04:04:45 | 2022/08/16 04:17:43 | 0.22 | OK | 17184 | 1.Bin51&53&54&55&56&58 不同機卡皆有此問題續測 | NO | NO | NO | NO | YES | NO
EQPRepairJ22108507 | T6373-47 | CWC8J0753-01.002 | OverKill | JB1HKS07 | NA01687F(COG48BA)-A-AT-C | GGV504.1-BC12 | 21 | ALREADY FOLLOW OVERKILL PROCEDURE BUT STILL NOT EFFECTIVE TOO MANY O/S | A6937 | Over Kill | O/D不足 | 2022/08/16 04:16:34 | 2022/08/16 04:56:53 | 0.67 | OK | 20182 | 1.清針 OD:10->15 pass 2.PM(37,0) | NO | NO | NO | YES | YES | NO
EQPRepairJ22108508 | T6373-16 | CWC7J0L18-01 | OverKill | VB405S07 | PA5208KF-LT | J68QT | 12 | ALREADY FOLLOW OVERKILL PROCEDURE BUT STILL NOT EFFECTIVE TOO MANY O/S | A6937 | Over Kill | O/S誤宰 | 2022/08/16 04:24:03 | 2022/08/16 04:42:45 | 0.31 | OK | 15492 | REMOTE--- 1.針位貼下，清針、二分法(0，14)、RANGE:12 2.P/M(147，5)OK | NO | NO | NO | YES | YES | NO
EQPRepairJ22108509 | T6373-29 |  | Tester Down | JB1QNS06 | NA01687F(COG48BA)-A-AT-C | GGS941.1-1BC22 |  | 測試中大量bin13 function清調針無效點測不穩，更換配件04 上機全面bin13，拉下cal ABS轉掛Tester down | 17091 | Over Kill | Function誤宰 | 2022/08/16 04:33:47 | 2022/08/16 06:13:49 | 1.67 | OK | 15492 | 1.測試中大量bin13 function清調針無效點測不穩，更換配件04 上機全面bin13， 2. TESTER CAL ABS ，上機點測穩定 PASS 3.P/M(146,4)OK | NO | NO | NO | YES | YES | NO
EQPRepairJ22108513 | T6373-38 | CWC8J0752-01.001 | OverKill | JB1GKS10 | PA6543B | J6AK6 | 05 | EDA SHOW I OPEN AND SHORT MODE TEST-A6857 | CPAUTO | Over Kill | SHOW I | 2022/08/16 04:57:46 | 2022/08/16 05:13:48 | 0.27 | OK | 15492 | 1.[!!! WARNING !!!OPEN AND SHORT MODE TEST] 2.續測觀察 | NO | NO | NO | YES | YES | NO
EQPRepairJ22108526 | T6373-32 | CWC8J0562-01.004 | OverKill | KB1PNS01 | NA01761D(COF94AA)-A-Y-C | GGS330.1-BC1121 | 21 | p/m not good o/s.-A7048 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/16 05:59:06 | 2022/08/16 06:28:54 | 0.50 | OK | 17091 | 清針後OD+5UM 二分法R:-16~4 PM(198，7) | NO | NO | NO | YES | YES | NO
EQPRepairJ22108533 | T6373-111 | CWC8J0440-01 | OverKill | JB1Q4S23 | PA6542C | J6C0T | 03 | eda show 2h no run corrr-A6446 | CPAUTO | Over Kill | EDA SHOW 2H | 2022/08/16 06:05:38 | 2022/08/16 07:03:31 | 0.96 | OK | 10897 | 機台清針，測試一段OK續觀察 | NO | NO | NO | NO | YES | NO
EQPRepairJ22108535 | T6373-38 | CWC8J0752-01.001 | OverKill | JB1GKS10 | PA6543B | J6AK6 | 06 | TOO MANY OPEN/SHORT | A6599 | Over Kill | O/S誤宰 | 2022/08/16 06:11:15 | 2022/08/16 06:45:05 | 0.56 | OK | 17091 | 機台清針後重新對針 二分法R:-8~2 PM(158，6) | NO | NO | NO | YES | YES | NO
EQPRepairJ22108539 | T6373-86 | CWC8J0217-01 | OverKill | JB1Q4S15 | PA6542C | J68S2 | 11 | too many o/s | A6941 | Over Kill | O/S誤宰 | 2022/08/16 06:30:42 | 2022/08/16 07:00:51 | 0.50 | OK | 09108 | 1.O/S 針位貼右 調針PASS range:14um (43.5) | NO | NO | NO | YES | YES | NO
EQPRepairJ22108540 | T6373-90 |  | Probe Card Down | JB1BYS12 | PA6587B_15U90H | J62W7 |  | 1.配件#12協助上機驗證，上機一樣全面Bin42 Function，無備卡配件當 | 15492 | Other | 配件down改測 | 2022/08/16 06:36:39 | 2022/08/16 08:07:11 | 1.51 | OK | 15362 | 配件down改測PA5137F-LT1(TB1R8S82),#02 y range:8,p/m(157,4) | NO | NO | NO | YES | YES | NO
EQPRepairJ22108547 | T6373-91 |  | Setup Fail | JB1FDS07 | PA1558W-10A_NC | RDDXB1000 |  | setup fail tester alarm output leakage-A6118 | CPAUTO | Tester | Other | 2022/08/16 06:57:34 | 2022/08/16 09:53:02 | 2.92 | OK | 07764 | 1.改機PA1558W-10A_NC 點測RTE 秀 Hvdc Guard Alarm occurred. (Pin: 1072 DUT: 1) 2.TESTER REBOOT無效 , 只有一張配件 3.貨與卡拉至ND3-12 驗證PASS 4.故掛TESTER DOWN | NO | NO | NO | YES | YES | NO
EQPRepairJ22108550 | T6373-16 |  | Prober Down | VB405S07 | PA5208KF-LT | J68QT |  | Loader error | A6896 | Prober | Loader Down | 2022/08/16 07:08:41 | 2022/08/16 07:21:05 | 0.21 | OK | 20075 | loader init ok | NO | NO | NO | YES | YES | NO
EQPRepairJ22108564 | T6373-47 | CWC8J0753-01.003 | OverKill | JB1HKS07 | NA01687F(COG48BA)-A-AT-C | GGV504.1-BC112 | 22 | TOO MANY O/S-A6131 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/16 08:01:14 | 2022/08/16 08:30:06 | 0.48 | OK | 10210 | 重新對針(146.4) | NO | NO | NO | YES | YES | NO
EQPRepairJ22108573 | T6373-111 | CWC8J0440-01 | OverKill | JB1Q4S23 | PA6542C | J6C0T | 03 | too many fail bin6 | A6858 | Over Kill | Function誤宰 | 2022/08/16 08:29:35 | 2022/08/16 09:26:36 | 0.95 | OK | 14291 | bin6 磨針pass P/M(-1.2)RG:10 | NO | NO | NO | YES | YES | NO
EQPRepairJ22108576 | T6373-41 | CWC8J0809-01 | OverKill | JB8KZS16 | ND01657B(COF94AA)-A-C | S0VP70.00-BC | 14 | too many o/s | A6873 | Over Kill | O/S誤宰 | 2022/08/16 08:56:17 | 2022/08/16 09:01:17 | 0.08 | OK | 10210 | 調整針位(216.9) | NO | NO | NO | YES | YES | NO
EQPRepairJ22108577 | T6373-34 |  | Prober Down | JB1BYS10 | PA6587B_15U90H | J696F |  | AlarmCode:1058 AlarmMessage:MODEL NOT FOUND ( FINE ALIGNMENT ) !! | CPAUTO | Prober | Wafer Mode Not Found | 2022/08/16 08:57:47 | 2022/08/16 09:10:48 | 0.22 | OK | 15362 | 重新抓取,check die ok | NO | NO | NO | YES | YES | NO
EQPRepairJ22108585 | T6373-42 | CWC8J0508-01.004 | OverKill | JB1PNS07 | NA01761D(COF94AA)-A-Y-C | GGS421.1-BC112 | 14 | TOO MANY O/S | A6131 | Over Kill | O/D不足 | 2022/08/16 09:33:34 | 2022/08/16 09:51:42 | 0.30 | OK | 20075 | 重新對針range 10,mark -1,4 | NO | NO | NO | YES | YES | NO
EQPRepairJ22108588 | T6373-91 |  | Tester Down |  |  |  |  | 1.改機PA1558W-10A_NC 點測RTE 秀 Hvdc Guard Alarm occurred. (Pin: 1072 DUT: 1) 2.TESTER REBOOT無效 , 只有一張配件 3.貨與卡拉至ND3-12 驗證PASS 4.故掛TESTER DOWN | 07764 | Tester | Other | 2022/08/16 09:53:38 | 2022/08/16 13:11:43 | 3.30 | OK | 07764 | 型號：PA1558W-10A_NC / CP2 Tester Down--- 1.改機點測RTE秀Hvdc Guard Alarm occurred. (Pin: 1072 DUT: 1)。 2.把貨與卡拉至ND3-12驗證PASS ,改掛TESTER DOWN 3.LCD#05 SWAP LCD#04 / HVDC#05 SWAP HVDC#06 ,執行abs pass ,上機驗證FAIL轉移至HVDC#06 4.更換備品HVDC#06 上機abs pass，點測pass。 5.壞品放置備品室，請知悉 | NO | NO | NO | YES | YES | YES
EQPRepairJ22108593 | T6373-39 | CWC7J0H50-01 | Prober Down | TB1R8S78 | PA5137F-LT1 | P097F.1 | 13 | AlarmCode:1995 AlarmMessage:Faied to upload log file | CPAUTO | Prober | Faied to upload log file | 2022/08/16 10:06:50 | 2022/08/16 10:11:36 | 0.08 | OK | 10210 | 消音後續測 | NO | NO | NO | YES | YES | NO
EQPRepairJ22108600 | T6373-121 | CWC7J0H70-01.004 | OverKill | TB1INS29 | NA01404H-A-AT-C | GGT500.1-B1C21211 | 07 | o/s-11316 | CPAUTO | Over Kill | O/S | 2022/08/16 10:42:06 | 2022/08/16 10:57:08 | 0.25 | OK | 13222 | O/S 二分法RANGE=14 (25.5) | NO | NO | NO | YES | YES | NO
EQPRepairJ22108601 | T6373-13 | CWC8J0119-01.001 | OverKill | JB1GVS19 | NC01682A(COG48AA)-A-AT-N3-C | GGP620.1-2BC21 | 09 | #09 已依誤宰處理流程處置無效 (O/S過多) | 05131 | Over Kill | O/S誤宰 | 2022/08/16 10:47:41 | 2022/08/16 11:23:20 | 0.59 | OK | 20230 | 1.O/S 2.機台清針,二分法調針RG=11,PM(168.5) 3.PASS | NO | NO | NO | YES | YES | NO
EQPRepairJ22108608 | T6373-121 | CWC7J0H70-01.004 | OverKill | TB1INS29 | NA01404H-A-AT-C | GGT500.1-B1C21211 | 07 | too many o/s-A6873 | CPAUTO | Over Kill | O/S | 2022/08/16 11:15:46 | 2022/08/16 11:40:34 | 0.41 | OK | 20230 | 1.O/S 2.機台清針,二分法調針YRG=15,XRG=36,PM(55.6) 3.PASS | NO | NO | NO | YES | YES | NO
EQPRepairJ22108609 | T6373-152 |  | Tester Down | JB252S02 | TD7850-A0S-AZ-CP | SC1GT-B1121C |  | PLS CHECK TESTER BARCODE NG | A6132 | Tester | 鍵盤異常 | 2022/08/16 11:16:33 | 2022/08/16 11:58:25 | 0.70 | OK | 20075 | 插拔USB OK | NO | NO | NO | YES | YES | NO
EQPRepairJ22108618 | T6373-47 | CWC8J0753-01.003 | OverKill | JB1HKS07 | NA01687F(COG48BA)-A-AT-C | GGV504.1-BC112 | 24 | #24 pls check and confirm pm too many o/s-A6132 | CPAUTO | Over Kill | 機台清針後Pass | 2022/08/16 11:40:49 | 2022/08/16 11:59:15 | 0.31 | OK | 20075 | 機台清針後Pass,RANGE 10 | NO | NO | NO | YES | YES | NO
EQPRepairJ22108625 | T6373-32 | CWC8J0515-01.003 | OverKill | KB1PNS01 | NA01761D(COF94AA)-A-Y-C | GGS331.1-BC1221 | 18 | too many o/s-A6920 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/16 12:24:54 | 2022/08/16 13:01:07 | 0.60 | OK | 13581 | 1.使用機台清針+調針位pass(199.7) 2.二分法(109.4)Range:18um | NO | NO | NO | NO | YES | NO
EQPRepairJ22108626 | T6373-29 | CWC8J0390-01.002 | OverKill | JB1QNS04 | NA01687F(COG48BA)-A-AT-C | GGS941.1-1BC22 | 07 | #5.6 EDA 2H NO CORR.-A6920 | CPAUTO | Over Kill | 2H | 2022/08/16 12:30:06 | 2022/08/16 12:46:49 | 0.28 | OK | 15362 | 2H,續測 | NO | NO | NO | YES | YES | NO
EQPRepairJ22108639 | T6373-33 | CWC7J0676-01 | OverKill | JB1TAS26 | PA5208KF2 | J68G9 | 16 | o/s-11316 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/16 13:55:11 | 2022/08/16 14:07:28 | 0.20 | OK | 13222 | REMOTE---- 1. 機台清針 測試一段OK | NO | NO | NO | YES | YES | NO
EQPRepairJ22108642 | T6373-32 | CWC8J0515-01.003 | OverKill | KB1PNS01 | NA01761D(COF94AA)-A-Y-C | GGS331.1-BC1221 | 18 | REPROBING STILL O/S-A6564 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/16 14:03:34 | 2022/08/16 14:31:22 | 0.46 | OK | 15362 | 清針,二分法x/y range:24/22,p/m(50,8) | NO | NO | NO | YES | YES | NO
EQPRepairJ22108644 | T6373-38 |  | Setup Fail | JB1GKS09 | PA6543B | J6AC7 |  | setup pssed (87.4) dichotomy (0,-9) range:x0 y(-3,-15) run corr fresh wafer t32 change pc JB1GKS12 not release to machine 3-38,JB1GKS09 polish-A6143 | CPAUTO | Probe Card | bin2 open | 2022/08/16 14:14:15 | 2022/08/16 15:02:03 | 0.80 | OK | 15362 | 1.JB1GKS09 set up fail BIN2 OPEN固定PIN,清調針,換PIN無效,下卡磨針,上機一樣固定PIN 2.更換未Release配件 #12 上機 PASS，故配件#09送修,無備卡配件down | NO | NO | NO | YES | YES | NO
EQPRepairJ22108645 | T6373-12 |  | Prober Down | JB6H1S05 | NC01690F(COG46BA)-A-AT-C | GGY201.1-B12C122 |  | AlarmCode:1057 AlarmMessage:MODEL NOT FOUND ( ROUGH ALIGNMENT ) !! | CPAUTO | Prober | Wafer Edge/Mode Not Found | 2022/08/16 14:14:48 | 2022/08/16 14:24:54 | 0.17 | OK | 14291 | REMOTE處理-----重抓model and check die/針位 ok P/M(23.4) | NO | NO | NO | YES | YES | NO
EQPRepairJ22108652 | T6373-35 | CWC8J0622-01 | OverKill | JB1GVS24 | NC01682A(COG48AA)-A-AT-N2-C | GGT630.1-BC11 | 01 | p/c particle o/s bin-05131 | CPAUTO | Over Kill | O/S | 2022/08/16 14:39:26 | 2022/08/16 15:34:39 | 0.92 | OK | 13222 | 重新對針 RANGE=12 (18.5) | NO | NO | NO | YES | YES | NO
EQPRepairJ22108654 | T6373-34 | CWC7J0K25-01.001 | OverKill | JB1BYS10 | PA6587B_15U90H | J696F | 05 | EDA SHOW 2 R O/S-A6564 | CPAUTO | Over Kill | 2R | 2022/08/16 14:46:44 | 2022/08/16 15:27:11 | 0.67 | OK | 13222 | 針壓不足 重新對針 RANGE=12 (160.6) | NO | NO | NO | YES | YES | NO
EQPRepairJ22108654 | T6373-111 | CWC8J0440-01 | OverKill | JB1Q4S23 | PA6542C | J6C0T | 05 | BIN 6 TOO MANY-A6873 | CPAUTO | Over Kill | Function誤宰 | 2022/08/16 14:46:44 | 2022/08/16 15:11:45 | 0.42 | OK | 14291 | 換卡s14 pass P/M(43.5) | NO | NO | NO | YES | YES | NO
EQPRepairJ22108659 | T6373-38 |  | Probe Card Down | JB1GKS10 | PA6543B | J6AK6 |  | 1.JB1GKS09 set up fail BIN2 OPEN固定PIN,清調針,換PIN無效,下卡磨針,上機一樣固定PIN 2.更換未Release配件 #12 上機 PASS，故配件#09送修,無備卡配件down | 15362 | Probe Card | P/C DOWN | 2022/08/16 15:02:32 | 2022/08/16 16:18:21 | 1.26 | OK | 13222 | P/C DOWN--- 1. 配件送修回來上機PASS RANGE=10 (162.1) 2. 還機SETUP04 | NO | NO | NO | YES | YES | NO
EQPRepairJ22108668 | T6373-147 | CWC8J0618-01.004 | OverKill | JB1Q4S22 | PA6542C | J6C0W | 07 | o/s-07854 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/16 15:58:23 | 2022/08/16 16:20:16 | 0.36 | OK | 20075 | 1.程式當,手動RE-LOAD PROGRAM 2.清調針PASS,重新上貨 | NO | NO | NO | YES | YES | NO
EQPRepairJ22108672 | T6373-121 | CWC7J0H70-01 | OverKill | TB1INS29 | NA01404H-A-AT-C | GGT500.1-B1C1 | 01 | open/short PC with particles-A6858 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/16 16:13:18 | 2022/08/16 16:35:53 | 0.38 | OK | 07764 | open 清針 | NO | NO | NO | YES | YES | NO
EQPRepairJ22108679 | T6373-99 | CWC7J0L66-01 | OverKill | JB1TAS16 | PA5208K-LT | QT800 | 01 | o/s-11550 | CPAUTO | Over Kill | 機台清針後Pass | 2022/08/16 16:40:44 | 2022/08/16 17:35:23 | 0.91 | OK | 20075 | 機台清針後Pass | NO | NO | NO | YES | YES | NO
EQPRepairJ22108687 | T6373-16 | CWC7J0K83-01.001 | Prober Down | VB405S07 | PA5208KF-LT | J67HF | 03 | AlarmCode:6188 AlarmMessage:NO CLEANING SPACE !! | CPAUTO | Prober | NO CLEAN SPACE | 2022/08/16 17:30:23 | 2022/08/16 17:54:35 | 0.40 | OK | 10210 | 先關閉清針，待PM TEAM確認 | NO | NO | NO | YES | YES | NO
EQPRepairJ22108693 | T6373-121 | CWC7J0H70-01.004 | OverKill | TB1INS29 | NA01404H-A-AT-C | GGT500.1-B1C21211 | 07 | p/c particle o/s-05131 | CPAUTO | Over Kill | O/S | 2022/08/16 17:53:40 | 2022/08/16 18:16:28 | 0.38 | OK | 20230 | 1.O/S 2.機台清針,二分法調針YRG=17,XRG=38,PM(53.6) 3.PASS | NO | NO | NO | YES | YES | NO
EQPRepairJ22108696 | T6373-47 | CWC8J0124-01.001 | Corr Fail | JB1HKS07 | NA01687G(COG48BA)-A-AT-N2-C | GGS268.1-B2 | 14 | Corr Fail | 04950 | Corr | Other | 2022/08/16 18:01:42 | 2022/08/16 18:04:10 | 0.04 | OK | 10210 | 手動比對Pass | NO | NO | NO | YES | YES | NO
EQPRepairJ22108704 | T6373-16 | CWC7J0K83-01.001 | OverKill | VB405S07 | PA5208KF-LT | J67HF | 04 | #3 AE USE MACHINE-A6132 | CPAUTO | Prober | 設定3D清針 | 2022/08/16 18:22:54 | 2022/08/16 18:48:13 | 0.42 | OK | 10210 | 1.確認3D清針X Y為0 2.轉Prober down待PM TEAM確認 | NO | NO | NO | YES | YES | NO
EQPRepairJ22108710 | T6373-47 | CWC8J0348-01.016 | Prober Down | JB1HKS07 | NA01687G(COG48BA)-A-AT-N2-C | GGR356.1-B11C11122 | 21 | AlarmCode:1057 AlarmMessage:MODEL NOT FOUND ( ROUGH ALIGNMENT ) !! | CPAUTO | Prober | Wafer Edge/Mode Not Found | 2022/08/16 18:45:00 | 2022/08/16 18:57:41 | 0.21 | OK | 10210 | 重抓MODE OK | NO | NO | NO | YES | YES | NO
EQPRepairJ22108712 | T6373-16 |  | Prober Down | VB405S07 | PA5208KF-LT | J67HF |  | 1.確認3D清針X Y為0 2.轉Prober down待PM TEAM確認 | 10210 | Over Kill | O/S誤宰 | 2022/08/16 18:48:37 | 2022/08/16 20:00:21 | 1.20 | OK | 20277 | 輪工處理---- 1.請PM Team設定3D Clean參數OK 2.Setup PASS，Range:X>20，Y:16 | NO | NO | NO | YES | YES | NO
EQPRepairJ22108718 | T6373-86 |  | Setup Fail | JB1Q4S15 | PA6542C | J68S1 |  | Dew point temperature control failure. | A6938 | Prober | 溫度異常 (Temperature Error) | 2022/08/16 19:33:34 | 2022/08/16 19:51:27 | 0.30 | OK | 10897 | 改機降溫異常，Prober關機重開正常降溫 | NO | NO | NO | NO | YES | NO
EQPRepairJ22108726 | T6373-33 | CWC7J0676-01 | OverKill | JB1TAS26 | PA5208KF2 | J68G9 | 19 | too many O/S-A6857 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/16 20:22:49 | 2022/08/16 20:40:55 | 0.30 | OK | 10897 | 換片O/S過多，機台清針，針位確認，測試一段穩定，試打(44.0) | NO | NO | NO | YES | YES | NO
EQPRepairJ22108728 | T6373-47 | CWC8J0134-01.003 | Prober Down | JB1HKS07 | NA01687G(COG48BA)-A-AT-N2-C | GGS267.1-B11C12 | 06 | model not found-A6512 | CPAUTO | Prober | Wafer Edge/Mode Not Found | 2022/08/16 20:29:21 | 2022/08/16 20:39:50 | 0.17 | OK | 20277 | Prober Down----重抓Model OK(model not found) | NO | NO | NO | YES | YES | NO
EQPRepairJ22108729 | T6373-147 | CWC8J0618-01.004 | OverKill | JB1Q4S22 | PA6542C | J6C0W | 08 | Already follow Overkilll handling procedures but still not effective too many o/s | A7165 | Over Kill | O/S誤宰 | 2022/08/16 20:33:10 | 2022/08/16 20:51:47 | 0.31 | OK | 20182 | 1.清針調針PASS RG-10 PM(158.4) | NO | NO | NO | YES | YES | NO
EQPRepairJ22108732 | T6373-152 | CWC8J0455-01.002 | OverKill | JB252S02 | TD7850-A0S-AZ-CP | SC1GS-B1C | 11 | too many os-A6512 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/16 20:52:39 | 2022/08/16 21:14:47 | 0.37 | OK | 20182 | 1.清針，重對針PASS RG-8 PM(147.7) | NO | NO | YES | YES | YES | NO
EQPRepairJ22108745 | T6373-99 | CWC7J0L66-01 | OverKill | JB1TAS16 | PA5208K-LT | QT800 | 03 | TOO MANY O/S PM UP AND DOWN | A7023 | Over Kill | O/S誤宰 | 2022/08/16 21:34:07 | 2022/08/16 22:05:32 | 0.52 | OK | 20277 | Remote---- 1.Range:16 機台清針 PASS | NO | NO | NO | YES | YES | NO
EQPRepairJ22108753 | T6373-42 | CWC8J0855-01 | Other | JB8Q6S05 | NB01761A(COF94AA)-A-C | H2NF20.1-B1C | 01 | #1 Please check and confirm PM XY : 103.5 U/D | A6896 | Other | 針位 | 2022/08/16 21:55:51 | 2022/08/16 22:18:56 | 0.38 | OK | 20277 | Remote---- 1.Range:X:14，Y:20 試打(218.5)，確認針位、針壓無異常，還機生產 | NO | NO | NO | YES | YES | NO
EQPRepairJ22108756 | T6373-86 |  | Setup Fail | JB1Q4S15 | PA6542C | J68S2 |  | setup fail cleaning wafer not found.-A7048 | CPAUTO | Prober | Other | 2022/08/16 22:06:38 | 2022/08/16 22:30:53 | 0.40 | OK | 10897 | 將clean wafer關閉OK | NO | NO | NO | NO | YES | NO
EQPRepairJ22108760 | T6373-47 | CWC8J0348-01.012 | OverKill | JB1HKS07 | NA01687G(COG48BA)-A-AT-N2-C | GGR356.1-B11C12121 | 12 | #12 Already follow the overkilling procedure but still XY : 37.4 OD : 15 | A6896 | Over Kill | O/S誤宰 | 2022/08/16 22:28:59 | 2022/08/16 22:51:54 | 0.38 | OK | 20277 | Remote---- 1.Range:14 機台清針 PASS | NO | NO | NO | YES | YES | NO
EQPRepairJ22108768 | T6373-29 | CWC8J0753-01.004 | OverKill | JB1QNS04 | NA01687F(COG48BA)-A-AT-C | GGV504.1-BC22 | 09 | TOO MANY FAIL BIN 13 CONTINUOUS EDA SHOW 2H NO CORR | A6675 | Over Kill | Function誤宰 | 2022/08/16 22:51:25 | 2022/08/16 23:20:43 | 0.49 | OK | 10110 | 1.Bin13 Function不穩，已發生多次 2.待換片再執行cal機，先續測 | NO | NO | NO | YES | YES | NO
EQPRepairJ22108789 | T6373-16 | CWC7J0L58-01 | OverKill | VB405S07 | PA5208K-LT | QT7CH | 05 | too many os-A6512 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/17 00:19:27 | 2022/08/17 00:33:22 | 0.23 | OK | 20182 | 1.清針調針PASS RG-10 PM(147.5) | NO | NO | NO | YES | YES | NO
EQPRepairJ22108791 | T6373-29 | CWC8J0753-01.004 | OverKill | JB1QNS04 | NA01687F(COG48BA)-A-AT-C | GGV504.1-BC22 | 09 | too many failbin 13-A6936 | CPAUTO | Tester | BIN13 | 2022/08/17 00:26:56 | 2022/08/17 01:15:02 | 0.80 | OK | 17091 | 1.測試會有大量Bin13 Function問題，更換多張配件無效 2.Cal ABS會好但是無法久測，近期已發生多次 3.目前無方向維修，請協助改測，謝謝 | NO | NO | NO | YES | YES | NO
EQPRepairJ22108801 | T6373-21 |  | Prober Down | JB1GKS08 | PA6543B | J6AC8 |  | AlarmCode:4526 AlarmMessage:LoadPort1 safety error! | CPAUTO | Prober | Loader Down | 2022/08/17 01:01:19 | 2022/08/17 01:16:56 | 0.26 | OK | 20277 | 輪工處理---- 1.Loader init OK | NO | NO | NO | NO | YES | NO
EQPRepairJ22108807 | T6373-29 |  | Tester Down | JB1QNS04 | NA01687F(COG48BA)-A-AT-C | GGV504.1-BC22 |  | 1.測試會有大量Bin13 Function問題，更換多張配件無效 2.Cal ABS會好但是無法久測，近期已發生多次 3.目前無方向維修，請協助改測，謝謝 | 17091 | Tester | Bin13 Function | 2022/08/17 01:15:55 | 2022/08/17 06:10:21 | 4.91 | OK | 10110 | Tester Down--- 1.測試會有大量Bin13 Function問題，更換多張配件無效 2.Cal ABS會好但是無法久測，近期已發生多次 3.目前無方向維修，與產線協調改測 4.接手重新setup pass | NO | NO | NO | YES | YES | NO
EQPRepairJ22108810 | T6373-121 | CWC6J0880-01.009 | OverKill | TB1INS29 | NA01404H-A-AT-C | G9N306.1-2BC12212 | 23 | CONTINUOUS EDA SHOW 2H NO CORR WAFER | A6675 | Over Kill | Show H | 2022/08/17 01:20:54 | 2022/08/17 01:36:53 | 0.27 | OK | 09108 | 1.EDA Show H 續測 | NO | NO | NO | YES | YES | NO
EQPRepairJ22108822 | T6373-121 | CWC6J0880-01.009 | OverKill | TB1INS29 | NA01404H-A-AT-C | G9N306.1-2BC12212 | 25 | #22,23 (2h)(no corr)-03419 | CPAUTO | Over Kill | Show H | 2022/08/17 02:18:43 | 2022/08/17 02:23:40 | 0.08 | OK | 09108 | Remote處理 ----- 1.eda Show H 已有掛down紀錄 請產線續測 還機生產 | NO | NO | NO | YES | YES | NO
EQPRepairJ22108828 | T6373-86 |  | Prober Down | JB1Q4S15 | PA6542C | J68S2 |  | AlarmCode:1060 AlarmMessage:CANNOT MEASURE INDEX SIZE !! | CPAUTO | Prober | Wafer Edge/Mode Not Found | 2022/08/17 02:28:53 | 2022/08/17 02:35:39 | 0.11 | OK | 10897 | Mode Not Found/check die/針痕ok | NO | NO | NO | NO | YES | NO
EQPRepairJ22108833 | T6373-33 | CWC7J0676-01 | OverKill | JB1TAS26 | PA5208KF2 | J68G9 | 21 | o/s-03491 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/17 02:47:22 | 2022/08/17 02:59:36 | 0.20 | OK | 09108 | 1.O/S 清針 調針PASS range:14um (46.6) | NO | NO | NO | YES | YES | NO
EQPRepairJ22108840 | T6373-86 | CWC8J0217-01.001 | OverKill | JB1Q4S15 | PA6542C | J68S2 | 02 | many os-A7048 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/17 03:21:22 | 2022/08/17 03:43:50 | 0.37 | OK | 09108 | 1.o/s 清針 調針Pass range:12um (128.5) | NO | NO | NO | YES | YES | NO
EQPRepairJ22108852 | T6373-147 | CWC8J0618-01.004 | OverKill | JB1Q4S22 | PA6542C | J6C0W | 10 | Already follow Overkilll handling procedures but still not effective too many o/s | A7165 | Over Kill | O/S誤宰 | 2022/08/17 05:18:59 | 2022/08/17 06:01:46 | 0.71 | OK | 20182 | 1.清針 重對針PASS 2.RG-10 PM(13.4) | NO | NO | NO | YES | YES | NO
EQPRepairJ22108862 | T6373-15 | CWC8J0856-01 | OverKill | JB6H1S04 | NC01690I(COG46BA)-A-AT-C | GGS959.1-BC | 23 | EDA SHOW "I" NO CORR DUTTDIFF | A6675 | Over Kill | SHOW I | 2022/08/17 06:15:25 | 2022/08/17 06:24:50 | 0.16 | OK | 20182 | 單片秀I 前片無此問題，確認針痕PM(127.5) 續測 | NO | NO | NO | YES | YES | NO
EQPRepairJ22108879 | T6373-129 |  | Setup Fail | JB1QNS06 | NA01687G(COG48BA)-A-AT-N2-C | GGR356.1-B11 |  | too large needle --wait engr repair-A6942 | CPAUTO | Setup Fail | 對針異常 | 2022/08/17 07:42:54 | 2022/08/17 07:55:04 | 0.20 | OK | 20230 | 1.對針異常 2.重新對針ok | NO | NO | NO | YES | YES | NO
EQPRepairJ22108881 | T6373-21 | CWC8J0621-01.001 | OverKill | JB1GKS08 | PA6543B | J6AC8 | 16 | too many o/s | A6873 | Over Kill | O/S誤宰 | 2022/08/17 07:45:34 | 2022/08/17 07:55:39 | 0.17 | OK | 15362 | 清調針PASS,P/M(159,6) | NO | NO | NO | YES | YES | NO
EQPRepairJ22108884 | T6373-43 | CWC7J0M33-01.001 | Setup Fail | VB1BVS01 | PA5270D-LT | U38F8.00 | 15 | 07:20 #15 id/ passed. setup fail during run corr. Ae check confirm to MBD,use 1st wafer (76.3)dicho(0.-3) range x(0.0) y(-1.-10)preheat | A7086 | Setup Fail | 起測一顆PASS一顆OPEN | 2022/08/17 07:56:09 | 2022/08/17 08:09:49 | 0.23 | OK | 10210 | 1.Setup點測Pass，起測一顆PASS一顆OPEN。 2.開啟Double Contact驗證會Pass。 3.此機轉掛產品問題，請產品協助確認 | NO | NO | NO | YES | YES | NO
EQPRepairJ22108887 | T6373-16 | CWC8J0815-01 | OverKill | VB405S07 | PA5208K-LT | QT7YT | 02 | #2 PLS CHECK AND CONFIRM PM TOO MANY O/S-A6132 | CPAUTO | Over Kill | 配件有異物 | 2022/08/17 08:06:46 | 2022/08/17 08:48:37 | 0.70 | OK | 20075 | 換卡#S09,range 12 | NO | NO | NO | YES | YES | NO
EQPRepairJ22108889 | T6373-43 |  | Product Issue |  |  |  |  | 1.Setup點測Pass，起測一顆PASS一顆OPEN。 2.開啟Double Contact驗證會Pass。 3.此機轉掛產品問題，請產品協助確認 | 10210 | Product Issue | 一顆Pass一顆OPEN | 2022/08/17 08:10:16 | 2022/08/17 11:23:19 | 3.22 | OK | 10210 | 產品通知先借機，此機先IDLE | NO | NO | NO | YES | YES | NO
EQPRepairJ22108900 | T6373-129 |  | Setup Fail | JB1QNS06 | NA01687G(COG48BA)-A-AT-N2-C | GGR356.1-B11 |  | ARM 1VACUUM NG-A6942 | CPAUTO | Setup Fail | Prober異常 prober abnormal | 2022/08/17 08:49:31 | 2022/08/17 09:47:13 | 0.96 | OK | 14291 | 1.改機run corr時,機台自動清針後續造成arm1 vacuum error 2.Prober reset ok,查看結霜,升高溫烘烤 3.修改參數21515:35秒>60秒 4.T44 setup again (低溫型號) | NO | NO | NO | YES | YES | NO
EQPRepairJ22108919 | T6373-42 | CWC8J0855-01 | OverKill | JB8Q6S05 | NB01761A(COF94AA)-A-C | H2NF20.1-B1C | 09 | too many o/s, p/m Not Good.. please check | A7086 | Over Kill | 針位調整 | 2022/08/17 09:50:47 | 2022/08/17 10:11:48 | 0.35 | OK | 20075 | 二分法range 12,退內壓20um 跑soak time | NO | NO | NO | YES | YES | NO
EQPRepairJ22108923 | T6373-33 | CWC7J0676-01 | OverKill | JB1TAS26 | PA5208KF2 | J68G9 | 25 | #23.24 2R o/s-08889 | CPAUTO | Over Kill | 2R | 2022/08/17 10:05:13 | 2022/08/17 10:53:28 | 0.80 | OK | 20230 | 1.2R 2.更換備卡S28,RG=6,PM(47.6) 3.PASS | NO | NO | NO | YES | YES | NO
EQPRepairJ22108929 | T6373-39 | CWC7J0527-01 | OverKill | TB1R8S78 | PA5137F-LT1 | P07UM.1 | 09 | O/S-07854 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/17 10:25:59 | 2022/08/17 10:54:15 | 0.47 | OK | 10210 | Unit清針 | NO | NO | NO | YES | YES | NO
EQPRepairJ22108933 | T6373-42 | CWC8J0855-01 | OverKill | JB8Q6S05 | NB01761A(COF94AA)-A-C | H2NF20.1-B1C | 09 | TOO MANY O/S-A6131 | CPAUTO | Over Kill | 針位調整 | 2022/08/17 10:39:36 | 2022/08/17 11:43:10 | 1.06 | OK | 20075 | 針位調整,RANGE 12 | NO | NO | NO | YES | YES | NO
EQPRepairJ22108941 | T6373-01 | CWC7J0J00-01 | OverKill | JB1TAS23 | PA5208K2 | QT7WR | 15 | too many O/S | A6919 | Over Kill | O/S誤宰 | 2022/08/17 11:31:12 | 2022/08/17 13:09:17 | 1.63 | OK | 15362 | 原配件#23下卡磨針,配件室告知共面不佳,送修,更換配件#25 pass,y range:12,p/m(148,5) | NO | NO | NO | YES | YES | NO
EQPRepairJ22108944 | T6373-16 | CWC8J0815-01 | OverKill | VB405S09 | PA5208K-LT | QT7YT | 07 | #7 PLS CHECK AND CONFIRM PM TOO MANY O/S-A6132 | CPAUTO | Over Kill | 針位調整 | 2022/08/17 12:06:30 | 2022/08/17 12:21:46 | 0.25 | OK | 20075 | 針位調整,range 12 | NO | NO | NO | YES | YES | NO
EQPRepairJ22108962 | T6373-42 | CWC8J0855-01 | OverKill | JB8Q6S05 | NB01761A(COF94AA)-A-C | H2NF20.1-B1C | 09 | #9 PLS CHECK AND CONFIRM PM TOO MANY O/S-A6132 | CPAUTO | Probe Card | 熔針 | 2022/08/17 13:22:18 | 2022/08/17 15:01:45 | 1.66 | OK | 20075 | 1.測試中O/S不穩,下卡磨針。 2.配件室告知熔針,故將P/C送修,掛P/C DOWN | NO | NO | NO | YES | YES | NO
EQPRepairJ22108976 | T6373-147 | CWC8J0618-01.004 | OverKill | JB1Q4S22 | PA6542C | J6C0W | 13 | bin 15-07854 | CPAUTO | Over Kill | Function誤宰 | 2022/08/17 14:21:48 | 2022/08/17 14:55:13 | 0.56 | OK | 17229 | 1.依產品指示更換配件pass,原配件bin15 function送修,預溫重對針二分法range:19um落點(-1.3) | NO | NO | NO | YES | YES | NO
EQPRepairJ22108980 | T6373-13 | CWC8J0871-01 | OverKill | JB6K8S02 | NC01682B(COG48AA)-A-AT-N-C | GGR437.1-BC | 01 | TOO MANY OPEN SHORT-a6133 | CPAUTO | Over Kill | O/S | 2022/08/17 14:37:57 | 2022/08/17 15:44:42 | 1.11 | OK | 20230 | 1.O/S 2.更換備卡S03,RG=16,PM(45.6) 3.PASSS | NO | NO | NO | YES | YES | NO
EQPRepairJ22108986 | T6373-42 |  | Probe Card Down | JB8Q6S05 | NB01761A(COF94AA)-A-C | H2NF20.1-B1C |  | 1.測試中O/S不穩,下卡磨針。 2.配件室告知熔針,故將P/C送修,掛P/C DOWN | 20075 | Probe Card | 熔針 | 2022/08/17 15:03:01 | 2022/08/17 16:27:56 | 1.42 | OK | 20075 | 配件已外修，請產線協助改測或切IDLE | NO | NO | NO | YES | YES | NO
EQPRepairJ22108987 | T6373-129 | CWC8J0390-01.001 | Prober Down | JB1QNS06 | NA01687F(COG48BA)-A-AT-C | GGS941.1-1BC211 | 04 | AlarmCode:4827 AlarmMessage:ARM @@VACUUM ERROR !! | CPAUTO | Prober | VACUUM Error | 2022/08/17 15:06:59 | 2022/08/17 18:23:29 | 3.28 | OK | 20230 | 1.Slot #04測試中機台自動使用Clean Wafer清針時，發生Arm1 vacuum error，Prober 強制重開. 2.此型號STDF，請協助將Wafer#04 Hold To TPD. | NO | NO | NO | YES | YES | NO
EQPRepairJ22109016 | T6373-15 |  | Other | JB6H1S04 | NC01690G(COG48AA)-A-AT-N4-C | GGG485.1-BC1212 |  | 機台自動貨上貨 無法自動帶入超過32字元 待輪工確認 | 07854 | Other | Other | 2022/08/17 18:16:37 | 2022/08/17 19:00:48 | 0.74 | OK | 13222 | Other--- 1. 產線反應 換批 自動化上貨待入lot 資訊時 因程式超過32字元 無法自動待入 2. 查看前一批 可正常 自動化待入無異常，請產線先換批測試 待觀察 | NO | NO | NO | YES | YES | NO
EQPRepairJ22109026 | T6373-15 |  | Other | JB6H1S04 | NC01690G(COG48AA)-A-AT-N4-C | GGG485.1-BC1212 |  | 機台自動化上貨 無法自動帶入32字元 待工程師確認 | 07854 | Other | Other | 2022/08/17 19:10:17 | 2022/08/17 19:25:37 | 0.26 | OK | 20277 | 輪工處理---- 1.請產線使用Cbsetup上貨 | NO | NO | NO | YES | YES | NO
EQPRepairJ22109034 | T6373-21 | CWC8J0621-01.001 | Prober Down | JB1GKS08 | PA6543B | J6AC8 | 22 | AlarmCode:8630 AlarmMessage:FOCUS ERROR !! (INSTABILITY) | CPAUTO | Prober | focus error | 2022/08/17 19:54:06 | 2022/08/17 20:00:57 | 0.11 | OK | 17091 | 點start 重抓ok | NO | NO | NO | YES | YES | NO
EQPRepairJ22109039 | T6373-99 | CWC7J0L61-01.001 | OverKill | JB1TAS16 | PA5208K-LT | QT7CH | 23 | #23 Already follow the overkilling procedure but still OS | A6896 | Over Kill | O/S誤宰 | 2022/08/17 20:25:24 | 2022/08/17 20:51:46 | 0.44 | OK | 20182 | 1.清針，重對針 重mnt PASS 2.二分法RG-10 PM(45,6) | NO | NO | NO | YES | YES | NO
EQPRepairJ22109042 | T6373-21 | CWC8J0621-01.001 | OverKill | JB1GKS08 | PA6543B | J6AC8 | 22 | Too many o/s | A6898 | Over Kill | O/S誤宰 | 2022/08/17 20:49:43 | 2022/08/17 21:04:15 | 0.24 | OK | 20277 | Remote---- 1.Soaktime Range:X>20，Y:14 機台清針 起測一段PASS穩定，還機生產 | NO | NO | NO | YES | YES | NO
EQPRepairJ22109048 | T6373-38 | CWC8J0621-01.002 | OverKill | JB1GKS09 | PA6543B | J6AC8 | 10 | too many o/s-A6898 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/17 21:17:02 | 2022/08/17 21:26:56 | 0.17 | OK | 15492 | 1.零星 O/S，清、調針 PASS | NO | NO | NO | YES | YES | NO
EQPRepairJ22109054 | T6373-99 | CWC7J0L61-01.001 | OverKill | JB1TAS16 | PA5208K-LT | QT7CH | 23 | Already follow Overkilll handling procedures but still not effective, TOO MANY O/S | A7165 | Over Kill | O/S誤宰 | 2022/08/17 21:39:09 | 2022/08/17 22:41:54 | 1.05 | OK | 20277 | Remote---- 1.接手，重新對針 Range:16 機台清針 PASS | NO | NO | NO | YES | YES | NO
EQPRepairJ22109055 | T6373-152 |  | Setup Fail | JB2P2S06 | R63455-A1S-EZ-CP | NAHM83.00-B1 |  | setup fail due prober load NG-A6936 | CPAUTO | Tester | Tester Down | 2022/08/17 21:40:36 | 2022/08/17 23:08:18 | 1.46 | OK | 20277 | Tester Down---- 1.Setup load PGM Fail，Tester not init 2.改掛Tester Down，Cal ABS | NO | NO | NO | YES | YES | NO
EQPRepairJ22109075 | T6373-152 |  | Tester Down | JB2P2S06 | R63455-A1S-EZ-CP | NAHM83.00-B1 |  | Tester Down---- 1.Setup load PGM Fail，Tester not init 2.改掛Tester Down，Cal ABS | 20277 | Tester | Tester Down | 2022/08/17 23:08:43 | 2022/08/17 23:19:49 | 0.19 | OK | 20277 | 輪工處理---- 1.Cal ABS PASS，Load PGM OK | NO | NO | NO | NO | YES | NO
EQPRepairJ22109088 | T6373-70 | CWC7J0M28-01.004 | OverKill | JB1Q4S17 | PA6542C | J683J | 22 | #20,21 (2h)(corr pass)-03491 | CPAUTO | Over Kill | 2h | 2022/08/17 23:56:59 | 2022/08/18 00:11:28 | 0.24 | OK | 20182 | CORR PASS 續測 | NO | NO | NO | YES | YES | NO
EQPRepairJ22109103 | T6373-38 | CWC8J0615-01.003 | OverKill | JB1GKS09 | PA6543B | J6AC7 | 02 | too many o/s-A6938 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/18 01:01:52 | 2022/08/18 01:12:30 | 0.18 | OK | 17091 | soak time 後機台清針，二分法調針range:-18~-6 起測一段穩定還機 | NO | NO | NO | YES | YES | NO
EQPRepairJ22109117 | T6373-112 |  | Prober Down | JB1GVS25 | NC01682A(COG48AA)-A-AT-N2-C | GGS284.1-B1C221112 |  | AlarmCode:1057 AlarmMessage:MODEL NOT FOUND ( ROUGH ALIGNMENT ) !! | CPAUTO | Prober | Wafer Edge/Mode Not Found | 2022/08/18 01:25:53 | 2022/08/18 01:57:35 | 0.53 | OK | 20182 | 重抓水平ok | NO | NO | NO | YES | YES | NO
EQPRepairJ22109119 | T6373-111 | CWC7J0M28-01.003 | OverKill | JB1Q4S14 | PA6542C | J683J | 01 | TOO MANY O/S-A6294 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/18 01:37:39 | 2022/08/18 01:49:31 | 0.20 | OK | 10110 | Remote---清針調針pass | NO | NO | NO | YES | YES | NO
EQPRepairJ22109122 | T6373-152 |  | Corr Fail | JB2P2S06 | R63455-A1S-EZ-CP | NAHM83.00-B1 |  | Corr Fail | A6936 | Corr | Corr Fail | 2022/08/18 02:04:57 | 2022/08/18 02:08:33 | 0.06 | OK | 20277 | Corr Fail----手動比對PASS | NO | NO | NO | YES | YES | NO
EQPRepairJ22109126 | T6373-112 | CWC8J0189-01.007 | OverKill | JB1GVS25 | NC01682A(COG48AA)-A-AT-N2-C | GGS284.1-B1C2211121 | 05 | #5 Already follow the overkilling procedure but still OS | A6896 | Over Kill | O/S誤宰 | 2022/08/18 02:08:29 | 2022/08/18 03:59:40 | 1.85 | OK | 17091 | 接手處理TESTER螢幕不會亮+OS下卡摩針，查看後方線路螢幕線PIN斷裂，更換備品OK 壞品放置備品間，配件#25 下卡磨針上機固定PIN，清調針、重MNT相同，送修，更換配件#23 PASS，RANGE:-20~-10 PM(-2，3) | NO | NO | NO | YES | YES | NO
EQPRepairJ22109127 | T6373-111 | CWC7J0M28-01.003 | OverKill | JB1Q4S14 | PA6542C | J683J | 01 | OO MANY O/S- PARTICLE-A6294 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/18 02:09:53 | 2022/08/18 05:04:59 | 2.92 | OK | 10110 | 1.配件#14測試易跑open不穩，清針無效 2.更換備卡#22上機固定pin open，重mnt清針無效 3.再次更換備卡#16上機pass | NO | NO | NO | YES | YES | NO
EQPRepairJ22109133 | T6373-147 | CWC8J0217-01 | OverKill | JB1Q4S23 | PA6542C | J68S2 | 16 | Already follow Overkilll handling procedures but still not effective too many fail bin | A7165 | Over Kill | O/S誤宰 | 2022/08/18 02:30:50 | 2022/08/18 02:48:01 | 0.29 | OK | 17091 | remote---機抬清針pass | NO | NO | NO | YES | YES | NO
EQPRepairJ22109138 | T6373-33 | CWC7J0H51-01.001 | OverKill | JB1TAS28 | PA5208K2 | QT7WQ | 06 | too many o/s pm up and down | A6675 | Over Kill | O/S誤宰 | 2022/08/18 02:57:05 | 2022/08/18 03:10:41 | 0.23 | OK | 10897 | SOAKTIME與清針PASS，測試一段穩定(46.6) | NO | NO | NO | YES | YES | NO
EQPRepairJ22109149 | T6373-152 | CWC7J0H67-01 | Prober Down | JB2P2S06 | R63455-A1S-EZ-CP | NAHM72.00-B1C | 03 | AlarmCode:4827 AlarmMessage:ARM @@VACUUM ERROR !! | CPAUTO | Prober | 手臂異常(Arm Error) | 2022/08/18 03:29:43 | 2022/08/18 04:25:04 | 0.92 | OK | 20277 | 輪工處理---- 1.Prober Reset OK，Range:8 PASS，#03 Load again | NO | NO | NO | YES | YES | NO
EQPRepairJ22109181 | T6373-111 | CWC7J0M28-01.003 | OverKill | JB1Q4S16 | PA6542C | J683J | 01 | too many os-A6512 | CPAUTO | Prober | 結霜 | 2022/08/18 05:47:45 | 2022/08/18 05:50:20 | 0.04 | OK | 10110 | Prober Down--- 1.換卡後測試一段跑open 2.確認有些許結霜現象 3.DP:-62 升溫100度烘烤 4.先改掛Prober Down | NO | NO | NO | YES | YES | NO
EQPRepairJ22109183 | T6373-111 |  | Prober Down | JB1Q4S16 | PA6542C | J683J |  | Prober Down--- 1.換卡後測試一段跑open 2.確認有些許結霜現象 3.DP:-62 升溫100度烘烤 4.先改掛Prober Down | 10110 | Prober | Other | 2022/08/18 05:50:59 | 2022/08/18 07:00:15 | 1.15 | OK | 10110 | Prober Down--- 1.換卡後測試一段跑open 2.確認有些許結霜現象 3.DP:-62 升溫100度烘烤20min 4.重新降溫後setup pass range:8um | NO | NO | NO | YES | YES | NO
EQPRepairJ22109187 | T6373-47 | CWC6J0M86-01 | Corr Fail | JB1HKS07 | NA01687F(COG48BA)-A-AT-C | GGR359.1-B | 25 | ok-10851 | CPAUTO | Corr | Corr Fail | 2022/08/18 06:08:13 | 2022/08/18 06:11:42 | 0.06 | OK | 20277 | Corr Fail----系統比對異常，手動比對 Corr Pass | NO | NO | NO | YES | YES | NO
EQPRepairJ22109188 | T6373-152 | CWC7J0H67-01 | Prober Down | JB2P2S06 | R63455-A1S-EZ-CP | NAHM72.00-B1C | 05 | AlarmCode:4827 AlarmMessage:ARM @@VACUUM ERROR !! | CPAUTO | Prober | 手臂異常(Arm Error) | 2022/08/18 06:08:25 | 2022/08/18 07:54:56 | 1.78 | OK | 19771 | 接手處理 #2(56.8) | NO | NO | NO | YES | YES | NO
EQPRepairJ22109189 | T6373-112 | CWC8J0189-01.007 | OverKill | JB1GVS23 | NC01682A(COG48AA)-A-AT-N2-C | GGS284.1-B1C2211121 | 10 | #10 OS/ PM NG XY : 169.4 OD : 11 PM SHIFT U/D PM | A6896 | Over Kill | O/S誤宰 | 2022/08/18 06:09:10 | 2022/08/18 06:34:48 | 0.43 | OK | 20277 | 輪工處理---- 1.重新對針 Range:16 PASS | NO | NO | NO | YES | YES | NO
EQPRepairJ22109191 | T6373-86 | CWC7J0C21-01.001 | OverKill | JB1Q4S15 | PA6542C | J68S1 | 24 | too many o/s | A6675 | Over Kill | O/S誤宰 | 2022/08/18 06:18:30 | 2022/08/18 06:47:40 | 0.49 | OK | 10110 | Remote---重新對針後pass | NO | NO | NO | YES | YES | NO
EQPRepairJ22109193 | T6373-39 | CWC8J0087-01 | Tester Down | JB1UJS11 | PA5129LS_12U80H-LT | P078R.1 | 08 | AlarmCode:669 AlarmMessage:GP-IB COMMUNICATION TIMEOUT ERROR !! | CPAUTO | Tester | 離線(OCT/GPIB error) | 2022/08/18 06:21:48 | 2022/08/18 06:30:52 | 0.15 | OK | 17091 | 崇廉ok | NO | NO | NO | YES | YES | NO
EQPRepairJ22109197 | T6373-151 | CWC8J0038-01 | OverKill | JB2P2S08 | R63455-A1S-EZ-CP | NAPY65.00-B1C | 15 | bin23-03491 | CPAUTO | Over Kill | Function誤宰 | 2022/08/18 06:38:45 | 2022/08/18 06:44:55 | 0.10 | OK | 20277 | 輪工處理---- 1.Bin23 先前已驗證過換卡無效，換機回收也無法回收，前一片並無此問題，請產線Run Corr驗證，續測觀察 | NO | NO | NO | YES | YES | NO
EQPRepairJ22109207 | T6373-47 | CWC8J0291-01.003 | Prober Down | JB1HKS07 | NA01687F(COG48BA)-A-AT-C | GGU014.1-BC12111 | 08 | AlarmCode:1057 AlarmMessage:MODEL NOT FOUND ( ROUGH ALIGNMENT ) !! | CPAUTO | Prober | Wafer Edge/Mode Not Found | 2022/08/18 07:30:49 | 2022/08/18 07:40:19 | 0.16 | OK | 17680 | t44產品，先行切回idle | NO | NO | NO | YES | YES | NO
EQPRepairJ22109217 | T6373-111 | CWC7J0M28-01.003 | OverKill | JB1Q4S16 | PA6542C | J683J | 03 | TOO MANY O/S-20022 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/18 08:09:02 | 2022/08/18 08:40:48 | 0.53 | OK | 17680 | remote----換片o/s，調整針位position(0,4),mark(39,0)，OD→5um，起測穩定 | NO | NO | NO | YES | YES | NO
EQPRepairJ22109229 | T6373-129 |  | Prober Down | JB1QNS06 | NA01687G(COG48BA)-A-AT-X1-C | GGR354.1-B1 |  | MACHINE ALARM VACUUM 1 ERROR,setup pssed (103.4) dichotomy (3.4) range:x(-23.29) y(12.-4) no corr wafer-A6049 | CPAUTO | Prober | Loader Down | 2022/08/18 09:02:46 | 2022/08/18 09:16:25 | 0.23 | OK | 17680 | 開測蓋確認wafer位置，prober shutdown重開ok,升溫烘烤，轉掛prober down | NO | NO | NO | YES | YES | NO
EQPRepairJ22109231 | T6373-129 |  | Prober Down | JB1QNS06 | NA01687G(COG48BA)-A-AT-X1-C | GGR354.1-B1 |  | 開測蓋確認wafer位置，prober shutdown重開ok,升溫烘烤，轉掛prober down | 17680 | Prober | Loader Down | 2022/08/18 09:17:00 | 2022/08/18 10:45:20 | 1.47 | OK | 17680 | remote----resetup,position(0,0),mark(37,0) | NO | NO | NO | YES | YES | NO
EQPRepairJ22109242 | T6373-152 | CWC7J0H67-01 | Other | JB2P2S06 | R63455-A1S-EZ-CP | NAHM72.00-B1C | 06 | PLEASE CHECK MACHINE (THREE PIN DELAYING EVERY WAFER 1HR WAITING)-A6986 | CPAUTO | Other | Other | 2022/08/18 10:29:04 | 2022/08/18 12:30:29 | 2.02 | OK | 08774 | 測試時間異常 1.確認prober 參數 2.tester reboot 3.cal online init ，pass 還機產品借機 | NO | NO | NO | YES | YES | NO
EQPRepairJ22109255 | T6373-13 | CWC8J0558-01.004 | Corr Fail | JB1GVS19 | NC01682B(COG48AA)-A-AT-N-C | GGR939.1-B | 07 | setup run corr-05671 | CPAUTO | Corr | 產品問題 | 2022/08/18 11:39:00 | 2022/08/18 11:47:31 | 0.14 | OK | 11385 | Remote 處理----手動比對PASS | NO | NO | NO | YES | YES | NO
EQPRepairJ22109268 | T6373-111 | CWC7J0M28-01.003 | OverKill | JB1Q4S16 | PA6542C | J683J | 04 | O/S-17436 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/18 13:17:53 | 2022/08/18 16:41:05 | 3.39 | OK | 17680 | 1.replace #22，固定pin open 2.replace #14，o/s不穩，wafer易結霜升溫烘烤 3.轉掛prober down | NO | NO | NO | YES | YES | NO
EQPRepairJ22109270 | T6373-39 | CWC8J0087-01 | Tester Down | JB1UJS11 | PA5129LS_12U80H-LT | P078R.1 | 10 | AlarmCode:669 AlarmMessage:GP-IB COMMUNICATION TIMEOUT ERROR !! | CPAUTO | Tester | 離線(OCT/GPIB error) | 2022/08/18 13:22:10 | 2022/08/18 13:30:29 | 0.14 | OK | 19771 | 重連OK | NO | NO | NO | YES | YES | NO
EQPRepairJ22109276 | T6373-143 |  | Auto move in Issue | VB405S09 | PA5208K-LT | QT7CH |  | AE Testing...-13749 | CPAUTO | Other | Other | 2022/08/18 13:39:06 | 2022/08/18 13:46:56 | 0.13 | OK | 13749 | Check Machine EAP-Mail ，AE Check ok IDLE | NO | NO | NO | YES | YES | NO
EQPRepairJ22109283 | T6373-135 | CWC8J0851-01 | OverKill | TB1YLS38 | PA6485E(CP1-CP3) | P0CSF.1 | 07 | o/s-13081 | CPAUTO | Over Kill | O/D不足 | 2022/08/18 14:15:12 | 2022/08/18 14:34:00 | 0.31 | OK | 19771 | 重新對針 #7(180.6) | NO | NO | NO | NO | YES | NO
EQPRepairJ22109289 | T6373-121 | CMC8J0023-01 | Corr Fail | TB1IMS18 | NA01404E-A-AT-X2-C | GFWM86.1-1BC221212C | 12 | t44corr compare fail-17436 | CPAUTO | Corr | 產品問題 | 2022/08/18 14:53:00 | 2022/08/18 15:00:27 | 0.12 | OK | 11385 | Corr Fail---- T44 產品無 Corr Wafer Run Corr 比對 Fail | NO | NO | NO | YES | YES | NO
EQPRepairJ22109290 | T6373-128 | CWC7J0M33-01 | OverKill | VB1BVS02 | PA5270D-LT | U38F8.00 | 08 | cont. 2H - NO CORR-a6962 | CPAUTO | Over Kill | 2H | 2022/08/18 15:04:53 | 2022/08/18 15:07:42 | 0.05 | OK | 08774 | 2H 1.[CUSTOMER_HOLD_CRITERIA] [HW_BIN0<100%](Bin 1 < 798EA) 703(Hold Wafer) ----- 2.續測 | NO | NO | NO | YES | YES | NO
EQPRepairJ22109294 | T6373-135 | CWC8J0851-01 | OverKill | TB1YLS38 | PA6485E(CP1-CP3) | P0CSF.1 | 08 | too many o/s-A6671 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/18 15:21:52 | 2022/08/18 16:07:13 | 0.76 | OK | 08774 | Remote 處理----重新對針 機台清針7.4 | NO | NO | NO | YES | YES | NO
EQPRepairJ22109297 | T6373-47 | CWC8J0391-01.004 | OverKill | JB1HKS07 | NA01687G(COG48BA)-A-AT-N2-C | GGS419.1-B11C1112 | 25 | too many o/s, retest failo/s..pm check up/down (146,4)-A6986 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/18 15:33:25 | 2022/08/18 16:08:18 | 0.58 | OK | 16631 | 1.針位貼上，調整針位/CLEAN WAFER清針，Online init | NO | NO | NO | YES | YES | NO
EQPRepairJ22109299 | T6373-147 | CWC8J0217-01 | Other | JB1Q4S23 | PA6542C | J68S2 | 21 | change to MBD by leader-A6671 | CPAUTO | Probe Card | Function Fail更換P/C | 2022/08/18 15:43:22 | 2022/08/18 17:20:57 | 1.63 | OK | 08774 | 1.bin6過多RD指示換卡,配件送修 2.#21 (128.5) Y=17 | NO | NO | NO | YES | YES | NO
EQPRepairJ22109309 | T6373-111 |  | Prober Down | JB1Q4S16 | PA6542C | J683J |  | 1.replace #22，固定pin open 2.replace #14，o/s不穩，wafer易結霜升溫烘烤 3.轉掛prober down | 17680 | Over Kill | O/S | 2022/08/18 16:41:35 | 2022/08/18 17:46:11 | 1.08 | OK | 11385 | Remote 處理----1.replace #22，固定pin open 2.replace #14，o/s不穩，wafer易結霜升溫烘烤，重新降溫s/u pass穩定續測(131.0) | NO | NO | NO | YES | YES | NO
EQPRepairJ22109310 | T6373-47 |  | Prober Down | JB1HKS07 | NA01687G(COG48BA)-A-AT-N2-C | GGS419.1-B11C1112 |  | LOADPORT ERROR ! | A6916 | Prober | loadport error | 2022/08/18 16:43:22 | 2022/08/18 17:39:05 | 0.93 | OK | 19771 | loader init ok | NO | NO | NO | YES | YES | NO
EQPRepairJ22109326 | T6373-129 | CWC8J0391-01 | OverKill | JB1QNS06 | NA01687G(COG48BA)-A-AT-N2-C | GGS419.1-B11C1111 | 02 | TOO MANY O/S | A6897 | Over Kill | O/S | 2022/08/18 18:01:11 | 2022/08/18 18:20:37 | 0.32 | OK | 11385 | Remote 處理----重新對針調整針位 (146.0) | NO | NO | NO | YES | YES | NO
EQPRepairJ22109328 | T6373-147 | CWC8J0217-01 | Other | JB1Q4S16 | PA6542C | J68S2 | 19 | AE CHANGE PC | A6916 | Over Kill | O/S | 2022/08/18 18:02:59 | 2022/08/18 19:09:06 | 1.10 | OK | 19771 | 1.易跳O/S,清調針無效 2.配件磨針無效 3.查看log 易跳固定pin,配件送修,無備卡轉掛配件當 | NO | NO | NO | YES | YES | NO
EQPRepairJ22109342 | T6373-16 |  | Prober Down | VB405S09 | PA5208KF-LT | J6CCA |  | AlarmCode:1995 AlarmMessage:Faied to upload log file | CPAUTO | Other | Other | 2022/08/18 18:53:16 | 2022/08/18 18:56:28 | 0.05 | OK | 08774 | Remote 處理----確認後 續測 | NO | NO | NO | YES | YES | NO
EQPRepairJ22109344 | T6373-147 |  | Probe Card Down | JB1Q4S16 | PA6542C | J68S2 |  | 易跳固定pin ,配件送修,無備卡轉掛配件當 | 19771 | Probe Card | Other | 2022/08/18 19:09:58 | 2022/08/18 23:56:26 | 4.77 | OK | 17674 | p/c down 1. 配件#16修復上機Pass， 點測穩定Pass起測易OS不穩清/調針無效，故將配件再次送修無備卡 2. 更換備卡JB1Q4S22 上機Bin2-open點測亂跳pin，清調針無效，故將配件再次送修無備卡 | NO | NO | NO | YES | YES | NO
EQPRepairJ22109359 | T6373-43 | CWC7J0312-01 | Tester Down | JB1UJS03 | PA5129LS_12U80H-LT | P060N.1 | 13 | AlarmCode:669 AlarmMessage:GP-IB COMMUNICATION TIMEOUT ERROR !! | CPAUTO | Prober | 網路異常(Spooling Error) | 2022/08/18 20:10:56 | 2022/08/18 20:20:28 | 0.16 | OK | 17674 | 1. GP-IB 測試至 (45,4) Tester 正常，重新起測 | NO | NO | NO | YES | YES | NO
EQPRepairJ22109372 | T6373-39 | CWC8J0087-01 | OverKill | JB1UJS11 | PA5129LS_12U80H-LT | P078R.1 | 13 | too many o/s-A7033 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/18 21:49:04 | 2022/08/18 23:48:30 | 1.99 | OK | 17674 | 1. 配件#11 殘金 下卡磨針配件室告知針位不佳，更換備卡#13 PASSS， 2. 溫針 RANGE Y=10 (99,4)，休息 | NO | NO | NO | YES | YES | NO
EQPRepairJ22109377 | T6373-43 | CWC7J0312-01 | Tester Down | JB1UJS03 | PA5129LS_12U80H-LT | P060N.1 | 16 | AlarmCode:669 AlarmMessage:GP-IB COMMUNICATION TIMEOUT ERROR !! | CPAUTO | Tester | 離線(OCT/GPIB error) | 2022/08/18 22:11:13 | 2022/08/18 22:21:43 | 0.18 | OK | 13597 | 機台無訊息，無離線，測試正常，還機生產 | NO | NO | NO | YES | YES | NO
EQPRepairJ22109396 | T6373-147 |  | Probe Card Down | JB1Q4S16 | PA6542C | J68S2 |  | Probe Card Down----- 1. 配件#16修復上機Pass， 點測穩定Pass起測易OS不穩清/調針無效，故將配件再次送修無備卡 2. 更換備卡JB1Q4S22 上機Bin2-open點測亂跳pin，清調針無效，故將配件再次送修無備卡 | 17674 | Probe Card | Other | 2022/08/18 23:56:55 | 2022/08/19 07:01:19 | 7.07 | OK | 17674 | p/c down ----- 1. TB1R8S85 改測，c-lot s/u pass， range y=8 x=20 (16,5)#11，掛s/u 04 run corr | NO | NO | NO | YES | YES | NO
EQPRepairJ22109406 | T6373-99 | CWC7J0L19-01 | OverKill | JB1TAS16 | PA5208K-LT | QT7WS | 06 | MBD-A7164 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/19 01:20:25 | 2022/08/19 03:12:43 | 1.87 | OK | 17674 | 1. 測試中OS配件#16殘金，更換備卡#24 上機OPEN 固定Pin不過，清調針無效 2. 再更換備卡#27 Pass，故將配件#24送修.，RANGE Y=10 (34,m4) | NO | NO | NO | YES | YES | NO
EQPRepairJ22109409 | T6373-38 | CWC8J0615-01.003 | OverKill | JB1GKS09 | PA6543B | J6AC7 | 23 | Too many o/s-A6053 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/19 01:34:26 | 2022/08/19 01:47:02 | 0.21 | OK | 03426 | 機台清針PASS | NO | NO | NO | YES | YES | NO
EQPRepairJ22109413 | T6373-111 | CWC7J0M28-01.003 | OverKill | JB1Q4S14 | PA6542C | J683J | 09 | Already follow Overkilll handling procedures but still not effective too many o/s | A5548 | Over Kill | O/S誤宰 | 2022/08/19 01:41:41 | 2022/08/19 03:47:42 | 2.10 | OK | 11348 | <低溫下卡>O/S 下卡磨針上機等待20MIN降溫SETUP PASS | NO | NO | NO | YES | YES | NO
EQPRepairJ22109422 | T6373-37 | CWC7J0M17-01 | OverKill | TB2U7S09 | R69005A0FNCPJDQV-P | AYP072000-B2C | 13 | TOO MANY O/S PM CHECK 175.8 OD15-A7024 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/19 02:10:45 | 2022/08/19 02:25:12 | 0.24 | OK | 03426 | 針位調整 機台清針PASS | NO | NO | NO | YES | YES | NO
EQPRepairJ22109435 | T6373-47 | CWC5J0T09-01.007 | OverKill | JB1HKS07 | NA01687G(COG48BA)-A-AT-N2-C | GGM287.1-BC121 | 11 | too many fail bin 50-A6148 | CPAUTO | Over Kill | Function誤宰 | 2022/08/19 03:22:11 | 2022/08/19 03:59:35 | 0.62 | OK | 03426 | 針位調整PASS | NO | NO | NO | YES | YES | NO
EQPRepairJ22109450 | T6373-128 | CWC7J0M33-01.001 | OverKill | VB1BVS02 | PA5270D-LT | U38F8.00 | 17 | 2 continues eda show H no corr | A6926 | Over Kill | 2H | 2022/08/19 04:53:29 | 2022/08/19 05:45:09 | 0.86 | OK | 17674 | 1. 2H (Bin 1 < 798EA) 705(Hold Wafer)，不同FAIL BIN 過多，同批不同機相同問題 | NO | NO | NO | YES | YES | NO
EQPRepairJ22109456 | T6373-29 | CWC8J0291-01.009 | Setup Fail | JB1QNS01 | NA01687F(COG48BA)-A-AT-C | GGU014.1-B | 06 | 05:30 ST FAIL TOO LARGE NO AE CAN repair | A5929 | Setup Fail | Prober異常 prober abnormal | 2022/08/19 05:43:20 | 2022/08/19 06:01:09 | 0.30 | OK | 03426 | 重新對針ok 請產線SET UP | NO | NO | NO | YES | YES | NO
EQPRepairJ22109470 | T6373-90 | CWC8J0088-01.001 | Setup Fail | JB1BYS11 | PA6587B_15U90H | J696C | 03 | 06:13 #03 id ok, setup fail due to fail bin 12, no other available p/c polish at 6th floor | A6053 | Setup Fail | Function Fail | 2022/08/19 07:12:31 | 2022/08/19 07:59:19 | 0.78 | OK | 16631 | 1.磨針上機pass 133.7，setup 04 | NO | NO | NO | YES | YES | NO
EQPRepairJ22109504 | T6373-128 | CWC7J0M33-01.001 | OverKill | VB1BVS02 | PA5270D-LT | U38F8.00 | 19 | too many failbin 34-A6986 | CPAUTO | Over Kill | Function誤宰 | 2022/08/19 10:26:07 | 2022/08/19 10:38:53 | 0.21 | OK | 08774 | BIN34 處理後PASS，反應TPD確認 | NO | NO | NO | YES | YES | NO
EQPRepairJ22109507 | T6373-32 | CWC8J0508-01.002 | OverKill | KB1PNS01 | NA01761D(COF94AA)-A-Y-C | GGS421.1-BC2211 | 17 | o/s-20022 | CPAUTO | Over Kill | O/D不足 | 2022/08/19 10:37:55 | 2022/08/19 10:54:03 | 0.27 | OK | 16631 | 1.wafer邊緣o/s，調整針壓pass | NO | NO | NO | YES | YES | NO
EQPRepairJ22109517 | T6373-90 | CWC7J0M27-01 | OverKill | JB1BYS11 | PA6587B_15U90H | J696G | 02 | FAIL BIN 12-17436 | CPAUTO | Probe Card | Function Fail更換P/C | 2022/08/19 11:30:35 | 2022/08/19 12:04:01 | 0.56 | OK | 13749 | P/C Down--- 1.配件JB1BYS11 測試中跑大量 BIN12 Function 過多不穩. 2.確認此配件先前皆有多次相同送修紀錄. 3.故將配件#S11 先行送修，無備卡. | NO | NO | NO | YES | YES | NO
EQPRepairJ22109518 | T6373-38 | CWC8J0615-01.008 | Corr Fail | JB1GKS12 | PA6543B | J6AC7 | 05 | Corr Fail-A6916 | CPAUTO | Corr | Corr Fail | 2022/08/19 11:31:46 | 2022/08/19 11:35:38 | 0.06 | OK | 13749 | Corr Fail---T32 無Run corr 首片比對Corr Fail | NO | NO | NO | YES | YES | NO
EQPRepairJ22109522 | T6373-90 |  | Probe Card Down | JB1BYS11 | PA6587B_15U90H | J696G |  | P/C Down--- 1.配件JB1BYS11 測試中跑大量 BIN12 Function 過多不穩. 2.確認此配件先前皆有多次相同送修紀錄. 3.故將配件#S11 先行送修，無備卡. | 13749 | Probe Card | Function Fail | 2022/08/19 12:04:25 | 2022/08/19 15:51:12 | 3.78 | OK | 16631 | 1.配件當改測，setup pass 160.4，setup 04 | NO | NO | NO | YES | YES | NO
EQPRepairJ22109523 | T6373-16 | CWC8J0908-01 | OverKill | VB405S09 | PA5208KF-LT | J6G42 | 13 | TOO MANY O/S-A6526 | CPAUTO | Over Kill | O/D不足 | 2022/08/19 12:09:41 | 2022/08/19 12:29:17 | 0.33 | OK | 19771 | 重新對針ok | NO | NO | NO | YES | YES | NO
EQPRepairJ22109527 | T6373-128 | CWC7J0M33-01.001 | OverKill | VB1BVS02 | PA5270D-LT | U38F8.00 | 21 | too many fail bin34-A6526 | CPAUTO | Over Kill | Function誤宰 | 2022/08/19 12:15:27 | 2022/08/19 12:18:45 | 0.06 | OK | 08774 | 機台要pm idle | NO | NO | NO | YES | YES | NO
EQPRepairJ22109539 | T6373-38 | CWC8J0615-01.008 | OverKill | JB1GKS12 | PA6543B | J6AC7 | 07 | CONTINUOUS EDA SHOW 2H, NO CORR WAFER-A6671 | CPAUTO | Over Kill | 2H | 2022/08/19 13:25:44 | 2022/08/19 13:36:44 | 0.18 | OK | 16631 | 1.2H，BIN43，請產線更換貨批驗證 | NO | NO | NO | YES | YES | NO
EQPRepairJ22109547 | T6373-136 | CWC8J0969-01 | Setup Fail | JB8WXS08 | ND01733D(COF94AA)-A-C | H2PU18.3-22B | 21 | setup o/s 調針清針重mount無效 下卡磨針無備卡 | 05671 | Setup Fail | O/S Fail | 2022/08/19 13:53:06 | 2022/08/19 15:25:47 | 1.54 | OK | 19771 | 配件磨針上機pass #21(193.8) | NO | NO | NO | YES | YES | NO
EQPRepairJ22109587 | T6373-38 |  | Corr Fail | JB1GKS12 | PA6543B | J6AC7 |  | Corr Fail | A6916 | Corr | Corr Fail | 2022/08/19 16:57:49 | 2022/08/19 17:15:44 | 0.30 | OK | 13749 | Corr Fail--- 1.手動比對EDA 扣除BIN127 系統可忽略Fail Bin後比對PASS | NO | NO | NO | YES | YES | NO
EQPRepairJ22109602 | T6373-136 |  | Prober Down | JB8WXS08 | ND01733D(COF94AA)-A-C | H2PU18.3-22BC |  | AlarmCode:1057 AlarmMessage:MODEL NOT FOUND ( ROUGH ALIGNMENT ) !! | CPAUTO | Prober | Wafer Edge/Mode Not Found | 2022/08/19 18:03:46 | 2022/08/19 18:26:44 | 0.38 | OK | 08774 | Remote 處理---重抓 | NO | NO | NO | YES | YES | NO
EQPRepairJ22109612 | T6373-129 | CEC8J0033-01.001 | OverKill | JB1EBS10 | NA01644G(COG88AA)-A-AT-V-C | GGW352.1-2B1C2 | 04 | BIN11 | A6962 | Over Kill | Function誤宰 | 2022/08/19 18:53:21 | 2022/08/19 20:57:31 | 2.07 | OK | 11348 | <低溫換卡>BIN11 LEAK 更換配件#18 等待20MIN 降溫SETUP PASS | NO | NO | NO | YES | YES | NO
EQPRepairJ22109615 | T6373-151 | CWC8J0116-01 | Other |  |  |  | 19 | ae check-11933 | CPAUTO | Tester | Open / Short / Function Fail | 2022/08/19 18:58:31 | 2022/08/19 19:09:10 | 0.18 | OK | 08774 | Tester Down----R63455 CP1 BIN23 1.TPD要求SWAP PE CARD#1.3，2.4 SWAP CAL機中 | NO | NO | NO | YES | YES | NO
EQPRepairJ22109617 | T6373-151 |  | Tester Down |  |  |  |  | Tester Down----R63455 CP1 BIN23 1.TPD要求SWAP PE CARD#1.3，2.4 SWAP CAL機中 | 08774 | Tester | Other | 2022/08/19 19:09:36 | 2022/08/19 21:33:37 | 2.40 | OK | 17674 | Tester down ------ 1. R63455/CP1 Bin23 ，swap pe #01<>03 #02<>4，cal pass，c-lot s/u 上機pass ，#19(105,5) range y=10 (191,6)、(192,6) | NO | NO | NO | YES | YES | NO
EQPRepairJ22109623 | T6373-111 | CWC7J0M28-01.004 | OverKill | JB1Q4S14 | PA6542C | J683J | 15 | too many o/s-A6148 | CPAUTO | Prober | WAFER結霜 | 2022/08/19 19:58:10 | 2022/08/19 21:26:03 | 1.46 | OK | 11348 | <低溫下卡>O/S內層針黑點下卡磨針上機等待20MIN降溫SETUP依然有結霜問題升溫烘烤 | NO | NO | NO | YES | YES | NO
EQPRepairJ22109628 | T6373-135 | CWC8J0957-01 | OverKill | TB1YLS38 | PA6485E(CP1-CP3) | P0CSG.1 | 04 | too many o/s-A7033 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/19 20:06:09 | 2022/08/19 20:40:32 | 0.57 | OK | 17674 | 1. 下卡磨針pass，range y=10 (44,5)#04 | NO | NO | NO | YES | YES | NO
EQPRepairJ22109629 | T6373-136 | CWC8J0969-01 | OverKill | JB8WXS08 | ND01733D(COF94AA)-A-C | H2PU18.3-22BC | 23 | eda 2 I no corr wafer-A7033 | CPAUTO | Over Kill | 2I | 2022/08/19 20:13:21 | 2022/08/19 20:26:19 | 0.22 | OK | 17674 | 1. 2I Wafer 外圍有區域性fail bin ，同批不同機卡相同問題 | NO | NO | NO | YES | YES | NO
EQPRepairJ22109637 | T6373-34 | CWC5J0C35-01 | OverKill | JB1BYS10 | PA6587B_15U90H | J62W7 | 08 | TOO MANY O/S PM CHECK 18.1 OD8-A7024 | CPAUTO | Over Kill | O/D不足 | 2022/08/19 20:53:41 | 2022/08/19 21:14:01 | 0.34 | OK | 03426 | 重新對針PASS 43.0 43.7 | NO | NO | NO | YES | YES | NO
EQPRepairJ22109639 | T6373-37 |  | Other | TB2U7S09 | R69005A0FNCPJDQV-P | AYP685000-B1C |  | BROCODE槍 刷不出來 | 19697 | Other | Other | 2022/08/19 20:58:48 | 2022/08/19 21:25:19 | 0.44 | OK | 03426 | 只有晶舟條碼刷不出來 請產線用手動 | NO | NO | NO | YES | YES | NO
EQPRepairJ22109640 | T6373-08 | CMC1J0047-01 | Setup Fail | VB108S02 | PA5252N | AXM342000 | 02 | O/S FAIL，多根PIN不過，重MOUNT會跳PIN，調針、清針無效，特需單，無備卡，下卡磨針待輪工上機確認。 | 08438 | Prober | Other | 2022/08/19 20:59:00 | 2022/08/20 01:16:29 | 4.29 | OK | 13597 | Prober Down---- 1.改機一串OPEN不過，下卡磨針，換卡驗證一樣同區域OPEN不過 2.擦雲母片、多墊一片雲母片、轉配件角度無效 3.MNT頭時發現後方定位PIN偏移 4.量測DOCKING，上方區域有偏移現象 5.改掛Prober Down，待後續處理 | NO | NO | NO | YES | YES | NO
EQPRepairJ22109642 | T6373-38 | CWC8J0615-01.008 | OverKill | JB1GKS12 | PA6543B | J6AC7 | 12 | #9,11 CONTINUES EDA SHOW 2 H CORROLATION PASSED PM CHECK 157.6 OD5-A7024 | CPAUTO | Over Kill | Show H | 2022/08/19 21:01:11 | 2022/08/19 21:16:39 | 0.26 | OK | 03426 | Show H 請產線續測 | NO | NO | NO | YES | YES | NO
EQPRepairJ22109648 | T6373-111 | CWC7J0M28-01.004 | Prober Down | JB1Q4S14 | PA6542C | J683J | 13 | <低溫下卡>O/S內層針黑點下卡磨針上機等待20MIN降溫SETUP依然有結霜問題升溫烘烤 | 11348 | Prober | WAFER結霜 | 2022/08/19 21:27:08 | 2022/08/19 22:11:52 | 0.75 | OK | 11348 | WAFER結霜升溫烘烤降溫SETUP PASS | NO | NO | NO | YES | YES | NO
EQPRepairJ22109652 | T6373-135 | CWC8J0957-01 | OverKill | TB1YLS38 | PA6485E(CP1-CP3) | P0CSG.1 | 05 | too many o/s-A7033 | CPAUTO | Over Kill | 機台清針後Pass | 2022/08/19 21:41:42 | 2022/08/19 21:47:27 | 0.10 | OK | 17674 | 1. 機台清針後Pass (203,3)#05 | NO | NO | NO | YES | YES | NO
EQPRepairJ22109662 | T6373-33 | CWC7J0K83-01.001 | OverKill | JB1TAS28 | PA5208KF-LT | J67HF | 03 | too many o/s-A6148 | CPAUTO | Over Kill | O/S | 2022/08/19 22:23:46 | 2022/08/19 22:58:12 | 0.57 | OK | 15096 | O/S---調整針位PASS. | NO | NO | NO | YES | YES | NO
EQPRepairJ22109666 | T6373-135 | CWC8J0957-01 | OverKill | TB1YLS38 | PA6485E(CP1-CP3) | P0CSG.1 | 05 | always open short-A6786 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/19 22:28:55 | 2022/08/19 23:58:15 | 1.49 | OK | 17674 | 1. 下卡磨針PASS， RANGE Y=15 (55,5)#05 | NO | NO | NO | YES | YES | NO
EQPRepairJ22109682 | T6373-90 | CWC7J0G36-01.002 | OverKill | JB1BYS12 | PA6587B_15U90H | J696A | 22 | O/S TOO MANY-07658 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/19 23:48:17 | 2022/08/20 02:23:05 | 2.58 | OK | 03426 | 下卡磨針 43.0 45.7 | NO | NO | NO | YES | YES | NO
EQPRepairJ22109692 | T6373-135 | CWC8J0957-01 | OverKill | TB1YLS38 | PA6485E(CP1-CP3) | P0CSG.1 | 06 | O/S TOO MANY(191.3)CHECKED PM.OD15 NO PARTICLES-A6788 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/20 00:43:40 | 2022/08/20 00:57:18 | 0.23 | OK | 15171 | Remote處理 --- 清針 | NO | NO | NO | NO | YES | NO
EQPRepairJ22109695 | T6373-111 |  | Prober Down | JB1Q4S14 | PA6542C | J683J |  | AlarmCode:1060 AlarmMessage:CANNOT MEASURE INDEX SIZE !! | CPAUTO | Prober | Wafer Edge/Mode Not Found | 2022/08/20 00:48:33 | 2022/08/20 00:55:12 | 0.11 | OK | 15096 | Remote處理-----重抓model ok，check die ok. | NO | NO | NO | YES | YES | NO
EQPRepairJ22109700 | T6373-08 |  | Prober Down | TB1R8S79 | PA5223D1_12U60H-LT | PX3W9.1 |  | Prober Down---- 1.改機一串OPEN不過，下卡磨針，換卡驗證一樣同區域OPEN不過 2.擦雲母片、多墊一片雲母片、轉配件角度無效 3.MNT頭時發現後方定位PIN偏移 4.量測DOCKING，上方區域有偏移現象 5.改掛Prober Down，待後續處理 | 13597 | Tester | O/S | 2022/08/20 01:17:05 | 2022/08/20 08:15:09 | 6.97 | OK | 11385 | Tester down--- 1.改機一串OPEN不過，下卡磨針，換卡驗證一樣同區域OPEN不過 2.擦雲母片、多墊一片雲母片、轉配件角度無效 3.MNT頭時發現後方定位PIN偏移 4.量測DOCKING，上方區域有偏移現象 5.拉ND3-23 IDLE機台Hifix上機驗證PASS，確認Docking ok 6.因客戶包機故先行續測 | NO | NO | NO | YES | YES | NO
EQPRepairJ22109704 | T6373-128 | CWC7J0M33-01.001 | OverKill | VB1BVS02 | PA5270D-LT | U38F8.00 | 20 | too many fail bin34-A6786 | CPAUTO | Over Kill | Function誤宰 | 2022/08/20 01:29:29 | 2022/08/20 01:59:32 | 0.50 | OK | 17674 | 1. Function誤宰 34 重load pgm pass與先前T6373-43 相同問題 | NO | NO | NO | YES | YES | NO
EQPRepairJ22109711 | T6373-135 | CWC8J0957-01 | OverKill | TB1YLS38 | PA6485E(CP1-CP3) | P0CSG.1 | 06 | too many o/s-A7033 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/20 01:42:00 | 2022/08/20 02:01:31 | 0.33 | OK | 17674 | 1. 已下卡磨針兩次皆無法久測，查看前一批正常請產線測完#06 換批驗證 | NO | NO | NO | YES | YES | NO
EQPRepairJ22109717 | T6373-129 | CEC8J0033-01 | OverKill | JB1EBS18 | NA01644G(COG88AA)-A-AT-V-C | GGW352.1-2B1C1 | 03 | too many fail bin..-A7025 | CPAUTO | Over Kill | 針位調整 | 2022/08/20 01:56:38 | 2022/08/20 02:10:48 | 0.24 | OK | 20276 | 調整針位 Y軸 6um>8um P/M(156.0) | NO | NO | NO | YES | YES | NO
EQPRepairJ22109723 | T6373-34 | CWC7J0M27-01 | OverKill | JB1BYS10 | PA6587B_15U90H | J696G | 02 | O/S-19697 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/20 02:10:25 | 2022/08/20 03:17:53 | 1.12 | OK | 03426 | 下卡磨針 42.0 45.7 | NO | YES | YES | NO | YES | NO
EQPRepairJ22109726 | T6373-43 | CWC7J0L69-01.003 | Tester Down | JB1UJS03 | PA5129LS_12U80H-LT | P078M.1 | 15 | AlarmCode:669 AlarmMessage:GP-IB COMMUNICATION TIMEOUT ERROR !! | CPAUTO | Tester | 離線(OCT/GPIB error) | 2022/08/20 02:20:51 | 2022/08/20 02:32:35 | 0.20 | OK | 11348 | Remote---離線(OCT/GPIB error)重連OK | NO | NO | NO | YES | YES | NO
EQPRepairJ22109728 | T6373-12 | CWC8J0976-01 | OverKill | JB6K8S02 | NC01682B(COG48AA)-A-AT-N-C | GGP119.1-BC | 01 | too many o/s 166,4 pm checked od11-A6881 | CPAUTO | Over Kill | O/D不足 | 2022/08/20 02:32:14 | 2022/08/20 02:52:31 | 0.34 | OK | 13597 | 重對針PASS，RANGE 10 UM | NO | NO | NO | YES | YES | NO
EQPRepairJ22109736 | T6373-152 |  | Setup Fail | JB2P2S06 | R63455-A1S-EZ-CP | NAHM72.00-B1 |  | CHUCK TEMPERATURE CONTROL ERROR | A6788 | Setup Fail | Prober異常 prober abnormal | 2022/08/20 02:57:04 | 2022/08/20 03:34:30 | 0.62 | OK | 13597 | PROBER RESET OK | NO | NO | NO | YES | YES | NO
EQPRepairJ22109741 | T6373-33 | CWC7J0K83-01.001 | OverKill | JB1TAS28 | PA5208KF-LT | J67HF | 05 | TOO MANY O/S-A6924 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/20 03:35:29 | 2022/08/20 03:59:00 | 0.39 | OK | 13597 | 機台清針PASS | NO | NO | NO | YES | YES | NO
EQPRepairJ22109749 | T6373-128 | CWC7J0M33-01.001 | OverKill | VB1BVS02 | PA5270D-LT | U38F8.00 | 21 | too many fail bin34-A6786 | CPAUTO | Product Issue | bin34 | 2022/08/20 04:04:08 | 2022/08/20 04:17:23 | 0.22 | OK | 13597 | bin34 reload pgm pass | NO | NO | NO | YES | YES | NO
EQPRepairJ22109751 | T6373-99 | CWC7J0L22-01 | OverKill | JB1TAS27 | PA5208K-LT | QT7WS | 21 | EDA SHOW 2R-A7164 | CPAUTO | Over Kill | 2R | 2022/08/20 04:09:51 | 2022/08/20 05:34:45 | 1.42 | OK | 17674 | 1. 2R 下卡磨針PASS，溫針 RANGE Y=10 (21,4)#21 | NO | NO | NO | YES | YES | NO
EQPRepairJ22109780 | T6373-152 | CWC7J0H67-01 | OverKill | JB2P2S06 | R63455-A1S-EZ-CP | NAHM72.00-B1C | 07 | too many open short-A6786 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/20 06:23:37 | 2022/08/20 07:34:34 | 1.18 | OK | 16631 | 1.夜班處理O/S，CLEAN WAFER時ARM1 VACCUM ERROR，轉掛PROBER DOWN | NO | NO | NO | YES | YES | NO
EQPRepairJ22109792 | T6373-152 |  | Prober Down | JB2P2S06 | R63455-A1S-EZ-CP | NAHM72.00-B1C |  | Prober Down---- 1.夜班處理O/S，CLEAN WAFER時ARM1 VACCUM ERROR，轉掛PROBER DOWN | 16631 | Prober | Loader Down | 2022/08/20 07:35:05 | 2022/08/20 08:27:12 | 0.87 | OK | 16631 | 1.prober reset & setup pass，#07 load again | NO | NO | NO | YES | YES | NO
EQPRepairJ22109804 | T6373-43 | CWC7J0L69-01.003 | Tester Down | JB1UJS03 | PA5129LS_12U80H-LT | P078M.1 | 18 | AlarmCode:669 AlarmMessage:GP-IB COMMUNICATION TIMEOUT ERROR !! | CPAUTO | Tester | 離線(OCT/GPIB error) | 2022/08/20 08:21:20 | 2022/08/20 08:30:23 | 0.15 | OK | 19771 | 重新連線ok | NO | NO | NO | YES | YES | NO
EQPRepairJ22109810 | T6373-07 |  | Setup Fail | JB1TAS16 | PA5208KF-LT | J682P |  | 極低溫降溫異常無法降到-40度 維持在-25度上下 待輪工確認。 | 08341 | Prober | 溫度異常 (Temperature Error) | 2022/08/20 08:34:52 | 2022/08/20 08:36:57 | 0.03 | OK | 13749 | Prober Down--- 1.產線Setup Team 告知改機-40度型號，降溫無法達到目標溫到已超過1小時. 2.轉掛Prober Down 待確認. | NO | NO | NO | YES | YES | NO
EQPRepairJ22109811 | T6373-07 |  | Prober Down |  |  |  |  | Prober Down--- 1.產線Setup Team 告知改機-40度型號，降溫無法達到目標溫到已超過1小時. 2.轉掛Prober Down 待確認. | 13749 | Prober | 面板Down (Monitor Down) | 2022/08/20 08:37:20 | 2022/08/20 15:01:31 | 6.40 | OK | 13749 | 1.Prober 重開後無法正常開啟會跳SVC Call Error! 2.已通知假日值班設備工程協助前往維修OK. 3.還機產線重新掛Setup. | NO | NO | NO | YES | YES | NO
EQPRepairJ22109813 | T6373-21 |  | Prober Down | JB1GKS08 | PA6543B | J6AK6 |  | AlarmCode:1058 AlarmMessage:MODEL NOT FOUND ( FINE ALIGNMENT ) !! | CPAUTO | Prober | Wafer Edge/Mode Not Found | 2022/08/20 08:53:08 | 2022/08/20 09:02:04 | 0.15 | OK | 20231 | remote---- 1.重抓model OK，check die OK，PM(162,1) | NO | NO | NO | YES | YES | NO
EQPRepairJ22109819 | T6373-33 | CWC7J0K83-01.001 | OverKill | JB1TAS28 | PA5208KF-LT | J67HF | 07 | o/s fail-A6954 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/20 09:09:05 | 2022/08/20 09:31:47 | 0.38 | OK | 17680 | remote----使用機台清針，起測穩定 | NO | NO | NO | YES | YES | NO
EQPRepairJ22109827 | T6373-128 | CWC7J0M33-01.001 | OverKill | VB1BVS02 | PA5270D-LT | U38F8.00 | 23 | #21.22 2h-20022 | CPAUTO | Over Kill | 2H | 2022/08/20 09:39:12 | 2022/08/20 09:45:59 | 0.11 | OK | 16631 | 1.同批不同機卡有相同問題，續測 | NO | NO | NO | YES | YES | NO
EQPRepairJ22109832 | T6373-152 | CWC7J0H67-01 | Prober Down | JB2P2S06 | R63455-A1S-EZ-CP | NAHM72.00-B1C | 09 | AlarmCode:4827 AlarmMessage:ARM @@VACUUM ERROR !! | CPAUTO | Prober | 手臂異常(Arm Error) | 2022/08/20 09:48:08 | 2022/08/20 12:44:16 | 2.94 | OK | 19771 | 確認WAFER位置後重開Prober,重新SU OK#9(146.8) | NO | NO | NO | YES | YES | NO
EQPRepairJ22109838 | T6373-13 | CWC8J0871-01.002 | OverKill | JB1GVS19 | NC01682B(COG48AA)-A-AT-N-C | GGR437.1-BC2211 | 23 | too many bin29 | A6962 | Over Kill | Function誤宰 | 2022/08/20 10:04:11 | 2022/08/20 10:55:23 | 0.85 | OK | 17680 | remote----replace #27,position(0,0),mark(47,6) | NO | NO | NO | YES | YES | NO
EQPRepairJ22109848 | T6373-135 | CWC8J0957-01 | OverKill | TB1YLS38 | PA6485E(CP1-CP3) | P0CSG.1 | 17 | o/s-a6571 | CPAUTO | Over Kill | 機台清針後Pass | 2022/08/20 11:05:33 | 2022/08/20 11:13:40 | 0.14 | OK | 16631 | 1.機台清針後Pass | NO | NO | NO | YES | YES | NO
EQPRepairJ22109849 | T6373-34 | CWC7J0K25-01.004 | Corr Fail | JB1BYS10 | PA6587B_15U90H | J696F | 22 | run corr-14212 | CPAUTO | Corr | Other | 2022/08/20 11:05:50 | 2022/08/20 11:08:07 | 0.04 | OK | 17680 | 手動比對PASS | NO | NO | NO | YES | YES | NO
EQPRepairJ22109851 | T6373-13 | CWC8J0871-01.002 | Other | JB1GVS19 | NC01682B(COG48AA)-A-AT-N-C | GGR437.1-BC2211 | 23 | CANNOT MOVE IN, PLEASE CHECK PC JB1GVS27 IF RELEASE | A6668 | Over Kill | Other | 2022/08/20 11:18:56 | 2022/08/20 12:45:43 | 1.45 | OK | 11385 | 1.更換JB1GVS26 Bin30 Function 2.再更換JB1GVS22 PASS (17.5) | NO | NO | NO | YES | YES | NO
EQPRepairJ22109853 | T6373-33 | CWC7J0K83-01.001 | OverKill | JB1TAS28 | PA5208KF-LT | J67HF | 08 | o/s-a6962 | CPAUTO | Over Kill | 配件殘金 | 2022/08/20 11:29:49 | 2022/08/20 12:02:42 | 0.55 | OK | 13749 | 1.slot #08 原配件針尖黑點O/S. 2.更換備卡#S23->#S08 OK 預溫重對針 點測PASS溫針退針壓 PM(147.5) | NO | NO | NO | YES | YES | NO
EQPRepairJ22109859 | T6373-136 | CWC8J0358-01.004 | Prober Down | JB6H1S05 | NC01690F(COG46BA)-A-AT-C | GGY684.1-B1C111 | 22 | MODEL NOT FOUND | A6564 | Prober | Wafer Edge/Mode Not Found | 2022/08/20 12:00:59 | 2022/08/20 12:11:55 | 0.18 | OK | 19771 | 手動抓取OK | NO | NO | NO | YES | YES | NO
EQPRepairJ22109861 | T6373-90 | CWC8J0168-01.001 | OverKill | JB1BYS12 | PA6587B_15U90H | J6969 | 16 | 2R (o/s)-11933 | CPAUTO | Over Kill | O/D不足 | 2022/08/20 12:24:31 | 2022/08/20 12:31:44 | 0.12 | OK | 13749 | Remote--- 1.EDA Show 2R O/S [RW_BIN31>0.6%](Bin 32 > 6EA) 8ea. 2.Check EDA O/S Fail BIN 皆落於Wafer 最外圍 針壓不足造成，測試OD維持8um. 3.確認針況、真位皆無異常，重新對針後 pass 穩定 PLEASE NEXT WAFER CHECK Slot #16 PM(132.0) | NO | NO | NO | YES | YES | NO
EQPRepairJ22109862 | T6373-136 | CWC8J0358-01.004 | Prober Down | JB6H1S05 | NC01690F(COG46BA)-A-AT-C | GGY684.1-B1C112 | 22 | IF FAILED IN LOCK/UNLOCK OF THE MANUAL RING LEVER-A6049 | CPAUTO | Prober | Sensor Error | 2022/08/20 12:29:31 | 2022/08/20 12:37:10 | 0.13 | OK | 16631 | 1.調整OK，作動多次正常 | NO | NO | NO | YES | YES | NO
EQPRepairJ22109863 | T6373-135 | CWC8J0957-01 | OverKill | TB1YLS38 | PA6485E(CP1-CP3) | P0CSG.1 | 18 | too many os-A6564 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/20 12:31:33 | 2022/08/20 13:22:21 | 0.85 | OK | 16631 | 1.磨針 237.4 | NO | NO | NO | YES | YES | NO
EQPRepairJ22109875 | T6373-152 | CWC7J0H67-01 | Prober Down | JB2P2S06 | R63455-A1S-EZ-CP | NAHM72.00-B1C | 09 | temp.control error-a6571 | CPAUTO | Prober | 溫度異常 (Temperature Error) | 2022/08/20 13:16:14 | 2022/08/20 14:52:37 | 1.61 | OK | 16631 | 1.測試中溫控跳掉，重開Prober多次都無法進入主畫面 2.轉掛Prober Down，待確認 | NO | NO | NO | YES | YES | NO
EQPRepairJ22109884 | T6373-129 | CEC8J0033-01 | OverKill | JB1EBS18 | NA01644G(COG88AA)-A-AT-V-C | GGW352.1-2B1C1 | 12 | 2H o/s | A6894 | Over Kill | O/S | 2022/08/20 14:02:43 | 2022/08/20 14:21:53 | 0.32 | OK | 11385 | Remote 處理----EDA H 調整針位清針Range:12 (51.6) | NO | NO | NO | YES | YES | NO
EQPRepairJ22109898 | T6373-152 |  | Prober Down | JB2P2S06 | R63455-A1S-EZ-CP | NAHM72.00-B1C |  | Prober Down---- 1.測試中溫控跳掉，重開Prober多次都無法進入主畫面 2.轉掛Prober Down，待確認 | 16631 | Prober | 溫度異常 (Temperature Error) | 2022/08/20 14:53:15 | 2022/08/20 18:44:39 | 3.86 | OK | 19771 | 1.溫控異常，重開Prober無法進入，常日協助更換硬碟後OK 2.重新SU #25(147.8) | NO | NO | NO | YES | YES | NO
EQPRepairJ22109899 | T6373-129 | CEC8J0033-01 | OverKill | JB1EBS18 | NA01644G(COG88AA)-A-AT-V-C | GGW352.1-2B1C1 | 12 | too many o/s-A6526 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/20 15:02:00 | 2022/08/20 15:27:42 | 0.43 | OK | 20231 | 1.確認針位偏下，上調5um pass，PM(180,1) | NO | NO | NO | YES | YES | NO
EQPRepairJ22109913 | T6373-136 | CWC7J0A50-01.001 | OverKill | JB6H1S05 | NC01690I(COG46BA)-A-AT-C | GGK814.1-BC21 | 11 | probemark NG, probemark shift down | A6921 | Over Kill | PM | 2022/08/20 15:58:39 | 2022/08/20 16:30:39 | 0.53 | OK | 16631 | 1.換片O/S，確認OD不足，CP4查看已有上下貼針痕，重新對針PASS，試打針痕(151.0)(129.5) | NO | NO | NO | YES | YES | NO
EQPRepairJ22109915 | T6373-12 | CWC8J0976-01 | OverKill | JB6K8S02 | NC01682B(COG48AA)-A-AT-N-C | GGP119.1-BC | 17 | RDA SHOW I , CORR PASSED | A6897 | Over Kill | show I | 2022/08/20 16:08:19 | 2022/08/20 16:43:11 | 0.58 | OK | 11385 | Remote 處理----EDA I 溫針重新對針調整針位 (137.6) | NO | NO | NO | YES | YES | NO
EQPRepairJ22109916 | T6373-07 |  | Setup Fail | JB1GVS24 | NB01682F(COG48BA)-A-AT-Z-C | GFSW63.1-B1 |  | 溫控異常無法降至-40度 待輪工確認 | 08341 | Prober | 溫度異常 (Temperature Error) | 2022/08/20 16:11:19 | 2022/08/20 17:30:00 | 1.31 | OK | 13749 | Prober Down--- 1.先前為Prober 重開異常 SVC Call Error ，假日值班人員已修復完成. 2.還機產線改機極低溫-40 度型號時，Prober無法降到目標溫度 停留在-20度左右. 3.重新升溫後降溫仍無法排除，故暫掛Prober Down協調改測. | NO | NO | NO | YES | YES | NO
EQPRepairJ22109924 | T6373-90 | CWC8J0168-01.001 | OverKill | JB1BYS12 | PA6587B_15U90H | J6969 | 18 | O/S | A6921 | Over Kill | 配件殘金 | 2022/08/20 16:56:16 | 2022/08/20 18:12:05 | 1.26 | OK | 13749 | 1.SLOT #18 針尖黑點 o/s 下卡磨針上機pass (135.7) | NO | NO | NO | NO | YES | NO
EQPRepairJ22109925 | T6373-34 | CWC8J0168-01.005 | OverKill | JB1BYS10 | PA6587B_15U90H | J6969 | 07 | EDA SHOWS 2R O/S | A6921 | Probe Card | 針位異常 | 2022/08/20 17:00:17 | 2022/08/20 18:28:51 | 1.48 | OK | 13749 | P/C Down--- 1.JB1BYS10 測試中O/S不穩，磨針後上機仍BIN32 OPEN不穩. 2.點測2分法 Y Range =1um ，確認針痕已上下貼邊擠壓BUMP. 3.故將配件#S10先行送修，無備卡. | NO | NO | NO | YES | YES | NO
EQPRepairJ22109934 | T6373-07 |  | Prober Down | JB1GVS24 | NB01682F(COG48BA)-A-AT-Z-C | GFSW63.1-B1 |  | Prober Down--- 1.先前為Prober 重開異常 SVC Call Error ，假日值班人員已修復完成. 2.還機產線改機極低溫-40 度型號時，Prober無法降到目標溫度 停留在-20度左右. 3.重新升溫後降溫仍無法排除，故暫掛Prober Down協調改測. | 13749 | Prober | 溫度異常 (Temperature Error) | 2022/08/20 17:30:24 | 2022/08/20 18:41:13 | 1.18 | OK | 13749 | 1.已協調PP 改測 OK. 低溫-40 to 高溫105. 2.與產線Setup Team 確認貨卡已到 IDLE 掛Setup. | NO | NO | NO | YES | YES | NO
EQPRepairJ22109940 | T6373-135 | CWC8J0957-01 | OverKill | TB1YLS38 | PA6485E(CP1-CP3) | P0CSG.1 | 25 | TOO MANY O/S CHECK PM 218.5-A6564 | CPAUTO | Over Kill | 機台清針後Pass | 2022/08/20 18:04:07 | 2022/08/20 18:23:27 | 0.32 | OK | 16631 | 1.機台清針後Pass | NO | NO | NO | YES | YES | NO
EQPRepairJ22109948 | T6373-34 |  | Probe Card Down | JB1BYS10 | PA6587B_15U90H | J6969 |  | P/C Down--- 1.JB1BYS10 測試中O/S不穩，磨針後上機仍BIN32 OPEN不穩. 2.點測2分法 Y Range =1um ，確認針痕已上下貼邊擠壓BUMP. 3.故將配件#S10先行送修，無備卡. | 13749 | Probe Card | os | 2022/08/20 18:29:13 | 2022/08/21 20:14:02 | 25.75 | OK | 17091 | 交接早班配件送修回來上機結霜，升溫100 30min 後降溫，重對針*2，二分法調針range:-2~10 pm(173，5) 起測一段穩定還機 | NO | NO | NO | YES | YES | NO
EQPRepairJ22109951 | T6373-90 | CWC8J0168-01.001 | OverKill | JB1BYS12 | PA6587B_15U90H | J6969 | 19 | #17,18 EDA SHOW 2R(o/s)-20021 | CPAUTO | Over Kill | Other | 2022/08/20 18:45:45 | 2022/08/20 18:49:03 | 0.06 | OK | 13749 | REMOTE---TA 補掛2r o/s ，輪工SLOT #19 已下卡磨針上機測試 pass 穩定 還機觀察. | NO | NO | NO | YES | YES | NO
EQPRepairJ22109956 | T6373-136 | CWC8J0508-01.001 | Setup Fail | JB8Q6S05 | NA01761D(COF94AA)-A-Y-C | GGS421.1-B | 06 | DUT 4 BIN 53 (no other P/C available), P/C NEED TO VERIFY BY AE | A5548 | Setup Fail | Function Fail | 2022/08/20 19:29:49 | 2022/08/20 20:08:36 | 0.65 | OK | 17674 | 1. JB8Q6S05 待驗卡改機Dut4-Bin53 Function Fail，清調針、重MNT、重Load PGM 無效，點測固定Dut4-Bin53，故將配件送修無備卡 | NO | NO | NO | YES | YES | NO
EQPRepairJ22109968 | T6373-136 |  | Probe Card Down | JB8Q6S05 | NA01761D(COF94AA)-A-Y-C | GGS421.1-B |  | Probe Card Down----- 1. JB8Q6S05 待驗卡改機Dut4-Bin53 Function Fail，清調針、重MNT、重Load PGM 無效，點測固定Dut4-Bin53，故將配件送修無備卡 | 17674 | Probe Card | Other | 2022/08/20 20:09:11 | 2022/08/20 21:27:38 | 1.31 | OK | 17674 | P/C DOWN ---- 1. 產線通知 IDEL | NO | NO | NO | YES | YES | NO
EQPRepairJ22109974 | T6373-07 | CWC7J0G38-01 | Corr Fail | JB1TAS16 | PA5208K-LT | QT7YW | 01 | Corr Fail | A5929 | Corr | Corr Fail | 2022/08/20 20:44:18 | 2022/08/20 21:14:46 | 0.51 | OK | 17542 | Corr Fail --- 自動比對Fail，上機回收，回收後比對Pass | NO | NO | NO | YES | YES | NO
EQPRepairJ22109979 | T6373-29 | CWC8J0969-01 | OverKill | KB2BNS09 | ND01733D(COF94AA)-A-C | H2PU18.3-22BC1 | 23 | eda conts 2R no corr wafer | A6404 | Over Kill | SHOW H | 2022/08/20 21:08:53 | 2022/08/20 21:33:34 | 0.41 | OK | 13597 | 不同機台測試一樣，續測 | NO | NO | NO | YES | YES | NO
EQPRepairJ22109989 | T6373-128 |  | Prober Down | JB1FDS08 | PA1638K-10E-B-LT | RDKAA0000 |  | AlarmCode:1060 AlarmMessage:CANNOT MEASURE INDEX SIZE !! | CPAUTO | Prober | Wafer Edge/Mode Not Found | 2022/08/20 21:59:01 | 2022/08/20 22:48:24 | 0.82 | OK | 13510 | 重抓ok | NO | NO | NO | YES | YES | NO
EQPRepairJ22110001 | T6373-128 | CWC8J0186-01 | OverKill | JB1FDS08 | PA1638K-10E-B-LT | RDKAA0000 | 20 | FAIL BIN BECOME TO O/S | A6148 | Over Kill | O/S誤宰 | 2022/08/20 23:21:49 | 2022/08/20 23:35:47 | 0.23 | OK | 11348 | Remote 處理----調針PASS，Range:10um | NO | NO | NO | YES | YES | NO
EQPRepairJ22110005 | T6373-111 | CWC8J0770-01.001 | Prober Down | JB1Q4S14 | PA6542C | J6C10 | 01 | MODDEL NOT FOUND,ROUGH ALIGNMENT | A6926 | Prober | Wafer Edge/Mode Not Found | 2022/08/20 23:43:09 | 2022/08/20 23:58:40 | 0.26 | OK | 17542 | Remote 處理 ----- 重抓Model OK(model not found)，Check Die ok，P/M(11,4)ok | NO | NO | NO | YES | YES | NO
EQPRepairJ22110008 | T6373-47 |  | Prober Down | JB1HKS06 | NA01687G(COG48BA)-A-AT-N-C | GGR354.1-B21C2 |  | AlarmCode:1057 AlarmMessage:MODEL NOT FOUND ( ROUGH ALIGNMENT ) !! | CPAUTO | Prober | Wafer Edge/Mode Not Found | 2022/08/20 23:59:58 | 2022/08/21 00:08:01 | 0.13 | OK | 17674 | 1. 換片發生Prober down ，check die 針位貼下 (36,0) | NO | NO | NO | YES | YES | NO
EQPRepairJ22110022 | T6373-16 | CWC8J0908-01 | OverKill | VB405S09 | PA5208KF-LT | J6G42 | 11 | #09.10 eda show 2R o/s-A6148 | CPAUTO | Over Kill | Show R | 2022/08/21 01:12:07 | 2022/08/21 01:27:30 | 0.26 | OK | 17542 | Show R，[RW_BIN28>0.3%](Bin 29 > 2EA) 5，零星O/S，機台清針，起測一段PASS穩定，請產線先行續測待觀察 | NO | NO | NO | YES | YES | NO
EQPRepairJ22110037 | T6373-35 | CWC8J0871-01.004 | OverKill | JB1GVS24 | NC01682B(COG48AA)-A-AT-N-C | GGR437.1-BC2212 | 13 | #11.12 eda show "2H" no corr wafer | A6148 | Over Kill | SHOW H | 2022/08/21 02:48:17 | 2022/08/21 02:52:46 | 0.07 | OK | 13597 | 前流程造成，續測 | NO | NO | NO | YES | YES | NO
EQPRepairJ22110041 | T6373-141 | CWC8J0348-01.004 | Corr Fail | JB1HKS10 | NA01687G(COG48BA)-A-AT-N2-C | GGR356.1-B11 | 08 | too many os | A5929 | Over Kill | O/S誤宰 | 2022/08/21 03:17:09 | 2022/08/21 03:53:48 | 0.61 | OK | 13597 | 更換配件#09，RANGE 10 UM | NO | NO | NO | YES | YES | NO
EQPRepairJ22110052 | T6373-141 | CWC8J0348-01.004 | Corr Fail | JB1HKS09 | NA01687G(COG48BA)-A-AT-N2-C | GGR356.1-B11 | 08 | Corr Fail | A6963 | Corr | Corr Fail | 2022/08/21 04:35:55 | 2022/08/21 05:10:58 | 0.58 | OK | 20276 | EDA比對 Corr pass | NO | NO | NO | YES | YES | NO
EQPRepairJ22110060 | T6373-21 | CWC8J0752-01.002 | OverKill | JB1GKS08 | PA6543B | J6AK6 | 25 | TOO MANY O/S PM CHECK 157.6 OD11-A7024 | CPAUTO | Over Kill | O/D不足 | 2022/08/21 05:33:15 | 2022/08/21 05:45:04 | 0.20 | OK | 20276 | 增加OD 11um>13um range 8 P/M(41.7) | NO | NO | NO | YES | YES | NO
EQPRepairJ22110072 | T6373-35 | CWC8J0871-01.004 | Tester Down | JB1GVS24 | NC01682B(COG48AA)-A-AT-N-C | GGR437.1-BC2212 | 14 | AlarmCode:669 AlarmMessage:GP-IB COMMUNICATION TIMEOUT ERROR !! | CPAUTO | Tester | 離線(OCT/GPIB error) | 2022/08/21 06:20:50 | 2022/08/21 06:38:21 | 0.29 | OK | 13597 | 無訊息，續測 | NO | NO | NO | YES | YES | NO
EQPRepairJ22110074 | T6373-29 | CWC8J0189-01.009 | Prober Down | JB1GVS21 | NC01682A(COG48AA)-A-AT-N2-C | GGS284.1-B1 | 03 | MODEL NOT FOUND (ROUGH ALIGNMENT)-A7024 | CPAUTO | Prober | Wafer Edge/Mode Not Found | 2022/08/21 06:22:38 | 2022/08/21 06:32:04 | 0.16 | OK | 13597 | 重抓ok | NO | NO | NO | YES | YES | NO
EQPRepairJ22110081 | T6373-16 | CWC8J0815-01.001 | OverKill | VB405S09 | PA5208K-LT | QT7YT | 10 | cont EDA SHOW 2R-A6786 | CPAUTO | Over Kill | 2R | 2022/08/21 06:42:58 | 2022/08/21 07:00:55 | 0.30 | OK | 17674 | 1. 2R [RW_BIN28>0.3%](Bin 29 > 2EA) 5(retest bin29)，機台清針Pass 良率99% | NO | NO | NO | YES | YES | NO
EQPRepairJ22110085 | T6373-33 | CWC7J0K83-01.001 | OverKill | JB1TAS23 | PA5208KF-LT | J67HF | 16 | too many o/s-A6924 | CPAUTO | Over Kill | O/S | 2022/08/21 07:06:00 | 2022/08/21 07:37:21 | 0.52 | OK | 20230 | 1.O/S 2.機台清針,二分法調針YRG=14,XRG=40,PM(47.6) 3.PASS | NO | NO | NO | YES | YES | NO
EQPRepairJ22110088 | T6373-15 | CWC8J0A45-01 | OverKill | JB6H1S04 | NC01690I(COG46BA)-A-AT-C | GGT531.1-BC | 18 | eda 2 H o/s-A7033 | CPAUTO | Over Kill | EDA 2H | 2022/08/21 07:12:18 | 2022/08/21 07:56:00 | 0.73 | OK | 20230 | 1.EDA 2H,BIN 83(零星O/S) 2.機台清針,上調2UM,PM(127.5) 3.PASS | NO | NO | NO | YES | YES | NO
EQPRepairJ22110093 | T6373-91 | CWC2J1512-01.002 | Setup Fail | TB1R8S82 | PA5223M1_12U60H-LT | PXCG6.1 | 13 | contact fail due to o/s, already change pc | A6660 | Setup Fail | SETUP FAIL | 2022/08/21 07:40:36 | 2022/08/21 08:24:56 | 0.74 | OK | 13581 | 1.SETUP FAIL O/S不過，接手前往處理，確認機台結霜 2.升溫烘烤，轉掛PROBER DOWN | NO | NO | NO | NO | YES | NO
EQPRepairJ22110101 | T6373-91 |  | Prober Down | TB1R8S82 | PA5223M1_12U60H-LT | PXCG6.1 |  | 1.SETUP FAIL O/S不過，接手前往處理，確認機台 結霜 2.升溫烘烤，轉掛PROBER DOWN | 13581 | Prober | 改機fail o/s不過,機台結霜 | 2022/08/21 08:25:37 | 2022/08/21 09:53:55 | 1.47 | OK | 13581 | 1.改機fail o/s不過,產線掛當 2.接手處理前往處理確認機台結霜 3.升溫烘烤後，降溫重新setup PASS 4.請產線改 setup 04 | NO | NO | NO | NO | YES | NO
EQPRepairJ22110104 | T6373-16 |  | Prober Down | VB405S09 | PA5208K-LT | QT7YT |  | AlarmCode:1995 AlarmMessage:Faied to upload log file | CPAUTO | Prober | Other | 2022/08/21 08:31:27 | 2022/08/21 08:43:15 | 0.20 | OK | 13581 | 1.前往確認測試正常，續觀察 | NO | NO | NO | NO | YES | NO
EQPRepairJ22110110 | T6373-111 | CWC7J0C21-01.013 | Corr Fail | JB1Q4S14 | PA6542C | J68S1 | 06 | Corr Fail | A6660 | Corr | corr fail | 2022/08/21 08:57:23 | 2022/08/21 08:59:24 | 0.03 | OK | 20230 | 手動比對pass | NO | NO | NO | YES | YES | NO
EQPRepairJ22110118 | T6373-136 |  | Prober Down | JB6H1S05 | NC01690F(COG46BA)-A-AT-C | GGY201.1-B12C1222 |  | AlarmCode:1057 AlarmMessage:MODEL NOT FOUND ( ROUGH ALIGNMENT ) !! | CPAUTO | Prober | Wafer Edge/Mode Not Found | 2022/08/21 09:20:36 | 2022/08/21 09:33:45 | 0.22 | OK | 07764 | Remote處理----重抓 ok #06(2.3) | NO | NO | NO | YES | YES | NO
EQPRepairJ22110120 | T6373-152 | CWC8J0166-01 | OverKill | JB2P2S06 | R63455-A1S-EZ-CP | NAHM66.00-B1C | 24 | too many o/s-A6919 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/21 09:27:46 | 2022/08/21 09:51:19 | 0.39 | OK | 07764 | open 執行soaktime 後pass | NO | NO | NO | YES | YES | NO
EQPRepairJ22110124 | T6373-16 | CWC8J0815-01.001 | OverKill | VB405S09 | PA5208K-LT | QT7YT | 24 | TOO MANY O/S | A6919 | Over Kill | O/S誤宰 | 2022/08/21 09:35:17 | 2022/08/21 09:56:54 | 0.36 | OK | 13222 | REMOTE---o/s 1. 機台清針 測試 一段ok | NO | NO | NO | YES | YES | NO
EQPRepairJ22110129 | T6373-43 |  | Tester Down | JB1UJS03 | PA5129LS_12U80H-LT | P078M.1 |  | TESTER ng-A6919 | CPAUTO | Product Issue | 自動化異常 | 2022/08/21 10:08:54 | 2022/08/21 10:35:01 | 0.44 | OK | 13581 | 1.登出重新LOAD ok | NO | NO | NO | NO | NO | NO
EQPRepairJ22110132 | T6373-90 | CWC7J0M27-01.006 | OverKill | JB1BYS12 | PA6587B_15U90H | J696G | 07 | EDA SHOW 2R-A6131 | CPAUTO | Over Kill | 2R | 2022/08/21 10:27:35 | 2022/08/21 10:46:20 | 0.31 | OK | 13222 | 機台清針OK (133.7) | NO | NO | NO | YES | YES | NO
EQPRepairJ22110133 | T6373-11 | CWC8J0295-01.006 | Setup Fail | JB6H1S06 | NC01690G(COG48AA)-A-AT-N4-C | GGG486.1-B | 13 | TOO LARGE NEEDLE, CHANGE TO MBD AS PER AE | A6660 | Setup Fail | 對針異常 | 2022/08/21 10:29:14 | 2022/08/21 12:57:09 | 2.47 | OK | 13581 | 1.改機fail(JB6H1S06)對針異常, 2.接收確認對針時1-1對會FAIL,重試多次相同 3.更換#07配件上機對針ok，請產線重新掛SETUP 02 | NO | NO | NO | NO | NO | NO
EQPRepairJ22110136 | T6373-111 | CWC8J0618-01.004 | OverKill | JB1Q4S14 | PA6542C | J6C0W | 02 | O/S-07854 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/21 10:39:15 | 2022/08/21 11:32:28 | 0.89 | OK | 13581 | 1.固定PIN O/S 清調針無效,重MNT pass(128.5) | NO | NO | NO | NO | YES | NO
EQPRepairJ22110144 | T6373-136 | CWC8J0503-01.004 | OverKill | JB6H1S05 | NC01690F(COG46BA)-A-AT-C | GGY201.1-B12C1222 | 08 | CONTINUES WAFER EDA SHOW 2I,,WARNING OPEN SHORT MODE TEST-a6133 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/21 11:19:23 | 2022/08/21 11:55:12 | 0.60 | OK | 07764 | EDA I , 確認針痕/加壓5UM #8(180.3) | NO | NO | NO | YES | YES | NO
EQPRepairJ22110147 | T6373-152 |  | Prober Down | JB2P2S06 | R63455-A1S-EZ-CP | NAHM66.00-B1C |  | cassette error-A6919 | CPAUTO | Prober | Other | 2022/08/21 11:39:51 | 2022/08/21 12:28:18 | 0.81 | OK | 07764 | 重放CASSETTE無效 開蓋後 OK , 續測等測試完在請輪工幫忙關蓋 | NO | NO | NO | YES | YES | NO
EQPRepairJ22110154 | T6373-29 | CWC8J0189-01.007 | OverKill | JB1GVS21 | NC01682A(COG48AA)-A-AT-N2-C | GGS284.1-B1C2211121 | 10 | TOO MANY OPEN SHORT-a6133 | CPAUTO | Over Kill | o/s | 2022/08/21 12:09:44 | 2022/08/21 13:09:27 | 1.00 | OK | 13222 | o/s 機台清針、調針pass 確認wafer 有針偏狀況，尚未測試位子也有，通知站長 初判前流程造成，針位無異常(47.6)(4.4) | NO | NO | NO | YES | YES | NO
EQPRepairJ22110159 | T6373-136 |  | Prober Down | JB6H1S05 | NC01690F(COG46BA)-A-AT-C | GGY201.1-B12C1222 |  | AlarmCode:1057 AlarmMessage:MODEL NOT FOUND ( ROUGH ALIGNMENT ) !! | CPAUTO | Prober | Wafer Edge/Mode Not Found | 2022/08/21 12:45:47 | 2022/08/21 13:00:45 | 0.25 | OK | 07764 | Remote處理----重抓model ok check die ok #10(165.4) | NO | NO | NO | YES | YES | NO
EQPRepairJ22110164 | T6373-43 | CWC7J0L69-01.003 | Tester Down | JB1UJS03 | PA5129LS_12U80H-LT | P078M.1 | 20 | AlarmCode:669 AlarmMessage:GP-IB COMMUNICATION TIMEOUT ERROR !! | CPAUTO | Tester | 離線(OCT/GPIB error) | 2022/08/21 13:08:30 | 2022/08/21 13:22:55 | 0.24 | OK | 07764 | Remote處理----重新連線 ok | NO | NO | NO | YES | YES | NO
EQPRepairJ22110168 | T6373-39 | CWC7J0M32-01 | Prober Down | JB1UJS13 | PA5129LS_12U80H-LT | P060Q.1 | 09 | AlarmCode:1995 AlarmMessage:Faied to upload log file | CPAUTO | Prober | Other | 2022/08/21 13:13:52 | 2022/08/21 13:25:13 | 0.19 | OK | 07764 | Remote處理----alarm取消後續測 | NO | NO | NO | YES | YES | NO
EQPRepairJ22110174 | T6373-91 | CWBBJ0N96-01.001 | OverKill | TB1R8S82 | PA5223M1_12U60H-LT | PX0R2.1 | 08 | TOO MANY O/S-A6317 | CPAUTO | Over Kill | O/S | 2022/08/21 13:35:48 | 2022/08/21 13:58:46 | 0.38 | OK | 20230 | 1.O/S 2.清調針,PM(119.6),退15UM 3.PASS | NO | NO | NO | YES | YES | NO
EQPRepairJ22110177 | T6373-152 | CWC8J0165-01 | OverKill | JB2P2S06 | R63455-A1S-EZ-CP | NAHN27.00-B2C | 18 | too many o/s-A6919 | CPAUTO | Prober | Other | 2022/08/21 13:52:54 | 2022/08/21 14:53:57 | 1.02 | OK | 07764 | 1.open 修機執行clean wafer 造成arm1 vacuum error 2.查看維修紀錄發生多次 , 轉掛prober down | NO | NO | NO | YES | YES | NO
EQPRepairJ22110182 | T6373-43 | CWC7J0L69-01.003 | Tester Down | JB1UJS03 | PA5129LS_12U80H-LT | P078M.1 | 20 | AlarmCode:669 AlarmMessage:GP-IB COMMUNICATION TIMEOUT ERROR !! | CPAUTO | Tester | 離線(OCT/GPIB error) | 2022/08/21 14:08:31 | 2022/08/21 14:45:04 | 0.61 | OK | 20075 | 測試至#20(62,4)RTE,重新連線ok | NO | NO | NO | YES | YES | NO
EQPRepairJ22110188 | T6373-147 | CWC8J0216-01 | OverKill | JB1Q4S18 | PA6542C | J683M | 20 | o/s-11550 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/21 14:28:29 | 2022/08/21 15:04:45 | 0.60 | OK | 17229 | 1.針壓過重,調整針壓CLEAN WAFER二分法調整RANGE:15UM落點-2.2 PASS. | NO | NO | NO | YES | YES | NO
EQPRepairJ22110190 | T6373-151 | CWC7J0J89-01 | OverKill | JB2P2S09 | R63455-A1S-EZ-CP | NAHN28.00-B1C | 07 | TOO MANY O/S-A6919 | CPAUTO | Over Kill | O/S | 2022/08/21 14:34:06 | 2022/08/21 15:16:04 | 0.70 | OK | 20230 | 1.O/S 2.清調針,RG=11,PM(56.8) 3.PASSSS | NO | NO | NO | YES | YES | NO
EQPRepairJ22110196 | T6373-91 | CWBBJ0N96-01.001 | OverKill | TB1R8S82 | PA5223M1_12U60H-LT | PX0R2.1 | 08 | TOO MANY O/S-A6317 | CPAUTO | Over Kill | O/S | 2022/08/21 14:48:07 | 2022/08/21 14:58:17 | 0.17 | OK | 20230 | 1.O/S 2.清調針RG=8,PM(17.5) 3.PASS | NO | NO | NO | YES | YES | NO
EQPRepairJ22110197 | T6373-152 |  | Prober Down | JB2P2S06 | R63455-A1S-EZ-CP | NAHN27.00-B2C |  | 1.open 修機執行clean wafer 造成arm1 vacuum error 2.查看維修紀錄發生多次 , 轉掛prober down | 07764 | Prober | arm1 vacuum error | 2022/08/21 14:54:34 | 2022/08/21 15:41:19 | 0.78 | OK | 13222 | 1.open 修機執行clean wafer 造成arm1 vacuum error 2.查看維修紀錄發生多次 , 轉掛prober down 3.更換Clean wafer 設定清針參數OK 上機嚐試清針無異常 ，重新對針SETUP OK (56.8) | NO | NO | NO | YES | YES | NO
EQPRepairJ22110198 | T6373-47 | CWC8J0124-01.001 | Corr Fail | JB1HKS06 | NA01687G(COG48BA)-A-AT-N2-C | GGS268.1-B2 | 14 | CORR FAIL | A7086 | Corr | CORR FAIL | 2022/08/21 14:58:58 | 2022/08/21 15:01:16 | 0.04 | OK | 20230 | 手動比對PASS | NO | NO | NO | YES | YES | NO
EQPRepairJ22110199 | T6373-39 | CWC7J0M32-01 | OverKill | JB1UJS13 | PA5129LS_12U80H-LT | P060Q.1 | 10 | o/s-11550 | CPAUTO | Over Kill | O/S | 2022/08/21 15:18:49 | 2022/08/21 15:33:29 | 0.24 | OK | 20230 | 1.O/S 2.清調針,RG=14.PM(115.5) 3.PASS | NO | NO | NO | YES | YES | NO
EQPRepairJ22110203 | T6373-43 | CWC7J0L69-01.003 | Tester Down | JB1UJS03 | PA5129LS_12U80H-LT | P078M.1 | 20 | AlarmCode:669 AlarmMessage:GP-IB COMMUNICATION TIMEOUT ERROR !! | CPAUTO | Tester | 離線(OCT/GPIB error) | 2022/08/21 15:27:52 | 2022/08/21 15:52:52 | 0.42 | OK | 20230 | 1.測試至#20(69,4)RTE,重新連線ok 2.存log,Tester無異常訊息 | NO | NO | NO | YES | YES | NO
EQPRepairJ22110210 | T6373-47 | CWC8J0391-01 | OverKill | JB1HKS06 | NA01687G(COG48BA)-A-AT-N2-C | GGS419.1-B11C1111 | 02 | o/s-11316 | CPAUTO | Over Kill | Other | 2022/08/21 16:04:04 | 2022/08/21 16:39:38 | 0.59 | OK | 07764 | Remotet處理---open 修改清針參數 | NO | NO | NO | YES | YES | NO
EQPRepairJ22110212 | T6373-52 | CWC7J0K72-01.005 | OverKill | TB1D3S04 | R63319C0EQCPQV-P | AYP075000-B11C | 11 | too many failed bin 36-A6919 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/21 16:06:56 | 2022/08/21 17:14:47 | 1.13 | OK | 20075 | 退內壓25UM,RUN SOAK TIME OK | NO | NO | NO | YES | YES | NO
EQPRepairJ22110215 | T6373-43 | CWC7J0L69-01.003 | Prober Down | JB1UJS03 | PA5129LS_12U80H-LT | P078M.1 | 20 | AlarmCode:9192 AlarmMessage:Prober got abnormal values of Touch-Sensor. | CPAUTO | Prober | 清針異常(Clean error) | 2022/08/21 16:20:36 | 2022/08/21 16:36:32 | 0.27 | OK | 07764 | Remotet處理----重新執行一次清針後 ok | NO | NO | NO | YES | YES | NO
EQPRepairJ22110220 | T6373-07 | CWC8J0866-01 | OverKill | JB1TAS16 | PA5208KF-LT | J6A1F | 01 | TOO MANY OPEN SHORT-a6133 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/21 16:43:27 | 2022/08/21 17:27:41 | 0.74 | OK | 17229 | 1.清針二分法調針RANGE:11um落點(5.4)起測pass. | NO | NO | NO | YES | YES | NO
EQPRepairJ22110232 | T6373-135 | CWC8J0865-01 | OverKill | TB1YLS38 | PA6485E(CP1-CP3) | P0CSA.1 | 02 | particle-11550 | CPAUTO | Over Kill | O/S | 2022/08/21 17:25:39 | 2022/08/21 17:42:05 | 0.27 | OK | 13222 | 機台清針OK (181.6) | NO | NO | NO | YES | YES | NO
EQPRepairJ22110235 | T6373-43 | CWC7J0L69-01.003 | Tester Down | JB1UJS03 | PA5129LS_12U80H-LT | P078M.1 | 21 | AlarmCode:669 AlarmMessage:GP-IB COMMUNICATION TIMEOUT ERROR !! | CPAUTO | Tester | 離線(OCT/GPIB error) | 2022/08/21 17:33:47 | 2022/08/21 17:51:59 | 0.30 | OK | 13222 | 1.測試至#21(103.1)RTE,重新連線ok 2.存log,Tester無異常訊息 | NO | NO | NO | YES | YES | NO
EQPRepairJ22110246 | T6373-07 | CWC8J0866-01 | OverKill | JB1TAS16 | PA5208KF-LT | J6A1F | 01 | TOO MANY O/S-A6131 | CPAUTO | Over Kill | 機台清針後Pass | 2022/08/21 18:01:54 | 2022/08/21 18:23:17 | 0.36 | OK | 13222 | 機台清針後Pass (148.5) | NO | NO | NO | YES | YES | NO
EQPRepairJ22110247 | T6373-52 | CWC7J0K72-01.005 | OverKill | TB1D3S04 | R63319C0EQCPQV-P | AYP075000-B11C | 11 | too many fail bin36-A7086 | CPAUTO | Over Kill | bin36 | 2022/08/21 18:02:04 | 2022/08/21 19:03:20 | 1.02 | OK | 13222 | 配件#04測試中bin36 清、調針無效，更換配件#05 上機pass (259.6)，配件#04 送修 | NO | NO | NO | YES | YES | NO
EQPRepairJ22110250 | T6373-152 |  | Other | JB2P2S06 | R63455-A1S-EZ-CP | NAHN27.00-B2C |  | PLEASE HELP CLOSE THE CASSETTE - FOSB-A6919 | CPAUTO | Other | Other | 2022/08/21 18:11:05 | 2022/08/21 18:44:06 | 0.55 | OK | 07764 | 關蓋 ok | NO | NO | NO | YES | YES | NO
EQPRepairJ22110254 | T6373-13 | CWC8J0622-01.003 | OverKill | JB6K8S03 | NC01682A(COG48AA)-A-AT-N2-C | GGT630.1-BC12 | 12 | o/s-a6962 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/21 18:32:31 | 2022/08/21 18:59:02 | 0.44 | OK | 13581 | 1.換批od不足，重新對針PASS(137.6) | NO | NO | NO | NO | YES | NO
EQPRepairJ22110257 | T6373-21 |  | Tester Down | JB1GKS08 | PA6543B | J6AC8 |  | FAILED TO LOAD TESTER-a6133 | CPAUTO | Other | Other | 2022/08/21 18:40:30 | 2022/08/21 19:12:44 | 0.54 | OK | 13222 | Other--- 1. CBVP Load程式異常，開啟CBSETUP LOAD ，NOT INIT，執行DATA RECOVER 重LOAD程式OK | NO | NO | NO | YES | YES | NO
EQPRepairJ22110259 | T6373-01 | CWC7J0G38-01 | Corr Fail | TB1R7S28 | PA5137E1-LT | P0CLR.1 | 01 | corr-A6942 | CPAUTO | Corr | CORR | 2022/08/21 18:43:16 | 2022/08/21 19:59:16 | 1.27 | OK | 17091 | CORR FAIL 查看EDA僅差一顆，請產線load回後 確認針位OK，清針後回收無效，點測P<>P F<>F穩定，此CORR已RUN多次先行退片還機生產 | NO | NO | NO | YES | YES | NO
EQPRepairJ22110260 | T6373-47 | CWC8J0291-01 | OverKill | JB1HKS06 | NA01687F(COG48BA)-A-AT-C | GGU014.1-BC11 | 10 | open/short too many probemark up and down-A6858 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/21 18:43:51 | 2022/08/21 19:55:05 | 1.19 | OK | 20182 | 1.更換備卡#10，上機PASS 2.二分法RG-8 PM(35.4) | NO | NO | NO | YES | YES | NO
EQPRepairJ22110265 | T6373-111 | CWC8J0618-01.004 | OverKill | JB1Q4S14 | PA6542C | J6C0W | 07 | too many o/s-A6920 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/21 19:08:40 | 2022/08/21 19:38:37 | 0.50 | OK | 10897 | 固定pin，機台清針pass，針位確認ok(130.0) | NO | NO | NO | NO | YES | NO
EQPRepairJ22110275 | T6373-01 |  | Other | TB1R7S28 | PA5137E1-LT | P0CLR.1 |  | 3D CLEAN-A6898 | CPAUTO | Other | Other | 2022/08/21 20:04:45 | 2022/08/21 20:14:04 | 0.16 | OK | 20277 | 輪工處理---- 1.3D Setting ok | NO | NO | NO | NO | YES | NO
EQPRepairJ22110290 | T6373-90 | CWC7J0M27-01.006 | OverKill | JB1BYS12 | PA6587B_15U90H | J696G | 13 | continuous eda SHOW 2R O.S #11.12 | A6675 | Over Kill | O/S誤宰 | 2022/08/21 20:40:15 | 2022/08/21 21:04:04 | 0.40 | OK | 20277 | Remote---- 1.Wafer右下針壓不足 Range:14 機台清針 續測觀察 | NO | NO | NO | YES | YES | NO
EQPRepairJ22110292 | T6373-13 | CWC7J0600-01.001 | Prober Down | JB6K8S03 | NC01682A(COG48AA)-A-AT-N2-C | GGT630.1-BC12 | 21 | Cannot Measure Index Size-A318 | CPAUTO | Prober | Wafer Edge/Mode Not Found | 2022/08/21 20:46:08 | 2022/08/21 20:58:33 | 0.21 | OK | 10897 | Mode Not Found/check die/針痕確認 | NO | NO | NO | YES | YES | NO
EQPRepairJ22110299 | T6373-47 | CWC8J0134-01.003 | OverKill | JB1HKS10 | NA01687F(COG48BA)-A-AT-C | GGU014.1-BC11 | 06 | CANNOT MEASURE INDEX SIZE! | A6810 | Prober | Wafer Edge/Mode Not Found | 2022/08/21 20:59:10 | 2022/08/21 21:20:16 | 0.35 | OK | 17091 | remote----手動抓ok check die ok pm(147,4) | NO | NO | NO | YES | YES | NO
EQPRepairJ22110302 | T6373-121 | CWC8J0A25-01.001 | OverKill | TB1INS25 | NA01404H-A-AT-C | GGV030.1-BC2 | 01 | O/S TOO MANY-A6318 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/21 21:12:20 | 2022/08/21 21:38:00 | 0.43 | OK | 09108 | 1.O/S 重新對針 SETUP PASS range:10um (56.6) | NO | NO | NO | YES | YES | NO
EQPRepairJ22110303 | T6373-07 | CWC8J0866-01 | OverKill | JB1TAS16 | PA5208KF-LT | J6A1F | 02 | too many o/s-A6898 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/21 21:14:12 | 2022/08/21 21:49:35 | 0.59 | OK | 17091 | 機台清針+二分法調針R:-13~1 起測一段穩定還機 | NO | NO | NO | YES | YES | NO
EQPRepairJ22110304 | T6373-13 | CWC7J0600-01.001 | Prober Down | JB6K8S03 | NC01682A(COG48AA)-A-AT-N2-C | GGT630.1-BC12 | 21 | CANNOT MEASURE INDEX SIZE-A6599 | CPAUTO | Prober | Wafer Edge/Mode Not Found | 2022/08/21 21:20:47 | 2022/08/21 22:04:13 | 0.72 | OK | 09108 | Prober Down----重抓Model OK(model not found) (140.6) | NO | NO | NO | YES | YES | NO
EQPRepairJ22110313 | T6373-111 |  | Prober Down | JB1Q4S14 | PA6542C | J6C0W |  | AlarmCode:1060 AlarmMessage:CANNOT MEASURE INDEX SIZE !! | CPAUTO | Prober | Wafer Edge/Mode Not Found | 2022/08/21 21:48:51 | 2022/08/21 22:09:25 | 0.34 | OK | 10897 | Mode Not Found/check die ok | NO | NO | NO | NO | YES | NO
EQPRepairJ22110317 | T6373-07 | CWC8J0866-01 | OverKill | JB1TAS16 | PA5208KF-LT | J6A1F | 02 | ALWAYS ALARM OS | A6142 | Over Kill | O/S誤宰 | 2022/08/21 22:41:27 | 2022/08/21 23:41:38 | 1.00 | OK | 10110 | 下卡磨針上機pass | NO | NO | NO | YES | YES | NO
EQPRepairJ22110319 | T6373-151 | CWC7J0A18-01.003 | OverKill | JB252S02 | TD7850-A0S-AZ-CP | SA2PK-B2C | 16 | #16 too many functin fail and o/s | A6441 | Over Kill | O/S誤宰 | 2022/08/21 22:50:01 | 2022/08/21 23:21:31 | 0.53 | OK | 20277 | Remote---- 1.Bin20 交叉點測Pass-->Pass，Fail-->Fail穩定，請產線Run Corr驗證，續測觀察 | NO | NO | NO | YES | YES | NO
EQPRepairJ22110323 | T6373-33 | CWC7J0K83-01.001 | OverKill | JB1TAS23 | PA5208KF-LT | J67HF | 22 | many o/s (147,5)od-10-A6118 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/21 22:55:46 | 2022/08/21 23:16:17 | 0.34 | OK | 09108 | 1.O/S 清針 調針PASS RANGE:12UM (45.6) | NO | NO | NO | YES | YES | NO
EQPRepairJ22110333 | T6373-121 | CWC8J0A25-01.001 | OverKill | TB1INS25 | NA01404H-A-AT-C | GGV030.1-BC2 | 02 | o/s too many-A6318 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/21 23:29:27 | 2022/08/22 00:02:32 | 0.55 | OK | 20182 | 1.清針，重對針PASS 2.RG-10 PM(140.6) | NO | NO | NO | YES | YES | NO
EQPRepairJ22110336 | T6373-70 |  | Setup Fail | JB1Q4S17 | PA6542C | J6C0W |  | setup failed due to soaktime correct Z up limit over-A6943 | CPAUTO | Other | Other | 2022/08/21 23:54:53 | 2022/08/22 00:10:01 | 0.25 | OK | 20182 | 機台自動抓取OK | NO | NO | NO | YES | YES | NO
EQPRepairJ22110339 | T6373-43 | CWC6J0747-01 | Tester Down | JB1UJS03 | PA5129LS_12U80H-LT | P047F.1 | 10 | AlarmCode:669 AlarmMessage:GP-IB COMMUNICATION TIMEOUT ERROR !! | CPAUTO | Tester | 離線(OCT/GPIB error) | 2022/08/22 00:20:27 | 2022/08/22 00:28:41 | 0.14 | OK | 20277 | Remote---- 1.無異常訊息 Soaktime後自動起測，無異常，續測 | NO | NO | NO | YES | YES | NO
EQPRepairJ22110342 | T6373-13 |  | Setup Fail | JB1GVS22 | NC01682B(COG48AA)-A-AT-N-C | GGR939.1-B |  | BARCODE NG. AE COMMAND TO MBD | A6563 | Tester | Other | 2022/08/22 00:24:35 | 2022/08/22 01:10:22 | 0.76 | OK | 09108 | 1.Tester barcode 無反應 更換備品ok 還機 setup | NO | NO | NO | YES | YES | YES
EQPRepairJ22110343 | T6373-12 | CWC8J0A86-01 | OverKill | JBCK3S02 | NB01853B(COG48AA)-A-AT-C | GGAN62.1-BC1 | 09 | EDA Show I (No Corr Wafer)-A6318 | CPAUTO | Over Kill | Show I | 2022/08/22 00:29:15 | 2022/08/22 00:32:51 | 0.06 | OK | 09108 | 1.eda Show I 續測 | NO | NO | NO | YES | YES | NO
EQPRepairJ22110358 | T6373-43 | CWC6J0747-01 | Tester Down | JB1UJS03 | PA5129LS_12U80H-LT | P047F.1 | 10 | AlarmCode:669 AlarmMessage:GP-IB COMMUNICATION TIMEOUT ERROR !! | CPAUTO | Tester | 離線(OCT/GPIB error) | 2022/08/22 01:20:29 | 2022/08/22 01:33:09 | 0.21 | OK | 17091 | Remote---- 1.無異常訊息 自動起測，無異常，續測 | NO | NO | NO | YES | YES | NO
EQPRepairJ22110360 | T6373-135 | CWC8J0865-01 | Prober Down | TB1YLS38 | PA6485E(CP1-CP3) | P0CSA.1 | 06 | AlarmCode:1995 AlarmMessage:Faied to upload log file | CPAUTO | Prober | FAIL TO UPLOAD | 2022/08/22 01:40:17 | 2022/08/22 01:42:39 | 0.04 | OK | 17091 | REMOTE---清除訊息oK | NO | NO | NO | YES | YES | NO
EQPRepairJ22110362 | T6373-47 | CWC8J0981-01.001 | OverKill | JB1HKS10 | NA01687G(COG48BA)-A-AT-N2-C | GGS939.1-B11C2 | 02 | Already follow Overkilll handling procedures but still not effective too many O/S. | A6810 | Over Kill | O/S誤宰 | 2022/08/22 01:43:31 | 2022/08/22 02:02:09 | 0.31 | OK | 16157 | Remote處理------O/S 機台清針，二分法調針Range:8um pass | NO | NO | NO | YES | YES | NO
EQPRepairJ22110364 | T6373-37 | CWC8J0119-01 | Corr Fail | JB1GVS21 | NC01682A(COG48AA)-A-AT-N3-C | GGP620.1-2B | 21 | Corr Fail | A7048 | Corr | Corr Fail | 2022/08/22 01:49:39 | 2022/08/22 01:58:46 | 0.15 | OK | 20277 | Corr Fail----比對PASS | NO | NO | NO | YES | YES | NO
EQPRepairJ22110369 | T6373-151 | CWC8J0565-01.002 | OverKill | JB252S02 | TD7850-A0S-AZ-CP | SC2QK-B2C | 13 | Already follow Overkilll handling procedures but still not effective too many O/S. | A6810 | Over Kill | O/S誤宰 | 2022/08/22 02:08:37 | 2022/08/22 02:24:36 | 0.27 | OK | 20277 | Remote---- 1.針位貼上 Range:8 機台清針 起測一段PASS穩定，還機生產 | NO | NO | NO | YES | YES | NO
EQPRepairJ22110376 | T6373-47 |  | Prober Down | JB1HKS10 | NA01687G(COG48BA)-A-AT-N2-C | GGS939.1-B11C2 |  | AlarmCode:1057 AlarmMessage:MODEL NOT FOUND ( ROUGH ALIGNMENT ) !! | CPAUTO | Prober | Wafer Edge/Mode Not Found | 2022/08/22 02:38:30 | 2022/08/22 02:42:12 | 0.06 | OK | 20182 | 機台重抓OK | NO | NO | NO | YES | YES | NO
EQPRepairJ22110378 | T6373-33 | CWC7J0K83-01.001 | OverKill | JB1TAS23 | PA5208KF-LT | J67HF | 24 | MANY O/S (5.4)OD=10 | A6118 | Over Kill | O/S誤宰 | 2022/08/22 03:06:09 | 2022/08/22 03:23:09 | 0.28 | OK | 10897 | 機台清針，針位確認，測試一段ok(46.6) | NO | NO | NO | YES | YES | NO
EQPRepairJ22110379 | T6373-111 | CWC8J0618-01.004 | OverKill | JB1Q4S14 | PA6542C | J6C0W | 10 | many o/s(157,4)od=15 | A6118 | Over Kill | O/S誤宰 | 2022/08/22 03:26:41 | 2022/08/22 03:53:56 | 0.45 | OK | 09108 | 1.o/s 清針 調針pass range:12um (131.0) | NO | NO | NO | YES | YES | NO
EQPRepairJ22110381 | T6373-128 | CWC7J0527-01 | OverKill | TB1R8S74 | PA5137F-LT1 | P07UM.1 | 16 | # 16 too many o/s | A6441 | Over Kill | O/S誤宰 | 2022/08/22 03:32:48 | 2022/08/22 03:51:19 | 0.31 | OK | 20277 | Remote---- 1.針位貼下 Range:8 PASS | NO | NO | NO | YES | YES | NO
EQPRepairJ22110387 | T6373-43 | CWC6J0747-01 | OverKill | JB1UJS03 | PA5129LS_12U80H-LT | P047F.1 | 12 | Continuous 2 wafers EDA shows "R". O/S | A6810 | Over Kill | O/S誤宰 | 2022/08/22 04:12:00 | 2022/08/22 04:30:34 | 0.31 | OK | 20182 | 2R O/S 清調針PASS PM(116.5) | NO | NO | NO | NO | YES | NO
EQPRepairJ22110388 | T6373-47 | CWC8J0981-01.001 | Other | JB1HKS10 | NA01687G(COG48BA)-A-AT-N2-C | GGS939.1-B11C2 | 04 | Already follow Overkilll handling procedures but still not effective too many O/S. | A6810 | Over Kill | O/S誤宰 | 2022/08/22 04:30:26 | 2022/08/22 04:44:24 | 0.23 | OK | 20182 | 清針調針PASS PM(35.4) | NO | NO | NO | YES | YES | NO
EQPRepairJ22110396 | T6373-43 | CWC6J0747-01 | Tester Down | JB1UJS03 | PA5129LS_12U80H-LT | P047F.1 | 12 | AlarmCode:669 AlarmMessage:GP-IB COMMUNICATION TIMEOUT ERROR !! | CPAUTO | Tester | 離線(OCT/GPIB error) | 2022/08/22 05:21:23 | 2022/08/22 05:30:56 | 0.16 | OK | 10110 | 重新連線OK | NO | NO | NO | YES | YES | NO
EQPRepairJ22110397 | T6373-90 | CWC7J0M27-01.001 | OverKill | JB1BYS12 | PA6587B_15U90H | J696G | 16 | TOO MANY O/S-A6898 | CPAUTO | Probe Card | p/c down | 2022/08/22 05:26:34 | 2022/08/22 08:09:13 | 2.71 | OK | 13222 | p/c down----測試中o/s過多 下卡磨針 上機針壓需較重針壓 >25um，且針位無明顯上下貼，但無range可調整，查看配件針尖黑點過多，配件送修，無備卡配件當 | NO | NO | NO | YES | YES | NO
EQPRepairJ22110399 | T6373-141 | CWC8J0974-01.001 | OverKill | JB1HKS09 | NA01687G(COG48BA)-A-AT-N2-C | GGS938.1-BC2 | 16 | TOO MANY OS/ALWAYS ALARM OS | A6142 | Over Kill | O/S誤宰 | 2022/08/22 05:36:34 | 2022/08/22 06:05:24 | 0.48 | OK | 20277 | Remote---- 1.Range:X>20，Y:16 機台清針 PASS | NO | NO | NO | YES | YES | NO
EQPRepairJ22110400 | T6373-43 | CWC6J0747-01 | Tester Down | JB1UJS03 | PA5129LS_12U80H-LT | P047F.1 | 12 | Tester Down | 03491 | Tester | 離線(OCT/GPIB error) | 2022/08/22 05:42:35 | 2022/08/22 05:56:19 | 0.23 | OK | 10110 | check ok | NO | NO | NO | YES | YES | NO
EQPRepairJ22110401 | T6373-33 | CWC7J0K83-01.001 | OverKill | JB1TAS23 | PA5208KF-LT | J67HF | 25 | many o/s (18,5)od=10 | A6118 | Over Kill | O/S | 2022/08/22 05:50:28 | 2022/08/22 08:06:43 | 2.27 | OK | 20230 | 1.夜班交接,原配件S23 OPEN SHORT, 2.更換再庫卡S18,上機驗證BIN 13,清調針無效 3.更換再庫卡S19,上機驗證PASS,YRG=18,XRG=40,PM(45.6) 4.故S18配件送修 | NO | NO | NO | YES | YES | NO
EQPRepairJ22110403 | T6373-70 | CWC8J0706-01.003 | OverKill | JB1Q4S19 | PA6542C | J6C10 | 15 | too many os-A6512 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/22 05:56:41 | 2022/08/22 06:09:00 | 0.21 | OK | 20182 | 1.清針PASS 起測一段穩定 PM(41.5) | NO | NO | NO | YES | YES | NO
EQPRepairJ22110405 | T6373-47 | CWC8J0981-01.001 | OverKill | JB1HKS10 | NA01687G(COG48BA)-A-AT-N2-C | GGS939.1-B11C2 | 04 | STILL O/S. PM NOT GOOD.-A7048 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/22 06:22:13 | 2022/08/22 08:30:13 | 2.13 | OK | 20075 | 下卡磨針,RNAGE 10M,MARK 182,2 | NO | NO | NO | YES | YES | NO
EQPRepairJ22110425 | T6373-90 |  | Probe Card Down | JB1BYS12 | PA6587B_15U90H | J696G |  | p/c down----測試中o/s過多 下卡磨針 上機針壓需較重針壓 >25um，且針位無明顯上下貼，但無range可調整，查看配件針尖黑點過多，配件送修，無備卡配件當 | 13222 | Probe Card | p/c down | 2022/08/22 08:09:46 | 2022/08/22 14:11:48 | 6.03 | OK | 13222 | p/c down----測試中o/s過多 下卡磨針 上機針壓需較重針壓 >25um，且針位無明顯上下貼，但無range可調整，查看配件針尖黑點過多，配件送修，無備卡配件當，送修回來上機PASS (173.5) RANGE=18 | NO | NO | NO | YES | YES | NO
EQPRepairJ22110441 | T6373-07 | CWC8J0866-01 | OverKill | JB1TAS16 | PA5208KF-LT | J6A1F | 06 | TOO MANY O/S-A6564 | CPAUTO | Over Kill | o/s | 2022/08/22 09:13:08 | 2022/08/22 09:40:17 | 0.45 | OK | 13222 | o/s 機台清針ok | NO | NO | NO | YES | YES | NO
EQPRepairJ22110456 | T6373-01 | CWC7J0L59-01 | OverKill | JB1TAS20 | PA5208K-LT | QT7CH | 13 | EDA SHOW 2 H ALREADY CORR PASSED-A6564 | CPAUTO | Over Kill | 2H | 2022/08/22 10:33:38 | 2022/08/22 11:07:00 | 0.56 | OK | 13222 | 2H BIN13過多 更換配件#14 上機PASS(47.6) | NO | NO | NO | YES | YES | NO
EQPRepairJ22110467 | T6373-121 | CWC8J0A25-01.002 | OverKill | TB1INS25 | NA01404H-A-AT-C | GGV030.1-BC12 | 10 | too many o/s-A6293 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/22 11:42:17 | 2022/08/22 12:15:21 | 0.55 | OK | 14291 | 清/調針pass P/M(24.5)RG:8 | NO | NO | NO | YES | YES | NO
EQPRepairJ22110475 | T6373-47 | CWC8J0981-01.001 | OverKill | JB1HKS10 | NA01687G(COG48BA)-A-AT-N2-C | GGS939.1-B11C2 | 10 | #10 eda continues show 2r os-A6132 | CPAUTO | Over Kill | Show R | 2022/08/22 12:03:50 | 2022/08/22 12:22:19 | 0.31 | OK | 10210 | Clean wafer清針 | NO | NO | NO | YES | YES | NO
EQPRepairJ22110477 | T6373-29 | CWC8J0291-01 | OverKill | JB1QNS04 | NA01687F(COG48BA)-A-AT-C | GGU014.1-BC11 | 24 | TOO MANY O/S-A6920 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/22 12:24:20 | 2022/08/22 12:34:32 | 0.17 | OK | 13222 | REMOTE--- 1. wafer左邊須較重針壓，od:8>15 還雞生產(145.0) | NO | NO | NO | YES | YES | NO
EQPRepairJ22110480 | T6373-07 | CWC8J0866-01 | OverKill | JB1TAS16 | PA5208KF-LT | J6A1F | 07 | EDA SHOW 2 CONTINUES R O/S-A6564 | CPAUTO | Over Kill | 2r | 2022/08/22 12:54:59 | 2022/08/22 13:49:06 | 0.90 | OK | 13222 | 2r 更換配件#21 bin3 short 送修 更換配件#23 上機pass (17.5) | NO | NO | NO | YES | YES | NO
EQPRepairJ22110481 | T6373-21 |  | Prober Down | JB1GKS08 | PA6543B | J6AC8 |  | AlarmCode:8630 AlarmMessage:FOCUS ERROR !! (INSTABILITY) | CPAUTO | Other | Other | 2022/08/22 12:59:05 | 2022/08/22 13:12:21 | 0.22 | OK | 15362 | 點start重抓ok | NO | NO | NO | YES | YES | NO
EQPRepairJ22110489 | T6373-29 | CWC8J0391-01 | Tester Down | JB1QNS04 | NA01687G(COG48BA)-A-AT-N2-C | GGS419.1-B11C1111 | 02 | GP-IB (MES_DATA NOT PROD)-A6920 | CPAUTO | Other | Other | 2022/08/22 13:52:58 | 2022/08/22 13:59:16 | 0.11 | OK | 15362 | #02 退片,請產線重上 | NO | NO | NO | YES | YES | NO
EQPRepairJ22110490 | T6373-129 | CWB9J1640-01.005 | Prober Down | JB1GVS17 | NC01682A(COG48AA)-A-AT-N3-C | GFX925.1-BC2121 | 20 | AlarmCode:4827 AlarmMessage:ARM @@VACUUM ERROR !! | CPAUTO | Prober | 氣體供應異常(Air/Vacuum Supply Error) | 2022/08/22 14:03:47 | 2022/08/22 17:41:35 | 3.63 | OK | 07764 | 1.測試中發生ARM VACUUM ERROR , 開啟測蓋查看CLEAN WAFER 造成 2.重新RESET ,先升溫100度烘烤 , 在降溫-40度 3.查看維修紀錄發生多次也修改DELAY TIME無效 , 試著更換新的CLEAN WAFER 4.重新SETUP , 手動執行CLEAN WAFER無異常 , 續測觀察 5.STDF此片#20 HOLD TO TPD | NO | NO | NO | YES | YES | NO
EQPRepairJ22110500 | T6373-111 | CWC8J0440-01 | OverKill | JB1Q4S14 | PA6542C | J6C0T | 18 | TOO MANY O/S-A6131 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/22 14:32:55 | 2022/08/22 15:11:31 | 0.64 | OK | 13581 | 1.使用機台清針+調針位PASS(128.5) | NO | NO | NO | NO | YES | NO
EQPRepairJ22110508 | T6373-07 | CWC8J0866-01 | OverKill | JB1TAS23 | PA5208KF-LT | J6A1F | 08 | too many os-A6564 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/22 15:00:42 | 2022/08/22 15:15:03 | 0.24 | OK | 15362 | 清調針pass,y range:10,p/m(17,5) | NO | NO | NO | YES | YES | NO
EQPRepairJ22110513 | T6373-47 | CWC8J0981-01.001 | OverKill | JB1HKS10 | NA01687G(COG48BA)-A-AT-N2-C | GGS939.1-B11C2 | 11 | p/m ng-11550 | CPAUTO | Prober | 確認針位 | 2022/08/22 15:25:13 | 2022/08/22 15:46:01 | 0.35 | OK | 10210 | 調整針位(146.4) | NO | NO | NO | YES | YES | NO
EQPRepairJ22110518 | T6373-07 | CWC8J0866-01 | OverKill | JB1TAS23 | PA5208KF-LT | J6A1F | 08 | TOO MANY O/S-A6564 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/22 15:43:35 | 2022/08/22 16:25:50 | 0.70 | OK | 20075 | 重新對針range 10m mark 164,3 | NO | NO | NO | YES | YES | NO
EQPRepairJ22110543 | T6373-43 | CWC6J0747-01 | OverKill | JB1UJS03 | PA5129LS_12U80H-LT | P047F.1 | 14 | too many o/s.. probemark up/down-A6487 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/22 17:17:14 | 2022/08/22 18:28:27 | 1.19 | OK | 20075 | 換卡#s18,RANGE 12,MARK -1,3 | NO | NO | NO | YES | YES | NO
EQPRepairJ22110550 | T6373-86 | CRABJ0004-01 | OverKill | JB1Q4S15 | PA6542D | J2MC4 | 25 | TOO MANY O/S-A6131 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/22 17:46:58 | 2022/08/22 18:21:09 | 0.57 | OK | 14291 | 清針pass | NO | NO | NO | YES | YES | NO
EQPRepairJ22110552 | T6373-39 | CWC7J0M32-01 | OverKill | JB1UJS13 | PA5129LS_12U80H-LT | P060Q.1 | 19 | TOO MANY O/S-A6487 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/22 17:49:37 | 2022/08/22 18:10:17 | 0.34 | OK | 13222 | REMOTE---機台清針ok | NO | NO | NO | YES | NO | NO
EQPRepairJ22110560 | T6373-01 | CWC8J0815-01.001 | OverKill | JB1TAS14 | PA5208K-LT | QT7YT | 02 | too many o/s-A6564 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/22 18:30:09 | 2022/08/22 18:44:02 | 0.23 | OK | 15362 | 清調針pass,p/m(17,5) | NO | NO | NO | YES | YES | NO
EQPRepairJ22110574 | T6373-111 | CWC7J0C21-01.005 | OverKill | JB1Q4S14 | PA6542C | J68S1 | 04 | TOO MANY OS | A6142 | Over Kill | O/S誤宰 | 2022/08/22 19:58:14 | 2022/08/22 20:21:07 | 0.38 | OK | 10897 | 換片連續O/S，清針，加針壓，針位確認，測試一段OK(127.5) | NO | NO | NO | NO | YES | NO
EQPRepairJ22110580 | T6373-47 | CWC8J0981-01.001 | OverKill | JB1HKS10 | NA01687G(COG48BA)-A-AT-N2-C | GGS939.1-B11C2 | 15 | TOO MANY OS-A6512 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/22 20:16:43 | 2022/08/22 20:47:41 | 0.52 | OK | 20277 | Remote---- 1.Range:10 機台清針 PASS | NO | NO | NO | YES | YES | NO
EQPRepairJ22110590 | T6373-135 | CWC8J0865-01 | OverKill | TB1YLS38 | PA6485E(CP1-CP3) | P0CSA.1 | 15 | Already follow Overkilll handling procedures but still not effective too many o/s | 20372 | Over Kill | 配件有異物 | 2022/08/22 21:04:36 | 2022/08/22 22:04:14 | 0.99 | OK | 20182 | 1.配件有異物，下卡清針 上機PASS 2.二分法RG-8 PM(25,5) | NO | NO | NO | YES | YES | NO
EQPRepairJ22110595 | T6373-47 | CWC8J0981-01.001 | Prober Down | JB1HKS10 | NA01687G(COG48BA)-A-AT-N2-C | GGS939.1-B11C2 | 15 | AlarmCode:1057 AlarmMessage:MODEL NOT FOUND ( ROUGH ALIGNMENT ) !! | CPAUTO | Prober | Wafer Edge/Mode Not Found | 2022/08/22 21:37:17 | 2022/08/22 21:47:10 | 0.16 | OK | 20277 | Remote---- 1.抓水平OK | NO | NO | NO | YES | YES | NO
EQPRepairJ22110606 | T6373-111 | CWC7J0C21-01.005 | OverKill | JB1Q4S14 | PA6542C | J68S1 | 04 | TOO MANY OS | A6142 | Over Kill | O/S誤宰 | 2022/08/22 22:12:01 | 2022/08/22 22:56:28 | 0.74 | OK | 10897 | 重新對針，與針位確認試打(130.5)測試一段ok | NO | NO | NO | NO | YES | NO
EQPRepairJ22110611 | T6373-38 | CWC8J0841-01.002 | OverKill | JB1GKS12 | PA6543B | J6AK2 | 23 | alraedy follow overkill but still too many o/s-A6944 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/22 22:29:52 | 2022/08/22 22:44:21 | 0.24 | OK | 17091 | remote----換片起測os 調整針位pass | NO | NO | NO | YES | YES | NO
EQPRepairJ22110619 | T6373-07 | CWC8J0866-01 | OverKill | JB1TAS23 | PA5208KF-LT | J6A1F | 10 | already fo9loow overkill but still too many o/s-A6944 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/22 23:05:29 | 2022/08/22 23:35:30 | 0.50 | OK | 10110 | 清針調針pass | NO | NO | NO | YES | YES | NO
EQPRepairJ22110632 | T6373-43 | CWC6J0747-01 | OverKill | JB1UJS18 | PA5129LS_12U80H-LT | P047F.1 | 16 | too many o/s od8 no particle-A6943 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/23 00:13:35 | 2022/08/23 00:39:14 | 0.43 | OK | 20277 | Remote---- 1.Range:14 機台清針 加針壓 PASS | NO | NO | NO | YES | YES | NO
EQPRepairJ22110637 | T6373-47 | CWC8J0981-01.001 | OverKill | JB1HKS10 | NA01687G(COG48BA)-A-AT-N2-C | GGS939.1-B11C2 | 18 | too many o/s od8 no particle-A6943 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/23 00:32:02 | 2022/08/23 01:11:30 | 0.66 | OK | 20277 | Remote---- 1.下卡磨針 上機PASS Range:12 | NO | NO | NO | YES | YES | NO
EQPRepairJ22110676 | T6373-38 | CMC8J0082-01 | Tester Down | JB1GKS12 | PA6543B | J6630 | 17 | tester shut down while running | A6944 | Other | Other | 2022/08/23 03:15:05 | 2022/08/23 03:45:59 | 0.51 | OK | 17091 | 誤掛 | NO | NO | NO | YES | YES | NO
EQPRepairJ22110679 | T6373-128 | CWC7J0H96-01 | OverKill | TB1R8S74 | PA5137F-LT1 | P094J.1 | 23 | TOO MANY BIN 9-A6516 | CPAUTO | Over Kill | Function誤宰 | 2022/08/23 03:21:39 | 2022/08/23 03:47:12 | 0.43 | OK | 20277 | Remote---- 1.針位貼下 Range:12 機台清針 起測一段PASS穩定，還機生產 | NO | NO | NO | YES | YES | NO
EQPRepairJ22110682 | T6373-147 | CWC8J0C40-01 | OverKill | TB1R7S31 | PA5223Q_12U60H-LT | P0CLN.1 | 04 | o/s-20372 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/23 03:56:40 | 2022/08/23 04:09:54 | 0.22 | OK | 20277 | Remote---- 1.針位貼下下下...Range:16 機台清針 PASS | NO | NO | NO | YES | YES | NO
EQPRepairJ22110684 | T6373-141 | CWC8J0975-01 | OverKill | JB1HKS09 | NA01687G(COG48BA)-A-AT-N2-C | GGS939.2-BC | 05 | too many o/s-A6673 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/23 04:10:12 | 2022/08/23 04:56:18 | 0.77 | OK | 10110 | 下卡磨針上機pass | NO | NO | NO | YES | YES | NO
EQPRepairJ22110688 | T6373-111 | CWC8J0770-01.001 | OverKill | JB1Q4S14 | PA6542C | J6C0Y | 07 | pass-fail pass-fail, fail- fail. too many o/s-A6938 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/23 04:42:28 | 2022/08/23 05:41:25 | 0.98 | OK | 17184 | 1.重新對針&清針(173.2) | NO | NO | NO | YES | YES | NO
EQPRepairJ22110707 | T6373-43 | CWC6J0747-01 | Tester Down | JB1UJS18 | PA5129LS_12U80H-LT | P047F.1 | 19 | AlarmCode:669 AlarmMessage:GP-IB COMMUNICATION TIMEOUT ERROR !! | CPAUTO | Tester | 離線(OCT/GPIB error) | 2022/08/23 06:33:27 | 2022/08/23 06:38:48 | 0.09 | OK | 09108 | Remote---- 1.無異常訊息 消音後自動起測，無異常，續測 | NO | NO | NO | YES | YES | NO
EQPRepairJ22110708 | T6373-47 | CWC8J0981-01.001 | OverKill | JB1HKS10 | NA01687G(COG48BA)-A-AT-N2-C | GGS939.1-B11C2 | 22 | o/s | 03491 | Over Kill | O/S誤宰 | 2022/08/23 06:43:16 | 2022/08/23 06:54:12 | 0.18 | OK | 20182 | Wafer Clean PASS | NO | NO | NO | YES | YES | NO
EQPRepairJ22110709 | T6373-39 | CWC7J0M32-01 | Prober Down | JB1UJS13 | PA5129LS_12U80H-LT | P060Q.1 | 24 | AlarmCode:1995 AlarmMessage:Faied to upload log file | CPAUTO | Prober | 網路異常(Spooling Error) | 2022/08/23 06:44:25 | 2022/08/23 06:48:40 | 0.07 | OK | 20277 | Remote---- 1.Check OK，續測 | NO | NO | NO | YES | YES | NO
EQPRepairJ22110715 | T6373-33 | CWC8J0867-01 | OverKill | JB1TAS19 | PA5208KF-LT | J6CCA | 09 | still open/short-A6858 | CPAUTO | Over Kill | 針位調整 | 2022/08/23 07:20:09 | 2022/08/23 08:13:52 | 0.90 | OK | 07764 | Remote處理---- 1.測試連續open , 查看針痕 , 針偏全面性貼左 2.重新對針抓取contact 與針痕pass #09(22.5) | NO | NO | NO | YES | YES | NO
EQPRepairJ22110720 | T6373-38 | CMC8J0082-01 | OverKill | JB1GKS12 | PA6543B | J6630 | 22 | EDA SHOW 2H ALREADY CORR PASSED-A6564 | CPAUTO | Over Kill | 2H | 2022/08/23 07:43:27 | 2022/08/23 08:02:44 | 0.32 | OK | 15362 | 2H 不同卡都有相同現像,CORR PASS,續測,待換批觀察 | NO | NO | NO | YES | YES | NO
EQPRepairJ22110739 | T6373-08 | CMC1J0048-01 | Setup Fail | VB108S02 | PA5252N | AXM342000 | 06 | 08:15 #06 ID/setup fail due to O/S. No other P/C release(CM retest lot). already remount tester. AE check & confirm MBD. use 1st wafer.3d clean by AE | A7086 | Prober | Prober down | 2022/08/23 08:55:24 | 2022/08/23 10:09:40 | 1.24 | OK | 13222 | Prober down--- 1.改機一串O/S，轉配件O/S會變多變少，查看LOG與昨天KEY交接差不多LOG 2.查看昨日有調整過docking，今日改機上機相同狀況 3.再請Prober team，協助確認docking謝謝~~ | NO | NO | NO | YES | YES | NO
EQPRepairJ22110740 | T6373-70 | CWC8J0706-01 | OverKill | JB1Q4S19 | PA6542C | J6C10 | 12 | TOO MANY O/S-A6942 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/23 09:00:13 | 2022/08/23 09:06:27 | 0.10 | OK | 10210 | 調整針位(127.5) | NO | NO | NO | YES | YES | NO
EQPRepairJ22110748 | T6373-50 | CWC4J0J76-01 | Other | JB140S06 | TC2058BG-A0A | J638L | 01 | TEST DATA NOT MATCH WITH MES-04950 | CPAUTO | Product Issue | 自動化異常 | 2022/08/23 09:52:32 | 2022/08/23 10:17:14 | 0.41 | OK | 10210 | 1.產線自動化無法進站，Show異常訊息如圖。 2.此機轉掛產品當，請產品協助確認 | NO | NO | NO | NO | YES | NO
EQPRepairJ22110753 | T6373-08 |  | Prober Down | VB108S02 | PA5252N | AXM342000 |  | Prober down--- 1.改機一串O/S，轉配件O/S會變多變少，查看LOG與昨天KEY交接差不多LOG 2.查看昨日有調整過docking，今日改機上機相同狀況 3.再請Prober team，協助確認docking謝謝~~ | 13222 | Tester | 一串O/S | 2022/08/23 10:10:54 | 2022/08/23 10:19:09 | 0.14 | OK | 13222 | TESTER down--- 1.改機一串O/S，轉配件O/S會變多變少，查看LOG與昨天KEY交接差不多LOG 2.查看昨日有調整過docking，今日改機上機相同狀況 3.查看夜班已有驗證過HIFIX 上機會PASS 4.轉掛Tester down待確認 | NO | NO | NO | YES | YES | NO
EQPRepairJ22110756 | T6373-50 |  | Product Issue | JB140S06 | TC2058BG-A0A | J638L |  | 1.產線自動化無法進站，Show異常訊息如圖。 2.此機轉掛產品當，請產品協助確認 | 10210 | Product Issue | 自動化異常 | 2022/08/23 10:17:42 | 2022/08/23 10:47:38 | 0.50 | OK | 10210 | 1.產品通知重印測試注意事項，此機改idle | NO | NO | NO | YES | YES | NO
EQPRepairJ22110758 | T6373-08 |  | Tester Down | VB108S02 | PA5252N | AXM342000 |  | TESTER down--- 1.改機一串O/S，轉配件O/S會變多變少，查看LOG與昨天KEY交接差不多LOG 2.查看昨日有調整過docking，今日改機上機相同狀況 3.查看夜班已有驗證過HIFIX 上機會PASS 4.轉掛Tester down待確認 | 13222 | Other | Other | 2022/08/23 10:19:30 | 2022/08/23 15:26:06 | 5.11 | OK | 07764 | 1.TESTER team 更換HIFIX變docking異常 , 請prober team協助調整 2.調整ok , setup pass ok 3.掛setup04 | NO | NO | NO | YES | YES | NO
EQPRepairJ22110762 | T6373-86 | CWC7J0J58-01 | Corr Fail | JB1Q4S15 | PA6542C | J6C0W | 10 | Corr Fail | 14095 | Over Kill | Other | 2022/08/23 10:39:26 | 2022/08/23 11:04:21 | 0.42 | OK | 07764 | Remote處理----corr fail open , 回收pass | NO | NO | NO | YES | YES | NO
EQPRepairJ22110764 | T6373-35 | CWC8J0A25-01.001 | OverKill | TB1INS26 | NA01404H-A-AT-C | GGV030.1-BC21 | 09 | TOO MANY FAIL BIN20-A6564 | CPAUTO | Over Kill | BIN20 LEAK | 2022/08/23 10:56:41 | 2022/08/23 11:11:50 | 0.25 | OK | 13222 | BIN20 LEAK 裝風管 待10分鐘 點測PASS 還機生產 | NO | NO | NO | YES | YES | NO
EQPRepairJ22110774 | T6373-07 | CWC8J0866-01 | OverKill | JB1TAS23 | PA5208KF-LT | J6A1F | 15 | too many o/s-A6919 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/23 11:51:32 | 2022/08/23 12:04:14 | 0.21 | OK | 15362 | 清調針pass,p/m(149,5) | NO | NO | NO | YES | YES | NO
EQPRepairJ22110775 | T6373-86 | CWC8J0706-01.006 | OverKill | JB1Q4S15 | PA6542C | J6C10 | 13 | pc has particles | A6785 | Over Kill | O/S誤宰 | 2022/08/23 11:54:52 | 2022/08/23 12:39:07 | 0.74 | OK | 14291 | 換卡s16 pass P/M(12.1)RG:10 | NO | NO | NO | YES | YES | NO
EQPRepairJ22110791 | T6373-33 | CWC8J0867-01 | OverKill | JB1TAS19 | PA5208KF-LT | J6CCA | 13 | CONTINUOUS 2 WAFER EDA SHOW R. O/S-A6487 | CPAUTO | Over Kill | 2R | 2022/08/23 13:29:25 | 2022/08/23 14:48:10 | 1.31 | OK | 14291 | 換卡s25 pass P/M(45.6)RG:10 | NO | NO | NO | YES | YES | NO
EQPRepairJ22110813 | T6373-21 | CWC8J0752-01 | Prober Down | JB1GKS08 | PA6543B | J6AK6 | 02 | AlarmCode:1058 AlarmMessage:MODEL NOT FOUND ( FINE ALIGNMENT ) !! | CPAUTO | Prober | Wafer Mode Not Found | 2022/08/23 15:10:18 | 2022/08/23 15:16:10 | 0.10 | OK | 15362 | 重新抓取,check die ok,p/m(172,5) | NO | NO | NO | YES | YES | NO
EQPRepairJ22110818 | T6373-38 | CWC8J0615-01.010 | Corr Fail | JB1GKS12 | PA6543B | J6AC7 | 07 | Corr Fail | 11316 | Corr | corr fail | 2022/08/23 15:21:43 | 2022/08/23 15:41:01 | 0.32 | OK | 13222 | corr fail -- 1. 扣掉bin127 比對 pass | NO | NO | NO | YES | YES | NO
EQPRepairJ22110825 | T6373-147 | CWC7J0L69-01.002 | Corr Fail | JB1UJS03 | PA5129LS_12U80H-LT | P078M.1 | 16 | run corr ok-A6143 | CPAUTO | Corr | Other | 2022/08/23 15:51:18 | 2022/08/23 15:57:36 | 0.11 | OK | 10210 | 手動比對Pass | NO | NO | NO | YES | YES | NO
EQPRepairJ22110826 | T6373-70 |  | Prober Down | JB1Q4S19 | PA6542C | J6C0W |  | AlarmCode:7214 AlarmMessage:Soaktime correct Z-UP limit over !! | CPAUTO | Prober | SOAK TIME ERROR | 2022/08/23 16:00:14 | 2022/08/23 16:16:18 | 0.27 | OK | 20075 | RE-RUN SOAK TIME OK | NO | NO | NO | YES | YES | NO
EQPRepairJ22110828 | T6373-38 | CWC8J0621-01 | OverKill | JB1GKS12 | PA6543B | J6AC8 | 01 | too many o/s | A6919 | Over Kill | O/D不足 | 2022/08/23 16:19:45 | 2022/08/23 16:35:09 | 0.26 | OK | 13222 | O/D不足 od:5>15 | NO | NO | NO | YES | YES | NO
EQPRepairJ22110835 | T6373-08 |  | Tester Down | VB108S02 | PA5252N | AXM342000 |  | tester NG inform PM ENG | A6858 | Product Issue | 自動化異常 | 2022/08/23 16:37:23 | 2022/08/23 17:35:24 | 0.97 | OK | 20352 | 產線上貨自動化異常,使用CBSETUP驗證OK | NO | NO | NO | YES | YES | NO
EQPRepairJ22110844 | T6373-70 |  | Prober Down | JB1Q4S19 | PA6542C | J6C0W |  | AlarmCode:7214 AlarmMessage:Soaktime correct Z-UP limit over !! | CPAUTO | Prober | SOAKTIME ERROR | 2022/08/23 17:35:27 | 2022/08/23 17:52:11 | 0.28 | OK | 10210 | 按START OK | NO | NO | NO | YES | YES | NO
EQPRepairJ22110845 | T6373-12 | CWC8J0C42-01 | OverKill | JB6K8S03 | NC01682B(COG48AA)-A-AT-N-C | GGU463.1-B1C | 12 | #11 EDA CONTINUOUS SHOW I.ALREADY CORR PASSED-A6487 | CPAUTO | Over Kill | Other | 2022/08/23 17:35:45 | 2022/08/23 17:56:46 | 0.35 | OK | 07764 | Remote處理----eda i , 沒有固定dut fail 確認針痕貼下 , 調針位續測 #12(135.7) | NO | NO | NO | YES | YES | NO
EQPRepairJ22110857 | T6373-07 | CWC8J0866-01 | OverKill | JB1TAS23 | PA5208KF-LT | J6A1F | 17 | too many o/s-A6919 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/23 18:17:03 | 2022/08/23 19:01:49 | 0.75 | OK | 15362 | 更換配件#28,y range:12,p/m(18,5) | NO | NO | NO | YES | YES | NO
EQPRepairJ22110858 | T6373-08 |  | Corr Fail | VB108S02 | PA5252N | AXM342000 |  | Corr Fail-A6858 | CPAUTO | Corr | corr fail | 2022/08/23 18:21:02 | 2022/08/23 19:51:03 | 1.50 | OK | 10897 | bin36過多點測會PASS，查看貨批非corr wafer，詢問夜班產線回覆回測批協助idle | NO | NO | NO | NO | YES | NO
EQPRepairJ22110867 | T6373-135 |  | Prober Down | TB1YLS38 | PA6485E(CP1-CP3) | P0CSA.1 |  | AlarmCode:1995 AlarmMessage:Faied to upload log file | CPAUTO | Prober | Other | 2022/08/23 19:11:54 | 2022/08/23 19:22:52 | 0.18 | OK | 20182 | 查看時已起測 | NO | NO | NO | YES | YES | NO
EQPRepairJ22110881 | T6373-16 | CWC8J0C58-01 | OverKill | VB405S09 | PA5208KF-LT | J6FGH | 02 | TOO MANY O/S | A6673 | Over Kill | O/S誤宰 | 2022/08/23 20:14:59 | 2022/08/23 20:34:30 | 0.33 | OK | 20277 | Remote---- 1.#02 O/S 清、調針無效 交叉點測Pass-->Pass，Fail-->Fail穩定，請產線Run Corr驗證，續測觀察 | NO | NO | NO | YES | YES | NO
EQPRepairJ22110882 | T6373-86 | CWC8J0706-01.006 | OverKill | JB1Q4S16 | PA6542C | J6C10 | 17 | too many o/s | A6898 | Over Kill | O/S誤宰 | 2022/08/23 20:30:40 | 2022/08/23 20:45:42 | 0.25 | OK | 10897 | Remote---清針，針位確認，測試一段ok | NO | NO | NO | YES | YES | NO
EQPRepairJ22110890 | T6373-43 | CWC6J0747-01 | OverKill | JB1UJS18 | PA5129LS_12U80H-LT | P047F.1 | 25 | TOO MANY O/S..ALREADY FOLLOW OVERKILLPROCEDURE | A6673 | Over Kill | O/S誤宰 | 2022/08/23 20:58:12 | 2022/08/23 21:12:56 | 0.25 | OK | 20182 | 1.清針PASS PM(12.4) | NO | NO | NO | NO | YES | NO
EQPRepairJ22110891 | T6373-151 | CWC8J0565-01.002 | OverKill | JB252S02 | TD7850-A0S-AZ-CP | SC2QK-B2C | 20 | too many o/s | A6673 | Over Kill | O/S誤宰 | 2022/08/23 21:00:55 | 2022/08/23 21:24:05 | 0.39 | OK | 09108 | 1.O/S 清針 調針PASS range:8um (148.7) | NO | NO | YES | YES | YES | NO
EQPRepairJ22110895 | T6373-129 | CWB9J1640-01.005 | OverKill | JB1GVS17 | NC01682A(COG48AA)-A-AT-N3-C | GFX925.1-BC2121 | 20 | TOO MANY O/S | A6898 | Over Kill | O/S誤宰 | 2022/08/23 21:07:14 | 2022/08/23 21:30:17 | 0.38 | OK | 10897 | 清針，調針，點測五點水平PASS，試打(44.6) | NO | NO | NO | NO | YES | NO
EQPRepairJ22110896 | T6373-42 | CWC8J0855-01 | OverKill | JB8Q6S04 | NB01761A(COF94AA)-A-C | H2NF20.1-B1C1 | 22 | o/s-03491 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/23 21:16:50 | 2022/08/23 21:38:36 | 0.36 | OK | 20182 | 1.清針PASS 2.二分法RGX-10 Y-10 PM(52,8) | NO | NO | NO | YES | YES | NO
EQPRepairJ22110898 | T6373-32 | CWC8J0855-01.002 | OverKill | JB1PNS06 | NB01761A(COF94AA)-A-C | H2NF20.1-B1C22 | 02 | too many o/s-A6810 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/23 21:25:13 | 2022/08/23 21:36:51 | 0.19 | OK | 09108 | 1.o/s 加針壓 5->8 pass range:20um (171.8) | NO | NO | NO | YES | YES | NO
EQPRepairJ22110900 | T6373-147 | CWC8J0301-01 | OverKill | JB1UJS03 | PA5129LS_12U80H-LT | P07CN.1 | 06 | Already follow Overkilll handling procedures but still not effective too many o/s | 20372 | Over Kill | O/S誤宰 | 2022/08/23 21:37:26 | 2022/08/23 22:47:24 | 1.17 | OK | 20277 | Remote---- 1.下卡磨針 上機PASS Range:16 | NO | NO | NO | YES | YES | NO
EQPRepairJ22110907 | T6373-12 |  | Prober Down | JB6K8S03 | NC01682B(COG48AA)-A-AT-N-C | GGU463.1-B1C |  | AlarmCode:1057 AlarmMessage:MODEL NOT FOUND ( ROUGH ALIGNMENT ) !! | CPAUTO | Prober | Wafer Edge/Mode Not Found | 2022/08/23 22:06:00 | 2022/08/23 22:27:32 | 0.36 | OK | 10897 | Mode Not Found/check die/2分法Range:10um(137.6) | NO | NO | NO | YES | YES | NO
EQPRepairJ22110910 | T6373-32 | CWC8J0855-01.002 | OverKill | JB1PNS06 | NB01761A(COF94AA)-A-C | H2NF20.1-B1C22 | 02 | too many o/s od8 no particle-A6943 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/23 22:22:47 | 2022/08/23 22:44:26 | 0.36 | OK | 09108 | 1.o/s 重新對針 setup pass range: x>20 y:18um (197.7) | NO | NO | NO | YES | YES | NO
EQPRepairJ22110916 | T6373-129 | CWBBJ0K38-01.004 | Prober Down | JB1GVS17 | NC01682A(COG48AA)-A-AT-N3-C | GFFV66.1-BC22112111 | 04 | cannot measure index size-A6294 | CPAUTO | Prober | Wafer Edge/Mode Not Found | 2022/08/23 22:52:46 | 2022/08/23 23:00:30 | 0.13 | OK | 20277 | Prober Down----重抓Model OK(model not found) | NO | NO | NO | YES | YES | NO
EQPRepairJ22110920 | T6373-129 | CWBBJ0K38-01.004 | Prober Down | JB1GVS17 | NC01682A(COG48AA)-A-AT-N3-C | GFFV66.1-BC2221 | 04 | gpib-A6936 | CPAUTO | Tester | 離線(OCT/GPIB error) | 2022/08/23 23:07:40 | 2022/08/23 23:26:46 | 0.32 | OK | 10897 | T44 ND3平台，狀態未切被起測，造成離線，退片請產線重新上貨 | NO | NO | NO | NO | YES | NO
EQPRepairJ22110929 | T6373-01 | CWC8J0908-01 | OverKill | JB1TAS14 | PA5208KF-LT | J6G42 | 03 | too many o/s..already follow overkill procedure-A6673 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/23 23:43:32 | 2022/08/23 23:51:59 | 0.14 | OK | 10110 | Remote---清針調針pass | NO | NO | NO | YES | YES | NO
EQPRepairJ22110934 | T6373-151 | CWC8J0565-01.002 | OverKill | JB252S02 | TD7850-A0S-AZ-CP | SC2QK-B2C | 21 | # Already follow Overkilll handling procedures but still not effective. Too Many O/S od15 no particle | A6943 | Over Kill | O/S誤宰 | 2022/08/24 00:06:43 | 2022/08/24 00:48:25 | 0.69 | OK | 20182 | 1.清針，重對針PASS 2.RG-10 PM(53.7) | NO | NO | YES | YES | YES | NO
EQPRepairJ22110950 | T6373-42 | CWC8J0855-01 | OverKill | JB8Q6S04 | NB01761A(COF94AA)-A-C | H2NF20.1-B1C1 | 24 | many o/s (197,6)od-10-A6118 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/24 00:53:36 | 2022/08/24 01:36:11 | 0.71 | OK | 20277 | Remote---- 1.機台清針 Range:X>20，Y:16 試打(218.5) 確認針位、針壓無異常，還機生產 | NO | NO | NO | YES | YES | NO
EQPRepairJ22110960 | T6373-70 |  | Prober Down | JB1Q4S19 | PA6542C | J6C0W |  | AlarmCode:7214 AlarmMessage:Soaktime correct Z-UP limit over !! | CPAUTO | Prober | Other | 2022/08/24 01:35:32 | 2022/08/24 01:39:03 | 0.06 | OK | 20277 | Remote---- 1.Start again ok | NO | NO | NO | YES | YES | NO
EQPRepairJ22110962 | T6373-43 | CEC7J0106-01 | OverKill | JB1UJS18 | PA5129LS_12U80H-LT | P08WT.1 | 02 | o/s-03491 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/24 01:40:52 | 2022/08/24 02:07:07 | 0.44 | OK | 20277 | Remote---- 1.加針壓 Range:16 機台清針 起測一段PASS穩定，還機生產 | NO | NO | YES | YES | YES | NO
EQPRepairJ22110965 | T6373-42 | CWC8J0855-01 | OverKill | JB8Q6S04 | NB01761A(COF94AA)-A-C | H2NF20.1-B1C1 | 24 | already follow overkill procedure but still not effecitve. o/s.-A7048 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/24 02:09:36 | 2022/08/24 02:23:57 | 0.24 | OK | 20277 | Remote---- 1.Soaktime Range:X>20，Y:16 PASS | NO | NO | NO | YES | YES | NO
EQPRepairJ22110971 | T6373-151 | CWC8J0565-01.002 | OverKill | JB252S02 | TD7850-A0S-AZ-CP | SC2QK-B2C | 22 | many o/s (146.7)od=15 | A6118 | Over Kill | O/S誤宰 | 2022/08/24 02:46:13 | 2022/08/24 03:15:45 | 0.49 | OK | 20277 | Remote---- 1.Range:14 機台清針 起測一段PASS穩定，還機生產 | NO | NO | NO | YES | YES | NO
EQPRepairJ22110990 | T6373-38 | CWC8J0615-01.010 | Corr Fail | JB1GKS12 | PA6543B | J6AC8 | 07 | Corr Fail | A6810 | Corr | corr fail | 2022/08/24 04:09:19 | 2022/08/24 04:12:34 | 0.05 | OK | 09108 | 1.corr fail 扣除bin127 比對pass | NO | NO | NO | YES | YES | NO
EQPRepairJ22110993 | T6373-38 |  | Prober Down | JB1GKS12 | PA6543B | J6AC8 |  | it failed to lock/unlock manual ring lever-A6943 | CPAUTO | Prober | Other | 2022/08/24 04:35:40 | 2022/08/24 04:41:39 | 0.10 | OK | 20182 | 調整ok 還機生產 | NO | NO | NO | NO | YES | NO
EQPRepairJ22110996 | T6373-136 | CWC8J0870-01.001 | OverKill | JB1PNS08 | NB01761A(COF94AA)-A-C | H2NF33.1-BC2 | 02 | many o/s (26,1)od-10-A6118 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/24 04:50:52 | 2022/08/24 05:03:58 | 0.22 | OK | 20182 | 1.機台清針PASS RGX-10 Y-10 PM(170.8) | NO | NO | NO | YES | YES | NO
EQPRepairJ22110997 | T6373-147 |  | Prober Down | TB1R7S31 | PA5223Q_12U60H-LT | P0CLN.1 |  | AlarmCode:1060 AlarmMessage:CANNOT MEASURE INDEX SIZE !! | CPAUTO | Prober | Wafer Edge/Mode Not Found | 2022/08/24 05:00:13 | 2022/08/24 05:08:24 | 0.14 | OK | 20182 | 機台重抓MoDe Ok Cheack Die OK | NO | NO | NO | YES | YES | NO
EQPRepairJ22111013 | T6373-08 |  | Tester Down | VB108S02 | PA5252N | AXM342000 |  | message appear on tester | A6944 | Tester | Other | 2022/08/24 06:00:19 | 2022/08/24 06:20:52 | 0.34 | OK | 10897 | 產線上貨自動化異常，使用CBSETUP OK | NO | NO | NO | NO | YES | NO
EQPRepairJ22111016 | T6373-121 | CWC8J0974-01.003 | OverKill | JB1QNS04 | NA01687G(COG48BA)-A-AT-N2-C | GGS938.1-BC22 | 05 | #03.04 CONTINOUS EDA SHOW H, NO CORR | A6898 | Over Kill | show H | 2022/08/24 06:25:57 | 2022/08/24 08:26:21 | 2.01 | OK | 11385 | EDA H 夜班交接配件下卡磨針，磨針上機續測 | NO | NO | NO | YES | YES | NO
EQPRepairJ22111031 | T6373-70 |  | Prober Down | JB1Q4S19 | PA6542C | J6C0Y |  | AlarmCode:7214 AlarmMessage:Soaktime correct Z-UP limit over !! | CPAUTO | Prober | soaktime Z-UP limit | 2022/08/24 07:08:08 | 2022/08/24 07:58:00 | 0.83 | OK | 19771 | 1.soaktime Z-UP limit 重新對針 固定pin 不過 2.執行soaktime後pass 點測不穩 3.機台清針後OK 4.#3(44.5) | NO | NO | NO | YES | YES | NO
EQPRepairJ22111037 | T6373-29 | CWC8J0969-01.002 | OverKill | KB2BNS09 | ND01733D(COF94AA)-A-C | H2PU18.3-22BC12 | 24 | #21.23 2H-20022 | CPAUTO | Over Kill | SHOW H | 2022/08/24 07:37:07 | 2022/08/24 07:47:58 | 0.18 | OK | 16631 | 1.#21 EDA Show [HW](CP3 $GOODBIN > CP2 $GOODBIN EA) 108 [HW](CP3 $FAILBIN(90) <> CP2 $FAILBIN EA) -117 #23 EDA Show [HW](CP3 $GOODBIN > CP2 $GOODBIN EA) 123 [HW](CP3 $FAILBIN(90) <> CP2 $FAILBIN EA) -127 #25 EDA Show [HW](CP3 $GOODBIN > CP2 $GOODBIN EA) 110 [HW](CP3 $FAILBIN(90) <> CP2 $FAILBIN EA) -113 2.轉掛產品問題，請DN&產品協助確認，謝謝。 | NO | NO | NO | YES | YES | NO
EQPRepairJ22111040 | T6373-29 |  | Product Issue | KB2BNS09 | ND01733D(COF94AA)-A-C | H2PU18.3-22BC12 |  | Product Issue---- 1.#21 EDA Show [HW](CP3 $GOODBIN > CP2 $GOODBIN EA) 108 [HW](CP3 $FAILBIN(90) <> CP2 $FAILBIN EA) -117 #23 EDA Show 2.轉掛產品問題，請DN&產品協助確認，謝謝。 | 16631 | Product Issue | Other | 2022/08/24 07:49:32 | 2022/08/24 09:11:39 | 1.37 | OK | 16631 | 1.產品MAIL回覆:客戶改Map需求，請忽略異常訊息，續測即可 | NO | NO | NO | YES | YES | NO
EQPRepairJ22111041 | T6373-147 |  | Prober Down | TB1R7S31 | PA5223Q_12U60H-LT | P06UN.1 |  | AlarmCode:1060 AlarmMessage:CANNOT MEASURE INDEX SIZE !! | CPAUTO | Prober | CANNOT MEASURE INDEX SIZE | 2022/08/24 07:59:34 | 2022/08/24 08:14:06 | 0.24 | OK | 07910 | 重抓Modle & check die ok，P/M(18.1)#3 | NO | NO | NO | YES | YES | NO
EQPRepairJ22111043 | T6373-43 | CEC7J0106-01 | OverKill | JB1UJS18 | PA5129LS_12U80H-LT | P08WT.1 | 05 | too many o/s | A6986 | Over Kill | 機台清針後Pass | 2022/08/24 08:08:11 | 2022/08/24 08:52:42 | 0.74 | OK | 16631 | 1.機台清針後Pass | NO | NO | NO | YES | YES | NO
EQPRepairJ22111047 | T6373-15 | CWC8J0123-01.001 | Corr Fail | JB1HKS03 | NA01687G(COG48BA)-A-AT-X1-C | GGS268.1-B1 | 09 | ok-A6668 | CPAUTO | Corr | Corr Fail | 2022/08/24 08:18:55 | 2022/08/24 08:26:54 | 0.13 | OK | 16631 | 1.Corr Fail，pass | NO | NO | NO | YES | YES | NO
EQPRepairJ22111053 | T6373-70 | CWC8J0770-01.002 | OverKill | JB1Q4S19 | PA6542C | J6C0Y | 03 | O/S TOO MANY-A6916 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/24 08:43:57 | 2022/08/24 09:12:17 | 0.47 | OK | 08774 | Remote 處理---重新對針 機台清針 44.5 | NO | NO | NO | YES | YES | NO
EQPRepairJ22111055 | T6373-33 | CWC8J0867-01 | OverKill | JB1TAS25 | PA5208KF-LT | J6CCA | 21 | O/S TOO MANY-A6916 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/24 08:54:22 | 2022/08/24 09:20:54 | 0.44 | OK | 17680 | remote----調整針位position(-1,12),mark(45,0)，clean unit，起測穩定 | NO | NO | NO | YES | YES | NO
EQPRepairJ22111061 | T6373-147 | CWC7J0L67-01 | OverKill | TB1R7S31 | PA5223Q_12U60H-LT | P06UN.1 | 07 | EDA 2H O/S | A6921 | Over Kill | O/S誤宰 | 2022/08/24 09:33:10 | 2022/08/24 09:50:05 | 0.28 | OK | 08774 | Remote 處理---重新對針 機台清針 1.無掛當紀錄 2.重新對針 機台清針 續測觀察一片 落針點x-3.y-10 46.0 | NO | NO | NO | YES | YES | NO
EQPRepairJ22111065 | T6373-38 | CWC8J0621-01 | OverKill | JB1GKS12 | PA6543B | J6AC8 | 19 | continuous EDA show 2H corr passed already-A6671 | CPAUTO | Over Kill | 2H | 2022/08/24 09:43:48 | 2022/08/24 09:48:38 | 0.08 | OK | 07910 | 2H，查EDA 資料[HW_BIN41>1%](Bin 42 > 11EA) 13，零星FAIL，且CORR PASS，觀察下片 | NO | NO | NO | YES | YES | NO
EQPRepairJ22111067 | T6373-70 | CWC8J0770-01.002 | OverKill | JB1Q4S19 | PA6542C | J6C0Y | 04 | O/S-A6571 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/24 09:51:06 | 2022/08/24 10:05:32 | 0.24 | OK | 08774 | 換片o/s 溫針 清針 加針壓 77.0 | NO | NO | NO | YES | YES | NO
EQPRepairJ22111069 | T6373-136 | CWC8J0870-01.001 | OverKill | JB1PNS08 | NB01761A(COF94AA)-A-C | H2NF33.1-BC2 | 04 | too many o/s, please help check have selected probemark shift (171,8)-A6986 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/24 10:02:13 | 2022/08/24 10:14:25 | 0.20 | OK | 08774 | Remote 處理---重新對針 機台清針 | NO | NO | NO | YES | YES | NO
EQPRepairJ22111075 | T6373-42 |  | Prober Down | JB8Q6S04 | NB01761A(COF94AA)-A-C | H2KX64.1-B111C |  | AlarmCode:1057 AlarmMessage:MODEL NOT FOUND ( ROUGH ALIGNMENT ) !! | CPAUTO | Prober | Wafer Edge/Mode Not Found | 2022/08/24 10:28:08 | 2022/08/24 10:35:50 | 0.13 | OK | 08774 | Remote 處理---重抓 | NO | NO | NO | YES | YES | NO
EQPRepairJ22111087 | T6373-43 | CEC7J0106-01 | Tester Down | JB1UJS18 | PA5129LS_12U80H-LT | P08WT.1 | 06 | AlarmCode:669 AlarmMessage:GP-IB COMMUNICATION TIMEOUT ERROR !! | CPAUTO | Tester | 離線(OCT/GPIB error) | 2022/08/24 11:23:36 | 2022/08/24 11:32:38 | 0.15 | OK | 07910 | REMOTE---重新連線 OK | NO | NO | YES | YES | YES | NO
EQPRepairJ22111101 | T6373-151 | CWC8J0572-01.003 | OverKill | JB252S02 | TD7850-A0S-AZ-CP | SC1GR-B1111C | 01 | o/s-20022 | CPAUTO | Over Kill | 針位調整 | 2022/08/24 12:49:14 | 2022/08/24 13:10:33 | 0.36 | OK | 16631 | 1.針位貼下，針位調整pass | NO | NO | YES | YES | YES | NO
EQPRepairJ22111102 | T6373-70 | CWC8J0770-01.002 | OverKill | JB1Q4S19 | PA6542C | J6C0Y | 05 | O/S TOO MANY-A6916 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/24 12:52:36 | 2022/08/24 14:20:49 | 1.47 | OK | 08774 | 1.測試中O/S，黑點過多，下卡磨針 2.上機依舊有黑點，O/S不過，配件送修 3.換卡pass range14 x0.y5 41.0 | NO | NO | NO | YES | YES | NO
EQPRepairJ22111103 | T6373-15 | CWC8J0291-01.009 | OverKill | JB1HKS03 | NA01687F(COG48BA)-A-AT-C | GGU014.1-BC121121 | 08 | o/s-A6921 | CPAUTO | Over Kill | 機台清針後Pass | 2022/08/24 13:01:50 | 2022/08/24 13:11:15 | 0.16 | OK | 19771 | 機台清針後pass | NO | NO | NO | YES | YES | NO
EQPRepairJ22111106 | T6373-43 | CEC7J0106-01 | Tester Down | JB1UJS18 | PA5129LS_12U80H-LT | P08WT.1 | 07 | AlarmCode:669 AlarmMessage:GP-IB COMMUNICATION TIMEOUT ERROR !! | CPAUTO | Product Issue | RTE | 2022/08/24 13:23:47 | 2022/08/24 13:29:33 | 0.10 | OK | 08774 | Remote 處理---無異常訊息，初判可能CAL ONLINE INIT | NO | NO | NO | YES | YES | NO
EQPRepairJ22111108 | T6373-33 | CWC8J0867-01 | OverKill | JB1TAS25 | PA5208KF-LT | J6CCA | 23 | so many o/s-08341 | CPAUTO | Over Kill | O/S | 2022/08/24 13:37:15 | 2022/08/24 14:19:52 | 0.71 | OK | 11385 | 更換JB1TAS20 | NO | NO | NO | YES | YES | NO
EQPRepairJ22111123 | T6373-70 | CWC8J0770-01.002 | Probe Card Down | JB1Q4S19 | PA6542C | J6C0Y | 05 | Probe Card Not Release-A6986 | CPAUTO | Probe Card | Other | 2022/08/24 14:28:37 | 2022/08/24 14:59:27 | 0.51 | OK | 08774 | 換卡23 13.4 159.1 | NO | NO | NO | YES | YES | NO
EQPRepairJ22111124 | T6373-141 | CWC8J0973-01.001 | OverKill | JB1HKS09 | NA01687F(COG48BA)-A-AT-C | GGV502.1-BC2 | 07 | 2r o/s-a6571 | CPAUTO | Over Kill | 2R | 2022/08/24 14:37:43 | 2022/08/24 15:05:42 | 0.47 | OK | 13749 | 1.eda show 2R O/S 殘金. 2.更換備卡#S09->#S07 OK 預溫重對針 PASS PM(10.1) | NO | NO | NO | YES | YES | NO
EQPRepairJ22111127 | T6373-29 | CWC8J0969-01 | OverKill | KB2BNS09 | ND01733D(COF94AA)-A-C | H2PU18.3-22BC11 | 22 | TOO MANY O/S-A6513 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/24 15:05:20 | 2022/08/24 15:17:27 | 0.20 | OK | 16631 | 1.調整針壓/針位，機台清針 | NO | NO | NO | YES | YES | NO
EQPRepairJ22111137 | T6373-112 |  | Prober Down | JB1GVS23 | NC01682B(COG48AA)-A-AT-N-C | GGU463.1-B1C2 |  | AlarmCode:1057 AlarmMessage:MODEL NOT FOUND ( ROUGH ALIGNMENT ) !! | CPAUTO | Prober | Wafer Edge/Mode Not Found | 2022/08/24 15:39:40 | 2022/08/24 15:48:31 | 0.15 | OK | 20231 | 1.重抓MODEL ok，CHECK DIE ok，PM(168,1) | NO | NO | NO | YES | YES | NO
EQPRepairJ22111138 | T6373-136 | CWC8J0870-01.001 | OverKill | JB1PNS08 | NB01761A(COF94AA)-A-C | H2NF33.1-BC2 | 06 | TOO MANY O/S, PLEASE HELP CHECK HAVE PROBEMARKSHIFT (171,8)-A6986 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/24 15:40:39 | 2022/08/24 16:13:21 | 0.54 | OK | 15897 | 針位貼上擠壓、下調Pass.偏重5um.重對Range:22/14um.#06(25.7) | NO | NO | NO | YES | YES | NO
EQPRepairJ22111139 | T6373-151 | CWC8J0572-01.003 | OverKill | JB252S02 | TD7850-A0S-AZ-CP | SC1GR-B1111C | 04 | too mnay o/s-A6986 | CPAUTO | Over Kill | O/S | 2022/08/24 15:45:20 | 2022/08/24 17:39:22 | 1.90 | OK | 19771 | 1.測試易跳O/S，清調針無效 2.配件磨針上機pass 3.#4(147.8) | NO | NO | NO | YES | YES | NO
EQPRepairJ22111153 | T6373-43 | CEC7J0106-01 | OverKill | JB1UJS18 | PA5129LS_12U80H-LT | P08WT.1 | 08 | TOO MANY O/S | A6897 | Over Kill | 機台清針後Pass | 2022/08/24 17:02:15 | 2022/08/24 17:21:48 | 0.33 | OK | 17712 | 機台清針Pass，#8(11.4)RG10 | NO | NO | YES | YES | YES | NO
EQPRepairJ22111158 | T6373-16 | CWC8J0A77-01 | OverKill | VB405S09 | PA5208KF-LT | J6CC9 | 17 | #17 o/s bin29-20021 | CPAUTO | Over Kill | O/D不足 | 2022/08/24 17:22:10 | 2022/08/24 17:38:46 | 0.28 | OK | 17712 | 溫針重對，#17(147.5)(147.2)RG20 | NO | NO | NO | YES | YES | NO
EQPRepairJ22111160 | T6373-21 |  | Prober Down | JB1GKS08 | PA6543B | J6AC7 |  | model not found | A6571 | Prober | Wafer Edge/Mode Not Found | 2022/08/24 17:41:22 | 2022/08/24 17:53:55 | 0.21 | OK | 19771 | 手動抓取OK | NO | NO | NO | YES | YES | NO
EQPRepairJ22111181 | T6373-70 |  | Prober Down | JB1Q4S23 | PA6542C | J6C0Y |  | AlarmCode:7214 AlarmMessage:Soaktime correct Z-UP limit over !! | CPAUTO | Prober | Other | 2022/08/24 19:13:31 | 2022/08/24 19:38:24 | 0.41 | OK | 17674 | Remote 處理 ------ 1. soaktime z軸 ，機台自動抓正常 (40,0)#09 | NO | NO | NO | YES | YES | NO
EQPRepairJ22111187 | T6373-70 | CWC8J0770-01.002 | OverKill | JB1Q4S23 | PA6542C | J6C0Y | 09 | TOO MANY O/S-A6148 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/24 19:57:35 | 2022/08/24 20:24:32 | 0.45 | OK | 11348 | Remote 處理----重新對針，Range:10um | NO | NO | NO | YES | YES | NO
EQPRepairJ22111200 | T6373-39 |  | Prober Down | JB1UJS13 | PA5129LS_12U80H-LT | P07CK.1 |  | AlarmCode:1995 AlarmMessage:Faied to upload log file | CPAUTO | Prober | 網路異常(Spooling Error) | 2022/08/24 20:40:30 | 2022/08/24 20:43:13 | 0.05 | OK | 11348 | Remote 處理-----Faied to upload log file 續測 | NO | NO | NO | YES | YES | NO
EQPRepairJ22111214 | T6373-151 | CWC8J0572-01.003 | OverKill | JB252S02 | TD7850-A0S-AZ-CP | SC1GR-B1111C | 14 | eda 2 I-A7033 | CPAUTO | Over Kill | 2I | 2022/08/24 21:52:33 | 2022/08/24 22:20:07 | 0.46 | OK | 15895 | EDA SHOW 2I，機台清針，並觀察一段穩定，還機續觀察 | NO | NO | NO | YES | YES | NO
EQPRepairJ22111222 | T6373-43 | CEC7J0106-01 | OverKill | JB1UJS18 | PA5129LS_12U80H-LT | P08WT.1 | 10 | TOO MANY O/S-A6148 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/24 22:24:48 | 2022/08/24 22:32:49 | 0.13 | OK | 15895 | o/s，機台清針，p/m(35.5) | NO | NO | YES | YES | YES | NO
EQPRepairJ22111228 | T6373-151 | CWC8J0572-01.003 | OverKill | JB252S02 | TD7850-A0S-AZ-CP | SC1GR-B1111C | 14 | too many o/s-A6148 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/24 23:04:23 | 2022/08/24 23:47:24 | 0.72 | OK | 17674 | 1. all-open 重對針PASS，溫針 RANGE Y=8 (174,4)#14 | NO | NO | NO | YES | YES | NO
EQPRepairJ22111234 | T6373-70 | CWC8J0770-01.002 | OverKill | JB1Q4S23 | PA6542C | J6C0Y | 11 | TOO MANY O/S-A6148 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/24 23:16:47 | 2022/08/25 02:08:55 | 2.87 | OK | 17674 | 1. 下卡磨針上機固定一小串不過 (2185~2202-L)，重MNT PASS 後又soaktime z 軸 2. 溫針重對針pass、再溫針 range Y=10 (81,1)#11 | NO | NO | NO | YES | YES | NO
EQPRepairJ22111239 | T6373-147 | CWC8J0301-01 | OverKill | JB1UJS03 | PA5129LS_12U80H-LT | P07CN.1 | 09 | MBD-A7164 | CPAUTO | Over Kill | Function誤宰 | 2022/08/24 23:33:34 | 2022/08/25 00:22:04 | 0.81 | OK | 17674 | 1. Function誤宰 6+OS 配件殘金 (金絲)，更換備卡#14 溫針 RANGE Y=10 (99,5)#09 | NO | NO | NO | YES | YES | NO
EQPRepairJ22111246 | T6373-103 | CWC8J0981-01.001 | Setup Fail | JB1QNS05 | NA01687G(COG48BA)-A-AT-N2-C | GGS939.1-B11 | 02 | 23:30 路艾 setup bin6 fail換卡JB1QNS05 .接手SETUP 一樣fail bin6.機清無效.掛當.待輪工確認 | 07079 | Tester | Open / Short / Function Fail | 2022/08/24 23:46:19 | 2022/08/25 00:16:37 | 0.51 | OK | 13597 | 型號:NA01687G/CP3 JB1QNS05，JB1QNS06 1.改機全面 bin6 fail，換卡無效 #06>05 2.查看之前有相同問題，轉掛Tester Down | NO | NO | NO | YES | YES | NO
EQPRepairJ22111250 | T6373-152 | CWC8J0455-01.004 | Corr Fail | JB252S01 | TD7850-A0S-AZ-CP | SC1GS-B1 | 09 | Corr Fail | A5929 | Corr | Corr fail | 2022/08/25 00:11:52 | 2022/08/25 00:27:35 | 0.26 | OK | 17674 | 1. corr pass | NO | NO | NO | YES | YES | NO
EQPRepairJ22111252 | T6373-103 |  | Tester Down | JB1QNS05 | NA01687G(COG48BA)-A-AT-N2-C | GGS939.1-B11 |  | 型號:NA01687G/CP3 JB1QNS05，JB1QNS06 1.改機全面 bin6 fail，換卡無效 #06>05 2.查看之前有相同問題，轉掛Tester Down | 13597 | Tester | 無排程改切待料 | 2022/08/25 00:17:08 | 2022/08/25 02:05:44 | 1.81 | OK | 13597 | TESTER DOWN---- 型號:NA01687G/CP3 JB1QNS05，JB1QNS06 1.改機全面 bin6 fail，換卡無效 #06>05 2.查看之前有相同問題，轉掛Tester Down 3.協調PP、排程改測 4.排程拉至其它機台測試，此機無排程改切待料 | NO | NO | NO | YES | YES | NO
EQPRepairJ22111258 | T6373-34 | CWC6J0117-01 | OverKill | TB1R8S85 | PA5223Q_12U60H-LT | P047W.1 | 15 | too many o/s | A6926 | Over Kill | 針位調整 | 2022/08/25 00:39:36 | 2022/08/25 00:55:34 | 0.27 | OK | 13597 | range 10 um | NO | NO | NO | YES | YES | NO
EQPRepairJ22111287 | T6373-34 | CWC6J0052-01.002 | OverKill | TB1R8S85 | PA5223Q_12U60H-LT | P047T.1 | 18 | too many o/s..bin58-A7025 | CPAUTO | Over Kill | O/D不足 | 2022/08/25 02:43:56 | 2022/08/25 03:03:48 | 0.33 | OK | 13597 | ta----- 1.加針壓pass | NO | NO | NO | YES | YES | NO
EQPRepairJ22111294 | T6373-43 | CEC7J0106-01 | OverKill | JB1UJS18 | PA5129LS_12U80H-LT | P08WT.1 | 12 | Too many o/s-A6053 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/25 03:02:47 | 2022/08/25 03:15:05 | 0.21 | OK | 15895 | 機台清針，觀察一段穩定，還機續觀察 | NO | NO | NO | YES | YES | NO
EQPRepairJ22111297 | T6373-70 | CWC8J0770-01.002 | OverKill | JB1Q4S23 | PA6542C | J6C0Y | 11 | Too many o/s-A6053 | CPAUTO | Over Kill | O/D不足 | 2022/08/25 03:18:19 | 2022/08/25 03:54:38 | 0.61 | OK | 17542 | O/D不足，OD+2um，起測一段PASS穩定，試打(12,4)無異常 | NO | NO | NO | YES | YES | NO
EQPRepairJ22111300 | T6373-38 | CWC8J0752-01 | OverKill | JB1GKS12 | PA6543B | J6AK6 | 03 | o/s too many-19697 | CPAUTO | Over Kill | 針位調整 | 2022/08/25 03:42:58 | 2022/08/25 03:57:36 | 0.24 | OK | 13597 | TA----- 1.TA不會調針位，針位貼上，調針PASS，RANGE 8 UM | NO | NO | NO | YES | YES | NO
EQPRepairJ22111306 | T6373-33 | CWC8J0918-01 | OverKill | JB1TAS20 | PA5208KF-LT | J6CC7 | 03 | too many o/s | A6926 | Over Kill | 機台清針後Pass | 2022/08/25 04:12:51 | 2022/08/25 04:27:00 | 0.24 | OK | 17542 | Remote處理 ----- 機台清針後Pass | NO | NO | NO | YES | YES | NO
EQPRepairJ22111314 | T6373-33 | CWC8J0918-01 | OverKill | JB1TAS20 | PA5208KF-LT | J6CC7 | 03 | o/s,od:10,p/m(46.6)-A6336 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/25 04:51:35 | 2022/08/25 05:27:19 | 0.60 | OK | 11348 | 更換配件#26 pass | NO | NO | NO | YES | YES | NO
EQPRepairJ22111329 | T6373-136 | CWC8J0870-01.001 | OverKill | JB1PNS08 | NB01761A(COF94AA)-A-C | H2NF33.1-BC21 | 10 | too many o/s-A6148 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/25 05:52:59 | 2022/08/25 06:09:36 | 0.28 | OK | 15895 | 輪工處理----二分法調針，y:16，p/m(171.8) | NO | NO | NO | YES | YES | NO
EQPRepairJ22111338 | T6373-01 | CWC8J0908-01 | OverKill | JB1TAS14 | PA5208KF-LT | J6G42 | 12 | TOO MANY O/S..BIN2-A7025 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/25 06:20:01 | 2022/08/25 06:33:31 | 0.23 | OK | 20276 | 機台清針後PASS | NO | NO | NO | YES | YES | NO
EQPRepairJ22111342 | T6373-32 | CWC8J0855-01.002 | OverKill | JB1PNS06 | NB01761A(COF94AA)-A-C | H2NF20.1-B1C22 | 13 | TOO MANY O/S..BIN2-A7025 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/25 06:30:25 | 2022/08/25 06:48:46 | 0.31 | OK | 20276 | 機台清針後PASS | NO | NO | NO | YES | YES | NO
EQPRepairJ22111347 | T6373-08 |  | Tester Down | VB104S01 | PA5252N | AXM342000 |  | MACHINE STOP USING PLEASE INFORM PM ENGINEER TO CHECK-A7024 | CPAUTO | Tester | 自動化異常 | 2022/08/25 06:46:22 | 2022/08/25 06:57:26 | 0.18 | OK | 11348 | 自動化異常，先請產線cbsetup測試 | NO | NO | NO | NO | YES | NO
EQPRepairJ22111351 | T6373-151 | CWC8J0449-01.003 | OverKill | JB252S02 | TD7850-A0S-AZ-CP | SC1GT-B1121C | 14 | too many bin 20-A7033 | CPAUTO | Over Kill | Function誤宰 | 2022/08/25 06:59:32 | 2022/08/25 08:28:45 | 1.49 | OK | 15897 | 產品問題---- 1.Dut1 Bin20 Function過多.點測先前Pass同樣.磨針無效. 2.Wafer alignment重測驗證Pass.產品問題.溫針重對Range:8um.退片請產線重新上貨#14(23.6). | NO | NO | YES | YES | YES | NO
EQPRepairJ22111368 | T6373-43 | CEC7J0106-01 | OverKill | JB1UJS18 | PA5129LS_12U80H-LT | P08WT.1 | 14 | O/S-A6571 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/25 08:11:25 | 2022/08/25 08:37:19 | 0.43 | OK | 08774 | Remote 處理---重新對針 機台清針 | NO | NO | YES | YES | YES | NO
EQPRepairJ22111375 | T6373-12 | CWC8J0C42-01 | Setup Fail | JB6K8S01 | NC01682B(COG48AA)-A-AT-N-C | GGU463.1-B1 | 10 | setup DUT2 o/s 調針清針重mount無效 換卡#04→#01 DUT4 o/s 調針重mount無效 下卡磨針 | 05671 | Setup Fail | O/S Fail | 2022/08/25 08:49:02 | 2022/08/25 09:32:17 | 0.72 | OK | 17680 | remote----range:23um,position(-5,-6),mark(17,5) | NO | NO | NO | YES | YES | NO
EQPRepairJ22111386 | T6373-147 | CWC8J0301-01 | OverKill | JB1UJS14 | PA5129LS_12U80H-LT | P07CN.1 | 12 | o/s, probemark up and down-a6921 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/25 09:26:32 | 2022/08/25 09:46:39 | 0.34 | OK | 07910 | REMOTE----機台清針 溫針 2分法PASS P/M(138.4)#12 RY12 | NO | NO | NO | YES | YES | NO
EQPRepairJ22111409 | T6373-112 | CWC8J0C42-01.001 | OverKill | JB1GVS23 | NC01682B(COG48AA)-A-AT-N-C | GGU463.1-B1C2 | 22 | o/s-A6921 | CPAUTO | Over Kill | O/D | 2022/08/25 11:25:39 | 2022/08/25 11:47:15 | 0.36 | OK | 19771 | 1.OD過重，重新對針 2.#22(137.6) | NO | NO | NO | YES | YES | NO
EQPRepairJ22111414 | T6373-08 |  | Corr Fail | VB104S01 | PA5252N | AXM342000 |  | #02 setup no corr,首片比對FAIL | 19617 | Corr | 產品問題 | 2022/08/25 11:55:02 | 2022/08/25 12:12:20 | 0.29 | OK | 11385 | Corr Fail---- T32 產品無 Corr Wafer Run Corr 比對 Fail | NO | NO | NO | YES | YES | NO
EQPRepairJ22111417 | T6373-147 | CWC8J0301-01 | OverKill | JB1UJS14 | PA5129LS_12U80H-LT | P07CN.1 | 13 | o/s (12. 4 pm check)-11933 | CPAUTO | Over Kill | O/S | 2022/08/25 12:14:26 | 2022/08/25 13:40:29 | 1.43 | OK | 19771 | 1.配件磨針#13(116.5) | NO | NO | NO | YES | YES | NO
EQPRepairJ22111423 | T6373-35 | CWC8J0973-01.001 | Setup Fail | JB1QNS06 | NA01687F(COG48BA)-A-AT-C | GGV502.1-B | 01 | 改機點測 起測皆跑BIN 6 重LOAD程式一樣 無配件可更換驗證 待輪工確認 | 08341 | Over Kill | Function誤宰 | 2022/08/25 12:39:13 | 2022/08/25 13:12:55 | 0.56 | OK | 13749 | 1.產線改機NA01687F / CP3 點測 起測皆BIN6 Function穩定. 2.確認測試頭後方熱排偏離，重新調整後 PASS Run corr PASS 退片還機生產. | NO | NO | NO | YES | YES | NO
EQPRepairJ22111429 | T6373-151 | CWC8J0449-01.003 | OverKill | JB252S02 | TD7850-A0S-AZ-CP | SC1GT-B1121C | 16 | dut1 o/s (23.6 pm check)-11933 | CPAUTO | Over Kill | O/D不足 | 2022/08/25 13:07:31 | 2022/08/25 13:33:26 | 0.43 | OK | 17712 | 溫針重對，#16(147.7)(148.7)rg10 | NO | NO | YES | YES | YES | NO
EQPRepairJ22111432 | T6373-12 | CWC7J0L30-01 | Corr Fail | JB6K8S01 | NC01682B(COG48AA)-A-AT-N-C | GGG498.1-B | 25 | setup ru corr-05671 | CPAUTO | Corr | 產品問題 | 2022/08/25 13:16:31 | 2022/08/25 13:18:02 | 0.03 | OK | 11385 | 手動比對PASS | NO | NO | NO | YES | YES | NO
EQPRepairJ22111441 | T6373-43 | CEC7J0106-01 | OverKill | JB1UJS18 | PA5129LS_12U80H-LT | P08WT.1 | 16 | o/s-a6571 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/25 14:08:42 | 2022/08/25 15:23:23 | 1.24 | OK | 08774 | 啟家處理 告知配件有黑點，換卡s/u 35.5 34.0 | NO | NO | YES | YES | YES | NO
EQPRepairJ22111448 | T6373-01 | CWC8J0908-01 | Other | JB1TAS14 | PA5208KF-LT | J6G42 | 16 | AE Change Probe Card-A6668 | CPAUTO | Other | Other | 2022/08/25 14:35:45 | 2022/08/25 15:05:25 | 0.49 | OK | 13749 | 1.slot #16 啟家指示將原配件JB1TAS14 下卡歸庫磨針 並量測D/C/L. 2.更換剛磨針完成配件 JB1TAS28 上機測試 OK. 3.預溫重對針 點測PASS 溫針退針壓 OK PM(148.5) X=-3 Y=-2 4.啟家指示修改3D Clean 參數 WA:2000ea/OD:10/10下_CL:300ea/OD:20/10下 | NO | NO | NO | YES | YES | NO
EQPRepairJ22111450 | T6373-08 | CMC1J0048-01 | OverKill | VB104S01 | PA5252N | AXM342000 | 06 | # 06 已依誤宰處理流程處置無效 (O/S,bin46 過多) | 13081 | Over Kill | O/S | 2022/08/25 14:38:16 | 2022/08/25 15:05:14 | 0.45 | OK | 11385 | 清針調整針位，點測Bin46 Function FAIL穩定續測 | NO | NO | NO | YES | YES | NO
EQPRepairJ22111463 | T6373-151 | CWC8J0449-01.003 | OverKill | JB252S02 | TD7850-A0S-AZ-CP | SC1GT-B1121C | 18 | #14.16 EDA SHOW 2I O/S AND FAILBIN #18 TOO MNAY O/S-A6986 | CPAUTO | Over Kill | Function | 2022/08/25 15:48:13 | 2022/08/25 16:45:31 | 0.96 | OK | 19771 | 1.Bin20過多,配件磨針無效 2.更換備卡上機仍相同 3.請TA更換Lot上機驗證 | NO | NO | NO | YES | YES | NO
EQPRepairJ22111479 | T6373-86 | CWC8J0216-01 | Corr Fail | JB1Q4S16 | PA6542C | J683M | 14 | Corr Fail | A6893 | Corr | 產品問題 | 2022/08/25 16:51:33 | 2022/08/25 17:11:05 | 0.33 | OK | 11385 | Corr Fail---- T32 產品無 Corr Wafer Run Corr 比對 Fail | NO | NO | NO | YES | YES | NO
EQPRepairJ22111484 | T6373-01 | CWC8J0908-01 | OverKill | JB1TAS28 | PA5208KF-LT | J6G42 | 17 | BIN 14 | A6571 | Probe Card | Function Fail更換P/C | 2022/08/25 17:13:48 | 2022/08/25 18:05:12 | 0.86 | OK | 13749 | 1.配件JB1TAS28 BIN14 Function過多送修. 2.更換備卡#S25 PASS PM(148.5) SLOT #17 | NO | NO | NO | YES | YES | NO
EQPRepairJ22111491 | T6373-32 | CWC8J0352-01.005 | OverKill | JB1PNS06 | NB01761A(COF94AA)-A-C | H2NC84.1-BC1221 | 14 | too many O/S-A6668 | CPAUTO | Over Kill | 機台清針後Pass | 2022/08/25 17:53:36 | 2022/08/25 18:11:40 | 0.30 | OK | 16631 | 1.機台clean wafer清針後Pass，212.6 | NO | NO | NO | YES | YES | NO
EQPRepairJ22111496 | T6373-08 | CMC1J0047-01 | Tester Down | VB104S01 | PA5252N | AXM342000 | 02 | 無法使用自動化上貨 | 13081 | Tester | 自動化異常 | 2022/08/25 18:08:48 | 2022/08/25 18:39:19 | 0.51 | OK | 07910 | REMOTE---查看維修記錄，8/25 也有發生，自動化異常，先請產線cbsetup測試 | NO | NO | NO | NO | YES | NO
EQPRepairJ22111502 | T6373-147 | CWC8J0301-01 | OverKill | JB1UJS14 | PA5129LS_12U80H-LT | P07CN.1 | 14 | o/s-A6921 | CPAUTO | Over Kill | O/D不足 | 2022/08/25 18:26:45 | 2022/08/25 18:32:03 | 0.09 | OK | 19771 | OD不足外壓+5 觀察一段穩定還機續測 | NO | NO | NO | YES | YES | NO
EQPRepairJ22111505 | T6373-70 | CWC7J0J58-01 | Corr Fail | JB1Q4S18 | PA6542C | J683L | 10 | 多3顆回不掉的function bin 15 | 08341 | Corr | CORR FAIL | 2022/08/25 19:07:47 | 2022/08/25 19:26:40 | 0.31 | OK | 15895 | CORR FAIL，手動比對PASS | NO | NO | NO | YES | YES | NO
EQPRepairJ22111510 | T6373-32 | CWC8J0352-01.005 | Prober Down | JB1PNS06 | NB01761A(COF94AA)-A-C | H2NC84.1-BC1221 | 14 | AlarmCode:1703 AlarmMessage:THETA AXIS SOFTWARE ERROR (DRIVE RANGE OVER) !! | CPAUTO | Prober | Wafer Edge/Mode Not Found | 2022/08/25 19:27:35 | 2022/08/25 19:40:07 | 0.21 | OK | 03426 | Wafer 測試中 Clean Wafer 清針 擺放角度水平差異過大 退片請產線重新上貨 | NO | NO | NO | YES | YES | NO
EQPRepairJ22111518 | T6373-147 | CWC8J0301-01 | OverKill | JB1UJS14 | PA5129LS_12U80H-LT | P07CN.1 | 14 | too many o/s-A7164 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/25 19:58:28 | 2022/08/25 20:37:06 | 0.64 | OK | 17674 | 1. 清針Pass，更換備卡#12 pass 溫針 range y=10 x=20 (33,5)#14 | NO | NO | NO | YES | YES | NO
EQPRepairJ22111525 | T6373-43 | CEC7J0106-01 | OverKill | JB1UJS19 | PA5129LS_12U80H-LT | P08WT.1 | 18 | TOO MANY O/S | A6148 | Over Kill | 機台清針後Pass | 2022/08/25 20:40:17 | 2022/08/25 20:58:16 | 0.30 | OK | 15895 | 機台清針後Pass，觀察一段穩定 | NO | NO | NO | YES | YES | NO
EQPRepairJ22111530 | T6373-147 | CWC8J0301-01 | OverKill | JB1UJS12 | PA5129LS_12U80H-LT | P07CN.1 | 15 | MBD 2H O/S-A7164 | CPAUTO | Over Kill | 2H | 2022/08/25 20:57:31 | 2022/08/25 21:06:02 | 0.14 | OK | 15895 | EDA SHOW 2H，先前已更換P/C，後續測試良好，還機續觀察 | NO | NO | NO | YES | YES | NO
EQPRepairJ22111555 | T6373-151 | CWC8J0448-01 | OverKill | JB252S01 | TD7850-A0S-AZ-CP | SC1GT-B111C | 08 | too many fail bin20-A6148 | CPAUTO | Over Kill | Function | 2022/08/25 22:52:26 | 2022/08/25 23:00:43 | 0.14 | OK | 15895 | 區域性BIN20，P-P F-F，還機續觀察 | NO | NO | NO | YES | YES | NO
EQPRepairJ22111560 | T6373-151 | CWC8J0448-01 | OverKill | JB252S01 | TD7850-A0S-AZ-CP | SC1GT-B111C | 08 | too many fail bin-A6786 | CPAUTO | Over Kill | Function誤宰 | 2022/08/25 23:28:40 | 2022/08/25 23:45:05 | 0.27 | OK | 17542 | Remote處理 ----- Function Fail Bin20過多(wafer中心)，點測驗證P<>P F<>F穩定，且與之前同批測試時配件不同，故續測 | NO | NO | NO | YES | YES | NO
EQPRepairJ22111570 | T6373-33 | CWC8J0918-01 | OverKill | JB1TAS26 | PA5208KF-LT | J6CC7 | 12 | EDA SHOW 3"R" O/S | A6296 | Over Kill | SHOW R | 2022/08/26 00:17:59 | 2022/08/26 00:52:24 | 0.57 | OK | 13597 | 更換配件#14，RANGE 10 UM | NO | NO | NO | YES | YES | NO
EQPRepairJ22111574 | T6373-136 | CWC8J0870-01.002 | OverKill | JB1PNS08 | NB01761A(COF94AA)-A-C | H2NF33.1-BC12 | 17 | too many open short-A6786 | CPAUTO | Over Kill | O/S | 2022/08/26 00:37:28 | 2022/08/26 00:51:12 | 0.23 | OK | 15096 | Remote處理----o/s，調針、清針pass. | NO | NO | NO | YES | YES | NO
EQPRepairJ22111589 | T6373-147 | CWC7J0L69-01.002 | Corr Fail | JB1UJS12 | PA5129LS_12U80H-LT | P078M.1 | 16 | Corr Fail | A6881 | Corr | CORR FAIL | 2022/08/26 01:56:15 | 2022/08/26 02:13:07 | 0.28 | OK | 15895 | CORR FAIL，調針回收pass，p/m(34.0) | NO | NO | NO | YES | YES | NO
EQPRepairJ22111592 | T6373-32 | CWC8J0869-01.002 | OverKill | JB1PNS06 | NB01761A(COF94AA)-A-C | H2NF30.1-B1C22 | 04 | continues 2H-A7025 | CPAUTO | Over Kill | Show H | 2022/08/26 02:04:12 | 2022/08/26 02:07:33 | 0.06 | OK | 03426 | Show H-[CUSTOMER_HOLD_CRITERIA] [HW_Bin1<=90.00％]1453(IF Bin1 ＜＝ 90 ％, Please HOLD) -----請產線續測 | NO | NO | NO | YES | YES | NO
EQPRepairJ22111606 | T6373-39 | CWC8J0620-01 | OverKill | JB1UJS13 | PA5129LS_12U80H-LT | P07CK.1 | 23 | too many o/s-A6148 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/26 02:54:21 | 2022/08/26 03:06:08 | 0.20 | OK | 15895 | 二分法調針Y:12，P/M (34.0) | NO | NO | NO | YES | YES | NO
EQPRepairJ22111618 | T6373-141 | CWC8J0F03-01.002 | OverKill | JB1HKS07 | NA01687G(COG48BA)-A-AT-N2-C | GGT208.1-B1C22 | 05 | O/S TOO MANY-19697 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/26 03:51:00 | 2022/08/26 04:00:11 | 0.15 | OK | 03426 | 針位調整 機台清針PASS | NO | NO | NO | YES | YES | NO
EQPRepairJ22111620 | T6373-07 | CWC7J0374-01 | Tester Down | JB1TAS20 | PA5208KF-LT | J682P | 20 | AlarmCode:669 AlarmMessage:GP-IB COMMUNICATION TIMEOUT ERROR !! | CPAUTO | Product Issue | RTE | 2022/08/26 03:55:59 | 2022/08/26 04:12:23 | 0.27 | OK | 03426 | 1.#20測試至(133.1)RTE 2.重新連線ok | NO | NO | NO | YES | YES | NO
EQPRepairJ22111623 | T6373-86 | CWC8J0216-01 | OverKill | JB1Q4S16 | PA6542C | J683M | 19 | 2 CONTINOUS EDA SHOW H---NO CORR WAFER | A6926 | Over Kill | Show H | 2022/08/26 04:11:04 | 2022/08/26 04:16:43 | 0.09 | OK | 17542 | Remote處理 ----- Show H，[HW_BIN45>0.75%](Bin 46 > 6EA) 8，不同批有相同問題(J683K-F)，續測 | NO | NO | NO | YES | YES | NO
EQPRepairJ22111648 | T6373-136 | CWC8J0870-01.002 | OverKill | JB1PNS08 | NB01761A(COF94AA)-A-C | H2NF33.1-BC12 | 18 | TOO MANY O/S-A6148 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/26 05:55:36 | 2022/08/26 06:07:21 | 0.20 | OK | 15895 | 機台清針，p/m(52.8) | NO | NO | NO | YES | YES | NO
EQPRepairJ22111661 | T6373-141 | CWC8J0F03-01.002 | OverKill | JB1HKS07 | NA01687G(COG48BA)-A-AT-N2-C | GGT208.1-B1C22 | 07 | EDA SHOW 2H NO CORR WAFER-A6963 | CPAUTO | Over Kill | 2H | 2022/08/26 07:05:46 | 2022/08/26 07:28:41 | 0.38 | OK | 13749 | 1.EDA Show 2H BIN29 Function過多穩定. 2.同LOT 拆批至ND3-47 測試 也會跑 BIN29 Fail. 3.確認針痕 同方向貼下 Y 上調5um pass穩定 PM(43.5) please next wafer check slot #07 | NO | NO | NO | YES | YES | NO
EQPRepairJ22111662 | T6373-11 | CWC8J0F15-01.001 | Prober Down | JB8Q6S05 | NB01761A(COF94AA)-A-C | H2NF30.1-B1C11 | 01 | CANNOT MEASURE INDEX SIZE | A6926 | Prober | Wafer Edge/Mode Not Found | 2022/08/26 07:10:09 | 2022/08/26 07:32:23 | 0.37 | OK | 17680 | remote----手動抓model,check die ok,mark(52,0) | NO | NO | NO | YES | YES | NO
EQPRepairJ22111664 | T6373-47 | CWC8J0F03-01 | OverKill | JB1HKS10 | NA01687G(COG48BA)-A-AT-N2-C | GGT208.1-B1C1 | 16 | continous eda shows 2R DUE TO O/S-A6788 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/26 07:14:05 | 2022/08/26 07:35:52 | 0.36 | OK | 16631 | 1.調整針位，CLEAN WAFER清針，續測觀察 | NO | NO | NO | YES | YES | NO
EQPRepairJ22111675 | T6373-70 | CWC8J0618-01.007 | OverKill | JB1Q4S18 | PA6542C | J6C0W | 03 | too many o/s-A6660 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/26 08:08:12 | 2022/08/26 08:21:54 | 0.23 | OK | 20231 | 1.清針起測一段穩定pass，PM(160,1) | NO | NO | NO | YES | YES | NO
EQPRepairJ22111690 | T6373-151 | CWC8J0907-01.002 | OverKill | JB252S01 | TD7850-A0S-AZ-CP | SC3MG-B12C | 16 | #14.15 EDA SHOW 2I-A6986 | CPAUTO | Over Kill | 2I | 2022/08/26 09:14:12 | 2022/08/26 10:07:03 | 0.88 | OK | 16631 | 1.調整針位，機台清針續測觀察，請產線RUN CORR | NO | NO | NO | YES | YES | NO
EQPRepairJ22111705 | T6373-151 | CWC8J0907-01.002 | OverKill | JB252S01 | TD7850-A0S-AZ-CP | SC3MG-B12C | 16 | too many failbin 20-A6986 | CPAUTO | Over Kill | Function誤宰 | 2022/08/26 10:50:10 | 2022/08/26 11:03:15 | 0.22 | OK | 16631 | 1.WAFER中心區域BIN20，點測P-P/F-F穩定，先前已有換卡無效，疑似WAFER ISSUE，續測待 RUN CORR | NO | NO | NO | YES | YES | NO
EQPRepairJ22111713 | T6373-147 | CWC8J0301-01 | Tester Down | JB1UJS12 | PA5129LS_12U80H-LT | P07CN.1 | 18 | AlarmCode:669 AlarmMessage:GP-IB COMMUNICATION TIMEOUT ERROR !! | CPAUTO | Tester | 離線(OCT/GPIB error) | 2022/08/26 11:14:42 | 2022/08/26 11:25:07 | 0.17 | OK | 17712 | GPIB，重新連線OK。 | NO | NO | NO | YES | YES | NO
EQPRepairJ22111718 | T6373-112 | CWC8J0C42-01 | OverKill | JB1GVS23 | NC01682B(COG48AA)-A-AT-N-C | GGU463.1-B1C11 | 10 | o/s-a6962 | CPAUTO | Over Kill | 針位調整 | 2022/08/26 11:50:35 | 2022/08/26 12:23:56 | 0.56 | OK | 13749 | Remote--- 1.slot #10 連續 O/S 輪工確認PM(18.1)NG 同方向貼下擠壓. 2.重新對針 PASS PM(140.0) X=14 Y=12 | NO | NO | NO | YES | YES | NO
EQPRepairJ22111732 | T6373-147 | CWC8J0301-01 | OverKill | JB1UJS12 | PA5129LS_12U80H-LT | P07CN.1 | 18 | CONT. 2H AFTER CORR PASSED-a6962 | CPAUTO | Over Kill | 2H | 2022/08/26 12:52:33 | 2022/08/26 13:34:18 | 0.70 | OK | 16631 | 1.換卡#18 138.4 | NO | NO | NO | YES | YES | NO
EQPRepairJ22111735 | T6373-07 | CWC8J0A77-01 | OverKill | JB1TAS20 | PA5208KF-LT | J6CC9 | 03 | 2H no corr | A6049 | Over Kill | 2H | 2022/08/26 12:58:30 | 2022/08/26 13:28:31 | 0.50 | OK | 20231 | [HW_BIN12>0.5%](Bin 13 > 4EA) 10(Hold Wafer) 1.MAP中央區域性BIN13，#03目前測試無BIN13，確認針況針位OK，PM(147,1)，續測再觀察 | NO | NO | NO | YES | YES | NO
EQPRepairJ22111745 | T6373-86 | CWC8J0216-01 | OverKill | JB1Q4S16 | PA6542C | J683M | 24 | 2R O/S-a6571 | CPAUTO | Over Kill | O/S | 2022/08/26 13:51:08 | 2022/08/26 20:03:57 | 6.21 | OK | 15096 | <低溫型號>----- 1.日班交接配件#16測試o/s range 過小，更換配件#22上機測試o/s不穩， 2.再更換配件#15固定pin o/s且結霜，升溫烘烤 3.再更換配件#19上機等待20分鐘交接夜班等待降溫 4.夜班接手降溫後setup pass. | NO | NO | NO | YES | YES | NO
EQPRepairJ22111756 | T6373-128 | CWC7J0L15-01 | OverKill | TB1R8S83 | PA5137E1-LT | P097C.1 | 05 | EDA SHOWS 2H (NO CORR) | A6660 | Over Kill | 2H | 2022/08/26 14:43:14 | 2022/08/26 14:47:35 | 0.07 | OK | 20231 | #03 [HW_BIN44>0.3%](Bin 45 > 2EA) 3(Hold Wafer) #04 [HW_BIN0<95.5%](Bin 1 < 814EA) 809(Hold Wafer)----- 1.兩片不同原因，續測 | NO | NO | NO | YES | YES | NO
EQPRepairJ22111759 | T6373-01 | CWC8J0908-01 | OverKill | JB1TAS25 | PA5208KF-LT | J6G42 | 25 | o/s-a6571 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/26 15:05:17 | 2022/08/26 15:17:49 | 0.21 | OK | 20231 | 1.清針起測一段穩定pass，PM(147,1) | NO | NO | NO | YES | YES | NO
EQPRepairJ22111772 | T6373-12 | CWC8J0G46-01 | OverKill | JBCK3S02 | NB01853E(COG48AA)-A-AT-X-C | GGP200.1-BC | 03 | #2 EDA SHOW I(T44) / NO CORR-19617 | CPAUTO | Over Kill | show I | 2022/08/26 16:18:25 | 2022/08/26 16:21:43 | 0.06 | OK | 17680 | single I,keep test | NO | NO | NO | YES | YES | NO
EQPRepairJ22111773 | T6373-47 | CWC8J0F03-01 | OverKill | JB1HKS10 | NA01687G(COG48BA)-A-AT-N2-C | GGT208.1-B1C1 | 21 | O/S-17436 | CPAUTO | Over Kill | 機台清針後Pass | 2022/08/26 16:20:04 | 2022/08/26 16:37:01 | 0.28 | OK | 13749 | Remote---Slot #21 Clean wafer pass PM Check (146.4) OK | NO | NO | NO | YES | YES | NO
EQPRepairJ22111775 | T6373-112 | CWC8J0C42-01 | OverKill | JB1GVS23 | NC01682B(COG48AA)-A-AT-N-C | GGU463.1-B1C11 | 15 | o/s-A6962 | CPAUTO | Over Kill | 配件殘金 | 2022/08/26 16:32:53 | 2022/08/26 17:19:04 | 0.77 | OK | 13749 | Remote--- 1.slot #15 O/S過多不穩，配件針髒+黑點. 2.更換備卡#S23->#S24 OK 預溫重對針 點測PASS 溫針退針壓OK. 3.PM(43.6) X=10 Y=14 | NO | NO | NO | YES | YES | NO
EQPRepairJ22111797 | T6373-29 |  | Prober Down | TB1INS26 | NA01404H-A-AT-C | GGV701.1-BC2 |  | AlarmCode:1058 AlarmMessage:MODEL NOT FOUND ( FINE ALIGNMENT ) !! | CPAUTO | Prober | Wafer Edge/Mode Not Found | 2022/08/26 18:06:01 | 2022/08/26 18:12:52 | 0.11 | OK | 13749 | Remote---Slot #02 重抓mode ok ，check die & PM(55.0) Check ok | NO | NO | NO | YES | YES | NO
EQPRepairJ22111799 | T6373-70 | CWC8J0618-01.007 | OverKill | JB1Q4S18 | PA6542C | J6C0W | 17 | o/s-a6571 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/26 18:20:40 | 2022/08/26 18:33:58 | 0.22 | OK | 17712 | 機台清針pass，確認針痕針壓ok，#17(158.1) | NO | NO | NO | YES | YES | NO
EQPRepairJ22111802 | T6373-151 | CWC8J0907-01.002 | OverKill | JB252S01 | TD7850-A0S-AZ-CP | SC3MG-B12C | 22 | CHANGE TO MBD AS PER LEADER DUE TO BIN10 AND BIN20-A6660 | CPAUTO | Over Kill | Function誤宰 | 2022/08/26 18:25:39 | 2022/08/26 18:59:27 | 0.56 | OK | 16631 | 1.產品盧思妤要求:Bin10,20偏多且集中於wafer中間，煩請協助swap PE Card 2.Swap PE Card #01<->#03，#02<->#04，Cal init | NO | NO | NO | YES | YES | NO
EQPRepairJ22111804 | T6373-147 | CWC8J0301-01 | OverKill | JB1UJS18 | PA5129LS_12U80H-LT | P07CN.1 | 20 | o/s-A6962 | CPAUTO | Over Kill | 機台清針後Pass | 2022/08/26 18:42:34 | 2022/08/26 19:03:53 | 0.36 | OK | 15897 | 針壓ok.修改參數清針Pass.#20(12.4).測一段ok. | NO | NO | NO | YES | YES | NO
EQPRepairJ22111809 | T6373-151 |  | Product Issue | JB252S01 | TD7850-A0S-AZ-CP | SC3MG-B12C |  | 1.產品盧思妤要求:Bin10,20偏多且集中於wafer中間，煩請協助swap PE Card 2.Swap PE Card #01<->#03，#02<->#04，Cal init | 16631 | Product Issue | Other | 2022/08/26 19:00:03 | 2022/08/26 21:31:46 | 2.53 | OK | 17674 | 產品當------ 1. 產品指示 bin20、10偏高， 日班 swap PE Card #01<>03、#02<>04 2. 接手 cal 機 pass，上機 Wafer 中間一樣有bin20 問題，點測 P-P F-F 3. 重新上機PASS， 溫針 RANGE Y=8 (201,4)#22 | NO | NO | NO | YES | YES | NO
EQPRepairJ22111817 | T6373-29 | CWC8J0C85-01.001 | OverKill | TB1INS26 | NA01404H-A-AT-C | GGV701.1-BC2 | 02 | Already follow Overkilll handling procedures but still not effective too many O/S. | A6937 | Over Kill | O/S誤宰 | 2022/08/26 19:57:18 | 2022/08/26 20:13:07 | 0.26 | OK | 03426 | 針位調整PASS | NO | NO | NO | YES | YES | NO
EQPRepairJ22111823 | T6373-136 | CWC8J0870-01.002 | OverKill | JB1PNS08 | NB01761A(COF94AA)-A-C | H2NF33.1-BC12 | 23 | TOO MANY O/S-A6148 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/26 20:15:46 | 2022/08/26 21:20:01 | 1.07 | OK | 17674 | 1. 下卡磨針Pass，#23 (112,4)溫針 range Y=10 x=20 (212,6) | NO | NO | NO | YES | YES | NO
EQPRepairJ22111836 | T6373-07 | CWC8J0A77-01 | OverKill | JB1TAS20 | PA5208KF-LT | J6CC9 | 05 | too many o/s..bin2-A7025 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/26 20:51:26 | 2022/08/26 21:20:33 | 0.49 | OK | 03426 | 針位調整PASS | NO | NO | NO | YES | YES | NO
EQPRepairJ22111842 | T6373-12 | CWC8J0G46-01 | OverKill | JBCK3S02 | NB01853E(COG48AA)-A-AT-X-C | GGP200.1-BC1 | 20 | #18 eda show i,no corr-A5717 | CPAUTO | Over Kill | Show I | 2022/08/26 21:15:27 | 2022/08/26 21:19:07 | 0.06 | OK | 17542 | Remote 處理 ----- Show I，T44單片Show I，確認EDA資料，不同Dut Low Yield造成(無固定Dut異常，且無O/S過多)，續測 | NO | NO | NO | YES | YES | NO
EQPRepairJ22111845 | T6373-35 |  | Corr Fail | JB1HKS06 | NA01687F(COG48BA)-A-AT-C | GGR359.1-B |  | Setup corr. failed | A6053 | Corr | Other | 2022/08/26 21:20:03 | 2022/08/26 21:21:56 | 0.03 | OK | 03426 | Corr 良率比 Corr Wafer 好，前流程 Fail Bin 回收成 Pass，手動比對 Corr Pass | NO | NO | NO | YES | YES | NO
EQPRepairJ22111847 | T6373-38 | CWC8J0841-01 | OverKill | JB1GKS09 | PA6543B | J6AK2 | 07 | too many o/s,(47.7)pm check,pm up/down,od8,p/c has particle-A5717 | CPAUTO | Over Kill | 配件卡毛 | 2022/08/26 21:27:45 | 2022/08/26 22:27:14 | 0.99 | OK | 03426 | 卡毛下卡磨針 5.5 | NO | NO | NO | YES | YES | NO
EQPRepairJ22111854 | T6373-91 |  | Tester Down | TB1R8S77 | PA5223M1_12U60H-LT | PX683.1 |  | failed to load main program-A6525 | CPAUTO | Tester | 無法Download程式 | 2022/08/26 22:15:32 | 2022/08/26 22:26:50 | 0.19 | OK | 17542 | 無法Download程式 --- 1.確認此型號是否有TDR Fail，無TDR異常，但Show Tester was not INIT 2.Cal Data recover PASS，重Load程式ok，請產線重新上貨 | NO | NO | NO | YES | YES | NO
EQPRepairJ22111858 | T6373-16 | CEC8J0085-01 | OverKill | VB405S07 | PA5208K-LT | QTC9L | 03 | too many open short-A6786 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/26 22:20:16 | 2022/08/26 22:47:50 | 0.46 | OK | 15895 | O/S誤宰，針壓不足，確認ok | NO | NO | NO | YES | YES | NO
EQPRepairJ22111860 | T6373-47 | CWC8J0F03-01 | OverKill | JB1HKS10 | NA01687G(COG48BA)-A-AT-N2-C | GGT208.1-B1C1 | 24 | WHILE REPROBING STILL OPEN SHORT-A6786 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/26 22:50:05 | 2022/08/26 23:13:06 | 0.38 | OK | 15895 | 機台清針，二分法調針y:10，p/m(35.4) | NO | NO | NO | YES | YES | NO
EQPRepairJ22111864 | T6373-91 |  | Tester Down | TB1R8S82 | PA5223M1_12U60H-LT | PXCG6.1 |  | long time load testprogram | A6926 | Other | Other | 2022/08/26 23:08:48 | 2022/08/26 23:14:49 | 0.10 | OK | 15096 | VP重新登入，請產線重load程式. | NO | NO | NO | YES | YES | NO
EQPRepairJ22111872 | T6373-38 | CWC8J0841-01 | OverKill | JB1GKS09 | PA6543B | J6AK2 | 07 | too many o/s..bin2-A7025 | CPAUTO | Over Kill | 機台清針後Pass | 2022/08/26 23:53:10 | 2022/08/27 00:09:39 | 0.27 | OK | 17674 | 機台清針後Pass (173,5) #07 | NO | NO | NO | YES | YES | NO
EQPRepairJ22111874 | T6373-47 | CWC8J0F03-01 | OverKill | JB1HKS10 | NA01687G(COG48BA)-A-AT-N2-C | GGT208.1-B1C1 | 25 | too many o/s-A6148 | CPAUTO | Over Kill | O/D不足 | 2022/08/26 23:55:13 | 2022/08/27 00:40:09 | 0.75 | OK | 17674 | 1. 重對針PASS，#25 (90,2)溫針 RANGE y=6 (145,0) | NO | NO | NO | YES | YES | NO
EQPRepairJ22111877 | T6373-39 | CWC8J0184-01.001 | Prober Down | JB1UJS13 | PA5129LS_12U80H-LT | P07CL.1 | 12 | AlarmCode:1995 AlarmMessage:Faied to upload log file | CPAUTO | Prober | Faied to upload log file | 2022/08/27 00:12:57 | 2022/08/27 00:18:00 | 0.08 | OK | 17542 | Remote 處理 ----- Faied to upload log file，確認機台狀況，關閉異常視窗，機台停機並起測(無異常問題)，請產線續測 | NO | NO | NO | YES | YES | NO
EQPRepairJ22111881 | T6373-91 | CWBBJ0N96-01.005 | Prober Down | TB1R8S77 | PA5223M1_12U60H-LT | PX683.1 | 17 | CANNOT MEASURE INDEX SIZE-A7024 | CPAUTO | Prober | Model Not Found | 2022/08/27 00:30:32 | 2022/08/27 00:38:47 | 0.14 | OK | 17542 | Remote 處理 ----- Model Not Found，重新抓取ok，check die ok，P/M(5,4)無異常 | NO | NO | NO | YES | YES | NO
EQPRepairJ22111883 | T6373-151 | CWC8J0907-01.002 | Other | JB252S01 | TD7850-A0S-AZ-CP | SC3MG-B12C | 22 | PLEASE CHECK P/M HAS P/M SHIFT (146,5) ALREADY ADJUST P/M DOWN-A6053 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/27 00:36:21 | 2022/08/27 00:47:22 | 0.18 | OK | 17674 | 1. 回收OPEN 針位貼上 (201,4)#22 | NO | NO | NO | YES | YES | NO
EQPRepairJ22111887 | T6373-136 |  | Prober Down | JB1PNS08 | NB01761A(COF94AA)-A-C | H2NF33.1-BC12 |  | model not found-A6148 | CPAUTO | Prober | Wafer Edge/Mode Not Found | 2022/08/27 00:59:10 | 2022/08/27 01:08:23 | 0.15 | OK | 15895 | Mode Not Found 手動抓取ok，check die ok，p/m51.8 | NO | NO | NO | YES | YES | NO
EQPRepairJ22111889 | T6373-38 | CWC8J0841-01 | OverKill | JB1GKS09 | PA6543B | J6AK2 | 08 | too many os-A6963 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/27 01:09:39 | 2022/08/27 01:24:18 | 0.24 | OK | 03426 | 重MNT PASS | NO | NO | NO | YES | YES | NO
EQPRepairJ22111891 | T6373-151 | CWC8J0907-01.002 | OverKill | JB252S01 | TD7850-A0S-AZ-CP | SC3MG-B12C | 23 | #20.22 EDA SHOW 2I-A5717 | CPAUTO | Over Kill | Show I | 2022/08/27 01:22:01 | 2022/08/27 01:35:10 | 0.22 | OK | 17542 | Remote 處理 ----- 1.Show I，[HW_BIN20>0.5%]366，Bin20過多 2.先前維修Tester Swap PE Card #01<>03、#02<>04，Cal Pass，上機 Wafer 中間一樣有Bin20 問題，點測 P-P F-F 3.請產線先行續測，待後續更換貨批確認 | NO | NO | NO | YES | YES | NO
EQPRepairJ22111919 | T6373-151 | CWC8J0907-01.002 | OverKill | JB252S01 | TD7850-A0S-AZ-CP | SC3MG-B12C | 23 | TOO MANY O/S PM NOT GOOD UP/DOWN | A6926 | Over Kill | 針位調整 | 2022/08/27 02:59:22 | 2022/08/27 03:10:24 | 0.18 | OK | 15895 | 二分法調針y:8，p/m(148.7) | NO | NO | YES | YES | YES | NO
EQPRepairJ22111922 | T6373-01 | CWC7J0L66-01 | OverKill | JB1TAS25 | PA5208K-LT | QT800 | 08 | too many o/s,(175pm,od10-A5717 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/27 03:07:04 | 2022/08/27 03:12:14 | 0.09 | OK | 03426 | 針位調整 機台清針PASS | NO | NO | NO | YES | YES | NO
EQPRepairJ22111924 | T6373-147 | CWC8J0301-01 | OverKill | JB1UJS18 | PA5129LS_12U80H-LT | P07CN.1 | 23 | too many o/s-A7160 | CPAUTO | Over Kill | 機台清針後Pass | 2022/08/27 03:30:43 | 2022/08/27 03:42:21 | 0.19 | OK | 15895 | 機台清針後Pass，p/m 116.5 | NO | NO | NO | YES | YES | NO
EQPRepairJ22111926 | T6373-86 | CWC8J0618-01.004 | Prober Down | JB1Q4S19 | PA6542C | J68S1 | 13 | model not found | A6926 | Prober | Wafer Edge/Mode Not Found | 2022/08/27 03:33:03 | 2022/08/27 03:39:37 | 0.11 | OK | 03426 | 重抓Module 查檢Check each ok | NO | NO | NO | YES | YES | NO
EQPRepairJ22111937 | T6373-34 | CWC8J0C40-01 | OverKill | TB1R8S85 | PA5223Q_12U60H-LT | P0CLN.1 | 05 | too many o/s..bin59-A7025 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/27 04:19:44 | 2022/08/27 04:34:47 | 0.25 | OK | 03426 | 針位調整PASS | NO | NO | NO | YES | YES | NO
EQPRepairJ22111940 | T6373-07 | CWC8J0A77-01 | OverKill | JB1TAS20 | PA5208KF-LT | J6CC9 | 07 | o/s..bin2-A7025 | CPAUTO | Over Kill | 機台清針後Pass | 2022/08/27 04:36:58 | 2022/08/27 05:49:22 | 1.21 | OK | 03426 | 機台清針PASS | NO | NO | NO | YES | YES | NO
EQPRepairJ22111955 | T6373-70 | CWC8J0618-01.007 | OverKill | JB1Q4S18 | PA6542C | J6C0W | 22 | too many o/s-A6148 | CPAUTO | Over Kill | 機台清針後Pass | 2022/08/27 05:29:24 | 2022/08/27 05:50:55 | 0.36 | OK | 15895 | 機台清針後Pass，P/M130.0 | NO | NO | NO | YES | YES | NO
EQPRepairJ22111963 | T6373-21 |  | Prober Down | JB1GKS08 | PA6543B | J6AK5 |  | AlarmCode:1058 AlarmMessage:MODEL NOT FOUND ( FINE ALIGNMENT ) !! | CPAUTO | Prober | Wafer Edge/Mode Not Found | 2022/08/27 06:28:35 | 2022/08/27 06:31:55 | 0.06 | OK | 03426 | 重抓Module 查檢Check each ok | NO | NO | NO | YES | YES | NO
EQPRepairJ22111966 | T6373-136 | CWC8J0855-01.001 | OverKill | JB1PNS08 | NB01761A(COF94AA)-A-C | H2NF20.1-B1C21 | 01 | too many o/s-A6148 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/27 06:32:34 | 2022/08/27 06:49:01 | 0.27 | OK | 17674 | 機台清針後Pass | NO | NO | NO | YES | YES | NO
EQPRepairJ22111969 | T6373-07 | CWC8J0A77-01 | OverKill | JB1TAS20 | PA5208KF-LT | J6CC9 | 07 | Too many o/s-A6053 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/27 06:35:55 | 2022/08/27 06:53:53 | 0.30 | OK | 03426 | 針位調整PASS | NO | NO | NO | YES | YES | NO
EQPRepairJ22111971 | T6373-147 | CWC8J0301-01 | OverKill | JB1UJS18 | PA5129LS_12U80H-LT | P07CN.1 | 24 | #24 Already follow Overkilll handling procedures but still not effective too many o/s 34.5 pm updown od 11, | A6881 | Over Kill | O/S誤宰 | 2022/08/27 06:45:32 | 2022/08/27 06:57:44 | 0.20 | OK | 15895 | 機台清針，二分法調針y:18，p/m(116.5) | NO | NO | NO | YES | YES | NO
EQPRepairJ22111972 | T6373-47 | CWC8J0F64-01.002 | OverKill | JB1HKS10 | NA01687G(COG48BA)-A-AT-N2-C | GGT523.1-BC22 | 05 | continous eda shows 2R due to o/s-A6788 | CPAUTO | Over Kill | 機台清針後Pass | 2022/08/27 06:49:07 | 2022/08/27 07:31:05 | 0.70 | OK | 20075 | 機台清針後Pass,range 10,mark 182,2 | NO | NO | NO | YES | YES | NO
EQPRepairJ22111996 | T6373-35 | CWC8J0347-01.005 | OverKill | JB1HKS06 | NA01687F(COG48BA)-A-AT-C | GGU954.1-BC12211 | 16 | open/short please help check PM not good-A6858 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/27 08:41:04 | 2022/08/27 09:43:02 | 1.03 | OK | 15362 | 二分法調˙針,y range:18,p/m(166,1) | NO | YES | NO | YES | YES | NO
EQPRepairJ22112015 | T6373-147 | CWC8J0301-01 | OverKill | JB1UJS18 | PA5129LS_12U80H-LT | P07CN.1 | 25 | O/S | 20427 | Over Kill | O/D不足 | 2022/08/27 09:40:10 | 2022/08/27 09:54:59 | 0.25 | OK | 20075 | 機台清針,od 11->15 | NO | NO | NO | YES | YES | NO
EQPRepairJ22112018 | T6373-16 |  | Prober Down | VB405S07 | PA5208K-LT | QTC9L |  | AlarmCode:1995 AlarmMessage:Faied to upload log file | CPAUTO | Prober | Faied to upload log file | 2022/08/27 09:46:52 | 2022/08/27 09:49:01 | 0.04 | OK | 10210 | 消音後OK | NO | NO | NO | YES | YES | NO
EQPRepairJ22112024 | T6373-07 | CWC8J0A77-01 | OverKill | JB1TAS20 | PA5208KF-LT | J6CC9 | 08 | o/s-11550 | CPAUTO | Over Kill | 機台清針後Pass | 2022/08/27 10:05:06 | 2022/08/27 10:50:05 | 0.75 | OK | 13222 | 機台清針後Pass | NO | NO | NO | YES | YES | NO
EQPRepairJ22112025 | T6373-42 | CWC8J0968-01 | OverKill | JB1HKS03 | NA01687G(COG48BA)-A-AT-N2-C | GGR354.1-B22C1 | 19 | too many fail bin 29-A6785 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/27 10:16:14 | 2022/08/27 10:23:20 | 0.12 | OK | 10210 | Unit清針 | NO | NO | NO | YES | YES | NO
EQPRepairJ22112041 | T6373-01 | CWC7J0L58-01 | OverKill | JB1TAS25 | PA5208K-LT | QT7CH | 07 | o/s-11550 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/27 11:23:11 | 2022/08/27 11:38:06 | 0.25 | OK | 15362 | 清調針,y range:10,p/m(147,5) | NO | NO | NO | YES | YES | NO
EQPRepairJ22112048 | T6373-103 | CWC8J0963-01.001 | OverKill | JB1CES10 | NA01248F(COG4)-A-AT-X2-C | NT2YC-1B2C2 | 01 | Already follow overkill handling procedures but still not effective too many o/s. | A6564 | Over Kill | O/S誤宰 | 2022/08/27 12:20:03 | 2022/08/27 12:31:00 | 0.18 | OK | 15362 | 清調針, y range:12,p/m(219,7) | NO | NO | NO | YES | YES | NO
EQPRepairJ22112067 | T6373-07 | CWC8J0A77-01 | OverKill | JB1TAS20 | PA5208KF-LT | J6CC9 | 09 | o/s-11550 | CPAUTO | Over Kill | 配件殘金 | 2022/08/27 13:22:23 | 2022/08/27 14:33:04 | 1.18 | OK | 15362 | 更換配件#26 pass,y range:18,p/m(46,6) | NO | NO | NO | YES | YES | NO
EQPRepairJ22112070 | T6373-152 | CWC8J0455-01.001 | OverKill | JB252S02 | TD7850-A0S-AZ-CP | SC1GS-B1C | 10 | TOO MANY O/S-A6873 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/27 13:39:27 | 2022/08/27 13:59:09 | 0.33 | OK | 20075 | 清調針PASS,RANGE 10 | NO | NO | NO | YES | YES | NO
EQPRepairJ22112074 | T6373-147 | CWC8J0227-01.001 | OverKill | JB1UJS18 | PA5129LS_12U80H-LT | P07CJ.1 | 02 | #2 O/S | 20427 | Over Kill | O/S誤宰 | 2022/08/27 13:58:29 | 2022/08/27 14:07:51 | 0.16 | OK | 20075 | 清調針PASS,RANGE 12,MARK -1,2 | NO | NO | NO | YES | YES | NO
EQPRepairJ22112096 | T6373-33 | CWC8J0C41-01 | OverKill | JB1TAS14 | PA5208KF-LT | J6CHG | 03 | too many o/s-A6293 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/27 15:41:57 | 2022/08/27 16:18:52 | 0.62 | OK | 13222 | REMOTE--- 1. 機台清針OK | NO | NO | NO | YES | YES | NO
EQPRepairJ22112100 | T6373-151 |  | Setup Fail | JB252S01 | TD7850-A0S-AZ-CP | SC1GT-B111 |  | setup fail due to bin20 dut 1and2 fail no other p/c already unmount the tester but sill same p/c polish-A6143 | CPAUTO | Tester | bin20 | 2022/08/27 15:54:20 | 2022/08/27 18:37:42 | 2.72 | OK | 20075 | 1. 機台SETUP 中間區域跑bin20,下卡磨針,換卡無效 2. 查看紀錄,之前有同樣情況 3. 目前無維修方向，請協助改測，謝謝 | NO | NO | YES | YES | NO | YES
EQPRepairJ22112103 | T6373-35 | CWC8J0291-01.009 | Tester Down | JB1QNS06 | NA01687F(COG48BA)-A-AT-C | GGU014.1-BC121121 | 06 | GP IB-A6858 | CPAUTO | Other | Other | 2022/08/27 16:10:10 | 2022/08/27 16:31:34 | 0.36 | OK | 15362 | #6退片請產線重上 | NO | NO | NO | YES | YES | NO
EQPRepairJ22112110 | T6373-47 |  | Prober Down | JB1HKS10 | NA01687G(COG48BA)-A-AT-N2-C | GGT523.1-BC22 |  | AlarmCode:1057 AlarmMessage:MODEL NOT FOUND ( ROUGH ALIGNMENT ) !! | CPAUTO | Prober | Mode Not Found | 2022/08/27 16:45:52 | 2022/08/27 16:52:48 | 0.12 | OK | 20075 | 抓MODE OK,CHECK DIE OK | NO | NO | NO | YES | YES | NO
EQPRepairJ22112121 | T6373-152 | CWC8J0455-01.001 | Other | JB252S02 | TD7850-A0S-AZ-CP | SC1GS-B1C | 15 | AE CHANGE PC-A6873 | CPAUTO | Other | 換卡驗證 | 2022/08/27 18:12:10 | 2022/08/27 18:27:16 | 0.25 | OK | 20075 | S02->S01 pass,range 12 | NO | NO | NO | YES | YES | NO
EQPRepairJ22112124 | T6373-12 |  | Prober Down | JB6H1S05 | NC01690I(COG46BA)-A-AT-C | GGW436.1-BC |  | AlarmCode:1057 AlarmMessage:MODEL NOT FOUND ( ROUGH ALIGNMENT ) !! | CPAUTO | Prober | Wafer Edge/Mode Not Found | 2022/08/27 18:18:04 | 2022/08/27 18:45:24 | 0.46 | OK | 14291 | REMOTE---重抓model and check die/針位 ok | NO | NO | NO | YES | YES | NO
EQPRepairJ22112127 | T6373-33 | CWC8J0C41-01 | OverKill | JB1TAS14 | PA5208KF-LT | J6CHG | 04 | CONTINOUS 2 WAFERS EDA SHOW R O/S-A6317 | CPAUTO | Over Kill | EDA 2R | 2022/08/27 18:24:53 | 2022/08/27 19:08:06 | 0.72 | OK | 20230 | REMOTE 1.EDA 2R 2.更換備卡S19,RG=17,PM(47.6) 3.PASS | NO | NO | NO | YES | YES | NO
EQPRepairJ22112130 | T6373-151 | CWC8J0448-01.006 | Tester Down | JB252S01 | TD7850-A0S-AZ-CP | SC3MG-B12C | 03 | 1. 機台SETUP 中間區域跑bin20,下卡磨針,換卡無效 2. 查看紀錄,之前有同樣情況 3. 目前無維修方向，請協助改測，謝謝 | 20075 | Over Kill | Function誤宰 | 2022/08/27 18:38:45 | 2022/08/27 21:53:01 | 3.24 | OK | 20277 | 輪工處理---- 1.Setup時發現Wafer中心區域性Bin20，交叉點測穩定，換卡無效 2.查看紀錄先前已Swap PE仍有相同問題 3.無維修方向，請DN & 產品協助確認後先行還機生產 4.Run Corr Run邊緣比對PASS | NO | NO | YES | YES | YES | NO
EQPRepairJ22112135 | T6373-128 | CWC7J0337-01.003 | OverKill | TB1R8S83 | PA5137F-LT1 | P07NU.1 | 18 | TOO MANY O/S-A6873 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/27 19:05:54 | 2022/08/27 19:33:45 | 0.46 | OK | 20277 | Remote---- 1.Range:X>20，Y:10 機台清針 起測一段PASS穩定，還機生產 | NO | NO | NO | YES | YES | NO
EQPRepairJ22112147 | T6373-42 | CWC8J0F02-01.003 | OverKill | JB1HKS03 | NA01687G(COG48BA)-A-AT-N2-C | GGT207.1-BC121 | 02 | too many bin29-A6512 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/27 20:00:02 | 2022/08/27 20:16:08 | 0.27 | OK | 18853 | 機台清針，R:12 #2(146,4) | NO | NO | NO | YES | YES | NO
EQPRepairJ22112161 | T6373-01 | CEC8J0078-01 | OverKill | JB1TAS25 | PA5208K-LT | QTAAS | 04 | #4 Already follow the overkilling procedure but still OS | A6896 | Over Kill | O/S誤宰 | 2022/08/27 20:36:14 | 2022/08/27 21:27:27 | 0.85 | OK | 10110 | 更換備卡#30上機PASS | NO | NO | NO | YES | YES | NO
EQPRepairJ22112195 | T6373-151 | CWC8J0572-01.007 | OverKill | JB252S02 | TD7850-A0S-AZ-CP | SC1GR-B1111C | 06 | Already follow Overkilll handling procedures but still not effective too many O/S od8 no particle | A6943 | Over Kill | O/S誤宰 | 2022/08/27 22:58:36 | 2022/08/27 23:12:38 | 0.23 | OK | 20277 | Remote---- 1.起測O/S 針位貼下 Range:10 測是一段PASS穩定，還機生產 | NO | NO | YES | YES | YES | NO
EQPRepairJ22112207 | T6373-42 | CWC8J0F66-01.001 | OverKill | JB1HKS03 | NA01687G(COG48BA)-A-AT-X1-C | GGT528.1-BC2 | 01 | too many os-A6512 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/27 23:42:30 | 2022/08/28 00:12:53 | 0.51 | OK | 18853 | 機台清針，r:15 #1(146,4) | NO | NO | NO | YES | YES | NO
EQPRepairJ22112213 | T6373-152 | CWC8J0455-01.001 | OverKill | JB252S01 | TD7850-A0S-AZ-CP | SC1GS-B1C | 22 | EDA SHOW 2H-A6857 | CPAUTO | Over Kill | Show H | 2022/08/28 00:28:24 | 2022/08/28 00:35:22 | 0.12 | OK | 20277 | 輪工處理---- 1.2H 不同批同機卡皆有相同問題 LOT:SA1MC-T 請產線Run Corr驗證，續測觀察 | NO | NO | NO | YES | YES | NO
EQPRepairJ22112216 | T6373-151 | CWC8J0572-01.007 | OverKill | JB252S02 | TD7850-A0S-AZ-CP | SC1GR-B1111C | 06 | too many os-A6512 | CPAUTO | Over Kill | 針位調整 | 2022/08/28 00:43:47 | 2022/08/28 00:59:26 | 0.26 | OK | 18853 | 二分法R:10 #6(147,7) | NO | NO | YES | YES | YES | NO
EQPRepairJ22112252 | T6373-07 | CWC8J0A77-01 | OverKill | JB1TAS26 | PA5208KF-LT | J6CC9 | 14 | #7 Already follow the overkilling procedure but still OS | A6896 | Over Kill | O/S誤宰 | 2022/08/28 03:27:04 | 2022/08/28 03:46:33 | 0.32 | OK | 17091 | 清針+二分法 | NO | NO | NO | YES | YES | NO
EQPRepairJ22112254 | T6373-70 | CWC8J0706-01.003 | OverKill | JB1Q4S18 | PA6542C | J6C10 | 11 | TOO MANY O/S..NO PARTICLE-A6516 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/28 03:46:03 | 2022/08/28 04:07:43 | 0.36 | OK | 20277 | Remote---- 1.加針壓 Range:14 機台清針 PASS | NO | NO | NO | YES | YES | NO
EQPRepairJ22112259 | T6373-33 | CWC8J0C41-01 | OverKill | JB1TAS19 | PA5208KF-LT | J6CHG | 06 | O/S TOO MANY-A6318 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/28 04:09:40 | 2022/08/28 04:29:07 | 0.32 | OK | 17184 | 1.清針&調針(3.2) | NO | NO | NO | NO | YES | NO
EQPRepairJ22112282 | T6373-38 | CWC8J0752-01.001 | Prober Down | JB1GKS09 | PA6543B | J6AK6 | 07 | AlarmCode:7273 AlarmMessage:Result of Die Origin Correction is Abnormal !! | CPAUTO | Prober | Other | 2022/08/28 06:19:21 | 2022/08/28 06:23:34 | 0.07 | OK | 17091 | remote--start重抓ok | NO | NO | NO | YES | YES | NO
EQPRepairJ22112284 | T6373-39 | CWC8J0369-01 | OverKill | JB1UJS13 | PA5129LS_12U80H-LT | P0GRF.1 | 06 | TOO MANY OS-A6512 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/28 06:28:33 | 2022/08/28 07:03:41 | 0.59 | OK | 18853 | 重新對針、溫針r:10 #6(116,5) | NO | NO | NO | YES | YES | NO
EQPRepairJ22112290 | T6373-152 | CWC8J0455-01.001 | OverKill | JB252S01 | TD7850-A0S-AZ-CP | SC1GS-B1C | 25 | EDA SHOW 2R O/S-A6857 | CPAUTO | Over Kill | 2r | 2022/08/28 06:54:50 | 2022/08/28 07:21:34 | 0.45 | OK | 20075 | 續測 | NO | NO | NO | YES | YES | NO
EQPRepairJ22112313 | T6373-16 | CWC8J0G90-01 | OverKill | VB405S07 | PA5208KF-LT | J6FGK | 04 | too many o/s-A6919 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/28 08:33:42 | 2022/08/28 08:47:14 | 0.23 | OK | 10210 | OD+5um Pass | NO | NO | NO | YES | YES | NO
EQPRepairJ22112314 | T6373-07 | CWC8J0A77-01 | OverKill | JB1TAS26 | PA5208KF-LT | J6CC9 | 16 | too many O/S-A6873 | CPAUTO | Over Kill | o/s | 2022/08/28 08:35:56 | 2022/08/28 08:55:15 | 0.32 | OK | 13222 | 機台清針ok | NO | NO | NO | YES | YES | NO
EQPRepairJ22112315 | T6373-152 | CWC8J0455-01.001 | OverKill | JB252S01 | TD7850-A0S-AZ-CP | SC1GS-B1C | 25 | TOO MANY O/S-A6919 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/28 08:36:27 | 2022/08/28 08:51:04 | 0.24 | OK | 20075 | 清調針pass,range 10 | NO | NO | NO | YES | YES | NO
EQPRepairJ22112317 | T6373-121 | CWC8J0F02-01.001 | OverKill | JB1QNS04 | NA01687G(COG48BA)-A-AT-N2-C | GGT207.1-BC21 | 14 | #14 eda continues show 2h no corr wafer-A6132 | CPAUTO | Over Kill | EDA 2H | 2022/08/28 08:53:36 | 2022/08/28 10:54:03 | 2.01 | OK | 20230 | REMOTE 1.EDA 2H 2.更換備卡S05,RG=20,PM(37.0) 3.PASSS | NO | NO | NO | YES | YES | NO
EQPRepairJ22112329 | T6373-136 | CWC8J0871-01.003 | Prober Down | JB1GVS23 | NC01682A(COG48AA)-A-AT-N3-C | GGP620.1-2BC11 | 08 | MODEL NOT FOUND | A6660 | Prober | Wafer Edge/Mode Not Found | 2022/08/28 09:35:59 | 2022/08/28 09:39:48 | 0.06 | OK | 10210 | 重抓MODE OK | NO | NO | NO | YES | YES | NO
EQPRepairJ22112337 | T6373-15 | CWC8J0G99-01 | Prober Down | JB6H1S04 | NC01690I(COG46BA)-A-AT-C | GGS958.1-2BC1 | 19 | 面板當 | 11550 | Prober | 面板Down (Monitor Down) | 2022/08/28 10:10:32 | 2022/08/28 10:44:21 | 0.56 | OK | 10210 | 1.Prober Reset ok 2.ReSetup Pass(22.4) | NO | NO | NO | YES | YES | NO
EQPRepairJ22112340 | T6373-38 | CWC8J0615-01.010 | Corr Fail | JB1GKS09 | PA6543B | J6AC7 | 07 | Corr Fail | A7086 | Corr | corr fail | 2022/08/28 10:22:36 | 2022/08/28 10:25:29 | 0.05 | OK | 13222 | 扣掉bin127 比對pass | NO | NO | NO | YES | YES | NO
EQPRepairJ22112348 | T6373-38 | CWC8J0841-01 | OverKill | JB1GKS09 | PA6543B | J6AK2 | 07 | too many O/S | A6873 | Over Kill | O/S誤宰 | 2022/08/28 11:25:50 | 2022/08/28 11:53:55 | 0.47 | OK | 15362 | 清調針PASS,P/M(132,7) | NO | NO | NO | YES | YES | NO
EQPRepairJ22112353 | T6373-43 | CEC7J0106-01 | OverKill | JB1UJS19 | PA5129LS_12U80H-LT | P08WT.1 | 24 | O/S-11550 | CPAUTO | Over Kill | 機台清針後Pass | 2022/08/28 11:48:41 | 2022/08/28 12:20:22 | 0.53 | OK | 20075 | 機台清針後Pass,RANGE 10 | NO | NO | NO | YES | YES | NO
EQPRepairJ22112378 | T6373-151 | CWC8J0955-01 | OverKill | JB252S02 | TD7850-A0S-AZ-CP | SC3MG-B11C | 04 | Continuous 2 wafers EDA shows "I" | A6919 | Over Kill | BIN20 ISSUE | 2022/08/28 14:29:02 | 2022/08/28 14:33:43 | 0.08 | OK | 20075 | BIN20 ISSUE,多次處理無效,續測 | NO | NO | NO | YES | YES | NO
EQPRepairJ22112392 | T6373-151 | CWC8J0955-01 | OverKill | JB252S02 | TD7850-A0S-AZ-CP | SC3MG-B11C | 04 | too many failed bin 20-A6919 | CPAUTO | Over Kill | Function誤宰 | 2022/08/28 15:41:25 | 2022/08/28 15:53:28 | 0.20 | OK | 10210 | Bin20請產線續測 | NO | NO | NO | YES | YES | NO
EQPRepairJ22112406 | T6373-70 | CWC8J0A76-01 | OverKill | JB1Q4S18 | PA6542C | J6F0H | 13 | o/s-11550 | CPAUTO | Over Kill | 機台清針後Pass | 2022/08/28 17:02:41 | 2022/08/28 17:27:18 | 0.41 | OK | 20075 | 機台清針後Pass | NO | NO | NO | YES | YES | NO
EQPRepairJ22112427 | T6373-152 |  | Other | JB252S01 | TD7850-A0S-AZ-CP | SC2QK-B2C |  | key board NG | A6919 | Tester | 鍵盤異常 | 2022/08/28 18:43:37 | 2022/08/28 18:48:42 | 0.08 | OK | 20075 | usb斷開,重插OK | NO | NO | NO | YES | YES | YES
EQPRepairJ22112430 | T6373-07 | CWC8J0A77-01 | OverKill | JB1TAS26 | PA5208KF-LT | J6CC9 | 20 | TOO MANY O/S-A6873 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/28 18:51:52 | 2022/08/28 19:01:22 | 0.16 | OK | 15362 | 清調針pass,p/m(46,6) | NO | NO | NO | YES | YES | NO
EQPRepairJ22112436 | T6373-128 | CWC7J0336-01.001 | Tester Down | TB1R8S83 | PA5137E1-LT | P07HG.1 | 01 | TESTER NG ( LONG TIME ) | A6919 | Tester | Other | 2022/08/28 19:10:38 | 2022/08/28 19:40:20 | 0.50 | OK | 20182 | 重新連線OK | NO | NO | NO | NO | YES | NO
EQPRepairJ22112461 | T6373-70 | CWC8J0A76-01 | OverKill | JB1Q4S18 | PA6542C | J6F0H | 15 | too many o/s-10851 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/28 20:22:16 | 2022/08/28 21:08:44 | 0.77 | OK | 20182 | 1.更換備卡#S16，上機PASS 2.RG-10 PM(41,0) | NO | NO | NO | YES | YES | NO
EQPRepairJ22112467 | T6373-21 | CWC8J0F99-01.002 | Prober Down | JB1GKS08 | PA6543B | J6AC7 | 13 | Model Not Found-A6318 | CPAUTO | Prober | Wafer Edge/Mode Not Found | 2022/08/28 20:36:39 | 2022/08/28 20:48:32 | 0.20 | OK | 17091 | remote----手動抓取ok check die ok pm(20，6) | NO | NO | NO | YES | YES | NO
EQPRepairJ22112468 | T6373-16 |  | Prober Down | VB405S07 | PA5208KF-LT | J6FGK |  | AlarmCode:1995 AlarmMessage:Faied to upload log file | CPAUTO | Prober | Other | 2022/08/28 20:39:40 | 2022/08/28 20:43:34 | 0.07 | OK | 17185 | 確認無異常續測 | NO | NO | NO | YES | YES | NO
EQPRepairJ22112472 | T6373-152 | CWC8J0455-01.004 | OverKill | JB252S01 | TD7850-A0S-AZ-CP | SC1GS-B1C | 09 | TOO MANY O/S..NO PARTICLE-A6516 | CPAUTO | Over Kill | O/D不足 | 2022/08/28 20:59:59 | 2022/08/28 21:26:49 | 0.45 | OK | 17185 | 重對 R:13um (176.6)(177.6) | NO | NO | NO | YES | YES | NO
EQPRepairJ22112482 | T6373-151 | CWC8J0955-01 | OverKill | JB252S02 | TD7850-A0S-AZ-CP | SC3MG-B11C | 06 | many o/s.-A7048 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/28 21:26:46 | 2022/08/28 21:45:07 | 0.31 | OK | 17185 | 清、調針Pass (52.7)(53.7) | NO | NO | NO | YES | YES | NO
EQPRepairJ22112494 | T6373-128 | CWC7J0336-01.001 | OverKill | TB1R8S83 | PA5137E1-LT | P07HF.1 | 12 | EDA SHOW 2H NO CORR WAFER | A6512 | Over Kill | 2H | 2022/08/28 22:50:45 | 2022/08/28 22:55:36 | 0.08 | OK | 17185 | 2H [HW_BIN0<95.5%](Bin 1 < 814EA) 804(Hold Wafer) 不同機卡組相同續測 | NO | NO | NO | YES | YES | NO
EQPRepairJ22112503 | T6373-16 | CEC8J0085-01 | OverKill | VB405S07 | PA5208K-LT | QTC9L | 08 | TOO MANY O.S-A6938 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/28 23:40:11 | 2022/08/28 23:57:32 | 0.29 | OK | 10110 | CLEAN WAFER清針PASS | NO | NO | NO | YES | YES | NO
EQPRepairJ22112517 | T6373-151 | CWC8J0955-01 | OverKill | JB252S02 | TD7850-A0S-AZ-CP | SC3MG-B11C | 08 | too many o/s...no psrticle-A6516 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/29 01:02:27 | 2022/08/29 02:00:50 | 0.97 | OK | 17185 | 下卡磨針 (147.7)(148.7) | NO | NO | NO | YES | YES | NO
EQPRepairJ22112535 | T6373-151 | CWC8J0955-01 | OverKill | JB252S02 | TD7850-A0S-AZ-CP | SC3MG-B11C | 08 | too many o/s..no particle-A6516 | CPAUTO | Over Kill | 針位調整 | 2022/08/29 03:11:25 | 2022/08/29 03:57:28 | 0.77 | OK | 17185 | 調整針位 (154.0)(155.0) | NO | NO | NO | YES | YES | NO
EQPRepairJ22112537 | T6373-16 |  | Prober Down | VB405S07 | PA5208KF2 | J66H4 |  | CANNOT MEASURE INDEX SIZE-A6512 | CPAUTO | Prober | Wafer Edge/Mode Not Found | 2022/08/29 03:14:19 | 2022/08/29 03:23:29 | 0.15 | OK | 17185 | 重抓 &check die ok | NO | NO | NO | YES | YES | NO
EQPRepairJ22112538 | T6373-43 | CWC8J0620-01.002 | Tester Down | JB1UJS19 | PA5129LS_12U80H-LT | P07CK.1 | 03 | AlarmCode:669 AlarmMessage:GP-IB COMMUNICATION TIMEOUT ERROR !! | CPAUTO | Tester | 離線(OCT/GPIB error) | 2022/08/29 03:20:46 | 2022/08/29 03:28:36 | 0.13 | OK | 09108 | Remote---- 1.無異常訊息 Soaktime後自動起測，無異常，續測 | NO | NO | NO | YES | YES | NO
EQPRepairJ22112541 | T6373-141 | CWC8J0F65-01.001 | OverKill | JB1HKS07 | NA01687G(COG48BA)-A-AT-N2-C | GGT209.1-BC21 | 21 | FAIL BIN TOO MANY-A6318 | CPAUTO | Over Kill | Function誤宰 | 2022/08/29 03:47:45 | 2022/08/29 04:28:28 | 0.68 | OK | 17091 | remote----bin29 clean unit pass+auto cal | NO | NO | NO | YES | YES | NO
EQPRepairJ22112543 | T6373-70 | CWC8J0A76-01 | OverKill | JB1Q4S16 | PA6542C | J6F0H | 21 | TOO MANY OS-A6512 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/29 04:04:03 | 2022/08/29 04:28:45 | 0.41 | OK | 20182 | 1.清針，重對針PASS 2.RG-10 PM(12.4) | NO | NO | NO | YES | YES | NO
EQPRepairJ22112551 | T6373-16 | CMC7J0154-01 | OverKill | VB405S07 | PA5208KF-LT | J66H3 | 18 | continous 2 wafers eda show H (o/s and fail bin) Too many o/s | A6898 | Over Kill | 2H | 2022/08/29 04:36:07 | 2022/08/29 04:52:20 | 0.27 | OK | 20182 | 2H，針髒清針 續測觀察 | NO | NO | NO | YES | YES | NO
EQPRepairJ22112555 | T6373-151 | CWC8J0955-01 | OverKill | JB252S02 | TD7850-A0S-AZ-CP | SC3MG-B11C | 09 | continous 2 wafers eda show I | A6898 | Over Kill | 2I | 2022/08/29 04:47:50 | 2022/08/29 05:00:02 | 0.20 | OK | 20182 | 日班交接，Bin20處理無效 先續測 | NO | NO | YES | YES | YES | NO
EQPRepairJ22112557 | T6373-07 | CWC8J0A77-01 | OverKill | JB1TAS26 | PA5208KF-LT | J6CC9 | 24 | o/s too many-A6318 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/29 04:51:44 | 2022/08/29 05:16:17 | 0.41 | OK | 20182 | 1.清針PASS PM(17,5) 起測一段穩定 | NO | NO | NO | YES | YES | NO
EQPRepairJ22112562 | T6373-151 | CWC8J0955-01 | OverKill | JB252S02 | TD7850-A0S-AZ-CP | SC3MG-B11C | 09 | too many o/s-A6938 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/29 05:09:59 | 2022/08/29 05:58:20 | 0.81 | OK | 20182 | 1.清針，調針PASS 2.RG-10 PM(192.5) 3.BIN20 續測 | NO | NO | YES | YES | YES | NO
EQPRepairJ22112569 | T6373-38 | CWC8J0841-01.006 | OverKill | JB1GKS09 | PA6543B | J6AK2 | 12 | too many o/s. pc has particle-A6938 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/29 05:30:46 | 2022/08/29 05:57:01 | 0.44 | OK | 17091 | remote---清針後二分法調針pass 測試一段穩定還機 | NO | NO | NO | YES | YES | NO
EQPRepairJ22112585 | T6373-33 | CWC8J0C41-01 | OverKill | JB1TAS19 | PA5208KF-LT | J6CHG | 17 | TOO MAY O/S-A7023 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/29 06:22:39 | 2022/08/29 06:50:13 | 0.46 | OK | 09108 | 1.o/s 針位差一個BUMP y:46 2.重新對針 setup pass range:16um (45.6) | NO | NO | NO | YES | YES | NO
EQPRepairJ22112591 | T6373-151 | CWC8J0955-01 | OverKill | JB252S02 | TD7850-A0S-AZ-CP | SC3MG-B11C | 09 | too many o/s-A6938 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/29 07:03:15 | 2022/08/29 07:24:14 | 0.35 | OK | 20075 | 清調針PASS,RANGE 10 | NO | NO | NO | YES | YES | NO
EQPRepairJ22112611 | T6373-16 | CWC8J0G90-01.002 | OverKill | VB405S07 | PA5208KF-LT | J6FGK | 03 | TOO MANY O/S-A6317 | CPAUTO | Over Kill | 配件殘金 | 2022/08/29 08:44:34 | 2022/08/29 09:14:44 | 0.50 | OK | 20075 | 換卡#S09,range 10 | NO | NO | NO | YES | YES | NO
EQPRepairJ22112612 | T6373-151 | CWC8J0955-01 | OverKill | JB252S02 | TD7850-A0S-AZ-CP | SC3MG-B11C | 10 | TOO MANY O/S-A6317 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/29 08:50:46 | 2022/08/29 09:12:43 | 0.37 | OK | 10210 | 上調4um Pass | NO | NO | NO | YES | YES | NO
EQPRepairJ22112622 | T6373-42 | CWC8J0G16-01.001 | OverKill | JB1HKS03 | NA01687G(COG48BA)-A-AT-N2-C | GGU455.1-BC2 | 04 | TOO MANY FAIL BIN29 | A6660 | Over Kill | Function誤宰 | 2022/08/29 09:29:10 | 2022/08/29 09:46:14 | 0.28 | OK | 20075 | bin29,清調針pass,range 8,mark 182,2 | NO | NO | NO | YES | YES | NO
EQPRepairJ22112630 | T6373-33 | CWC8J0C41-01 | OverKill | JB1TAS19 | PA5208KF-LT | J6CHG | 19 | O/S-14095 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/29 10:11:44 | 2022/08/29 10:46:37 | 0.58 | OK | 20230 | REMOTE 1.O/S 2.清調針,RG=15,PM(5.2) 3.PASSS | NO | NO | NO | YES | YES | NO
EQPRepairJ22112634 | T6373-42 | CWC8J0G16-01.001 | OverKill | JB1HKS03 | NA01687G(COG48BA)-A-AT-N2-C | GGU455.1-BC2 | 04 | # 4 too many failbin 29, pm checked ,od:10-A6143 | CPAUTO | Over Kill | Function誤宰 | 2022/08/29 10:20:33 | 2022/08/29 10:30:10 | 0.16 | OK | 10210 | 前清OK | NO | NO | NO | YES | YES | NO
EQPRepairJ22112637 | T6373-07 | CWC8J0C06-01 | OverKill | JB1TAS26 | PA5208KF-LT | J6CHC | 01 | TOO MANY O/S-A6873 | CPAUTO | Over Kill | O/S | 2022/08/29 10:48:34 | 2022/08/29 11:18:17 | 0.50 | OK | 13222 | O/S 機台清針OK | NO | NO | NO | YES | YES | NO
EQPRepairJ22112652 | T6373-42 | CWC8J0G16-01.001 | OverKill | JB1HKS03 | NA01687G(COG48BA)-A-AT-N2-C | GGU455.1-BC2 | 05 | 已依誤宰處理流程處置無效 (Bin29過多) | 14095 | Over Kill | Function誤宰 | 2022/08/29 13:26:42 | 2022/08/29 15:02:08 | 1.59 | OK | 16158 | Bin29，下卡磨針pass，Y15。(15.3) | NO | NO | NO | YES | YES | NO
EQPRepairJ22112662 | T6373-07 | CWC8J0C06-01 | OverKill | JB1TAS26 | PA5208KF-LT | J6CHC | 02 | TOO MANY O/S-A6873 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/29 14:44:46 | 2022/08/29 15:03:22 | 0.31 | OK | 10110 | 清調針pass,p/m(46,6) | NO | NO | NO | YES | YES | NO
EQPRepairJ22112671 | T6373-33 | CWC8J0C41-01 | OverKill | JB1TAS19 | PA5208KF-LT | J6CHG | 21 | CONTINOUS 2 WAFERS EDA SHOW R O/S-A6317 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/29 15:04:58 | 2022/08/29 15:49:10 | 0.74 | OK | 20230 | REMOTE 1.O/S 2.更換備卡S31,RG=13,PM(46.6) 3.PASS | NO | NO | NO | YES | YES | NO
EQPRepairJ22112675 | T6373-147 | CWC8J0227-01.001 | OverKill | JB1UJS18 | PA5129LS_12U80H-LT | P07CJ.1 | 24 | #24 O/S | 20427 | Over Kill | 機台清針後Pass | 2022/08/29 15:56:20 | 2022/08/29 16:18:40 | 0.37 | OK | 20075 | 機台清針後Pass | NO | NO | NO | YES | YES | NO
EQPRepairJ22112688 | T6373-07 | CWC8J0C06-01 | OverKill | JB1TAS26 | PA5208KF-LT | J6CHC | 03 | too many o/s overkill pass tehn o/s again-A6785 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/29 17:54:53 | 2022/08/29 18:13:47 | 0.32 | OK | 15362 | 清調針,p/\m(164,3) | NO | NO | NO | YES | YES | NO
EQPRepairJ22112689 | T6373-111 | CWC7J0J58-01 | Other | JB1Q4S14 | PA6542C | J683L | 10 | PLS HELP TO SET 3D CLEAN MODE MEMO-A6143 | CPAUTO | Other | 3D CLEAN | 2022/08/29 17:57:16 | 2022/08/29 18:11:12 | 0.23 | OK | 20230 | 3D CLEAN SET OK | NO | NO | NO | YES | YES | NO
EQPRepairJ22112690 | T6373-42 | CWC8J0G16-01.001 | OverKill | JB1HKS03 | NA01687G(COG48BA)-A-AT-N2-C | GGU455.1-BC2 | 07 | TOO MANY FAILBINS-A6317 | CPAUTO | Over Kill | BIN29 | 2022/08/29 18:00:27 | 2022/08/29 18:26:15 | 0.43 | OK | 13222 | BIN29 機台清針OK 點測穩定 續測 待觀察 (145.0) | NO | NO | NO | YES | YES | NO
EQPRepairJ22112698 | T6373-35 | CWC8J0F66-01.002 | OverKill | JB1QNS06 | NA01687G(COG48BA)-A-AT-X1-C | GGT528.1-BC22 | 06 | too many bin 6 | A6873 | Over Kill | Function誤宰 | 2022/08/29 18:55:19 | 2022/08/29 19:59:22 | 1.07 | OK | 15492 | 1.調整後方風管，點測PASS 2.續測觀察 | NO | NO | NO | YES | YES | NO
EQPRepairJ22112699 | T6373-01 | CWC7J0L22-01 | Tester Down | JB1TAS30 | PA5208K-LT | QTAAS | 22 | please check always need to click continue-A6785 | CPAUTO | Probe Card | Other | 2022/08/29 18:57:15 | 2022/08/29 20:17:37 | 1.34 | OK | 15492 | 1.實驗機台配件S30，產線換批後上貨Load 程式會跑TDR Fail,重MNT,更換POGO PIN 無效. 2.更換在庫配件S20,上機驗證PASS，與配件室確認此在庫配件已有登入數值 3.上機SETUP PASS，RANGE:12 | NO | NO | YES | YES | YES | NO
EQPRepairJ22112723 | T6373-111 | CWC7J0J58-01 | Tester Down | JB1Q4S14 | PA6542C | J683L | 10 | AlarmCode:669 AlarmMessage:GP-IB COMMUNICATION TIMEOUT ERROR !! | CPAUTO | Tester | 離線(OCT/GPIB error) | 2022/08/29 21:08:52 | 2022/08/29 21:16:50 | 0.13 | OK | 09108 | Remote---- 1.無異常rte訊息， 消除Alarm訊息 ，機台自動執行Soaktime後起測，無異常，續測 | NO | NO | NO | YES | YES | NO
EQPRepairJ22112726 | T6373-33 | CWC8J0C41-01 | OverKill | JB1TAS31 | PA5208KF-LT | J6CHG | 22 | #22 too many o/s | A6441 | Over Kill | O/S誤宰 | 2022/08/29 21:12:57 | 2022/08/29 21:32:07 | 0.32 | OK | 09108 | 1.o/s 清針 調針Pass range:10u (46.6) | NO | NO | NO | YES | YES | NO
EQPRepairJ22112727 | T6373-34 |  | Prober Down | TB1R8S85 | PA5223Q_12U60H-LT | P0CLN.1 |  | Model not found | A6896 | Prober | Wafer Edge/Mode Not Found | 2022/08/29 21:16:44 | 2022/08/29 21:33:45 | 0.28 | OK | 15492 | REMOTE--- 1.重抓 module ok 2.CHECK DIE (31，1)OK | YES | NO | NO | YES | YES | NO
EQPRepairJ22112728 | T6373-34 | CWC7J0L13-01 | Prober Down | TB1R8S85 | PA5223Q_12U60H-LT | P0CLN.1 | 01 | MODEL NOT FOUND | A7025 | Prober | Wafer Edge/Mode Not Found | 2022/08/29 21:17:47 | 2022/08/29 21:36:39 | 0.31 | OK | 15492 | REMOTE--- 1.重抓 MODULE OK 2.CHECK DIE (31,1)OK | NO | NO | NO | YES | YES | NO
EQPRepairJ22112732 | T6373-43 | CWC5J0C97-01 | OverKill | JB1UJS19 | PA5129LS_12U80H-LT | P01PT.1 | 01 | Already follow Overkilll handling procedures but still not effective too many o/s | A6810 | Over Kill | O/S | 2022/08/29 21:38:03 | 2022/08/29 21:46:33 | 0.14 | OK | 16157 | Remote處理--------O/S 機台清針pass | NO | NO | NO | YES | YES | NO
EQPRepairJ22112737 | T6373-07 | CWC8J0C06-01 | OverKill | JB1TAS26 | PA5208KF-LT | J6CHC | 05 | #5 Already follow theoverkilling procedure but still OS | A6896 | Over Kill | O/S誤宰 | 2022/08/29 22:23:09 | 2022/08/29 22:40:45 | 0.29 | OK | 17091 | 清針後確認針壓OK 二分法調針R:12 起測觀察一段穩定還機 | NO | NO | NO | YES | YES | NO
EQPRepairJ22112745 | T6373-35 | CWC8J0F66-01.002 | OverKill | JB1QNS06 | NA01687G(COG48BA)-A-AT-X1-C | GGT528.1-BC22 | 08 | #6.7 eda show h no corr wafer | A6512 | Over Kill | 2h | 2022/08/29 22:47:45 | 2022/08/29 23:12:00 | 0.40 | OK | 17091 | 2h bin6 先前修過bin6為風管不夠熱，重新調整風管位置後靜置10分鐘， 點測 10ea皆穩定pass，還機觀察一片 | NO | NO | NO | YES | YES | NO
EQPRepairJ22112756 | T6373-35 | CWC8J0F66-01.002 | OverKill | JB1QNS06 | NA01687G(COG48BA)-A-AT-X1-C | GGT528.1-BC22 | 08 | too many fail bin 6 od12 no particle-A6943 | CPAUTO | Over Kill | Function誤宰 | 2022/08/30 00:08:01 | 2022/08/30 00:21:16 | 0.22 | OK | 15492 | 1.BIN6 點測不穩，重新固定機台後方風管，點測穩定 PASS 2.續測觀察 | NO | NO | NO | YES | YES | NO
EQPRepairJ22112757 | T6373-42 | CWC8J0G16-01.001 | OverKill | JB1HKS03 | NA01687G(COG48BA)-A-AT-N2-C | GGU455.1-BC2 | 10 | TOO MANY FAIL BIN 29-A6673 | CPAUTO | Over Kill | Function誤宰 | 2022/08/30 00:14:47 | 2022/08/30 01:02:08 | 0.79 | OK | 16157 | BIN29 機台清針，調整針位PASS | NO | NO | NO | YES | YES | NO
EQPRepairJ22112761 | T6373-35 | CWC8J0F66-01.002 | OverKill | JB1QNS06 | NA01687G(COG48BA)-A-AT-X1-C | GGT528.1-BC22 | 08 | #8Continous bin 6 | A6896 | Over Kill | Function誤宰 | 2022/08/30 00:28:51 | 2022/08/30 01:07:26 | 0.64 | OK | 15492 | 1.BIN6，熱風較能使機台穩定PASS，故再次調整後方風管位置 | NO | NO | NO | YES | YES | NO
EQPRepairJ22112776 | T6373-141 | CWC8J0G95-01 | OverKill | JB1HKS07 | NA01687G(COG48BA)-A-AT-N2-C | GGU012.1-B1C1 | 24 | #24 Continous bin 29 | A6896 | Over Kill | O/S誤宰 | 2022/08/30 01:35:40 | 2022/08/30 01:50:35 | 0.25 | OK | 17091 | remote---clean unit pass | NO | NO | NO | YES | YES | NO
EQPRepairJ22112783 | T6373-07 | CWC8J0C06-01 | OverKill | JB1TAS26 | PA5208KF-LT | J6CHC | 06 | TOO MANY O/S | A6810 | Over Kill | O/S誤宰 | 2022/08/30 02:01:20 | 2022/08/30 02:42:57 | 0.69 | OK | 17091 | 下卡磨針上機PASS R:0~-16 | NO | NO | NO | YES | YES | NO
EQPRepairJ22112785 | T6373-21 | CWC8J0F99-01 | OverKill | JB1GKS08 | PA6543B | J6AK4 | 05 | TOO MANY O/S | A6941 | Over Kill | O/S誤宰 | 2022/08/30 02:21:12 | 2022/08/30 03:04:22 | 0.72 | OK | 17091 | 清針後二分法調針pass | NO | NO | NO | YES | YES | NO
EQPRepairJ22112788 | T6373-33 | CWC8J0C41-01 | OverKill | JB1TAS31 | PA5208KF-LT | J6CHG | 24 | REPROBING TOO MANY O/S.-A7048 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/30 02:32:17 | 2022/08/30 02:54:55 | 0.38 | OK | 10897 | 清針，調針，測試一段穩定 | NO | NO | NO | YES | YES | NO
EQPRepairJ22112796 | T6373-01 | CEC8J0085-01 | Tester Down | JB1TAS20 | PA5208K-LT | QTC9L | 02 | #2 GPIB | A6896 | Tester | 離線(OCT/GPIB error) | 2022/08/30 03:09:59 | 2022/08/30 03:16:13 | 0.10 | OK | 17091 | 重新連線OK | NO | NO | NO | YES | YES | NO
EQPRepairJ22112815 | T6373-39 | CWC8J0369-01 | Prober Down | JB1UJS13 | PA5129LS_12U80H-LT | P0GRF.1 | 23 | AlarmCode:1995 AlarmMessage:Faied to upload log file | CPAUTO | Prober | Faied to upload log file | 2022/08/30 04:49:42 | 2022/08/30 04:58:00 | 0.14 | OK | 20182 | 重新起測 無異常 | NO | NO | NO | YES | YES | NO
EQPRepairJ22112824 | T6373-15 |  | Setup Fail | JB6H1S04 | NC01690I(COG46BA)-A-AT-C | GGS959.1-B |  | too large needle --wait engr repair-A6118 | CPAUTO | Setup Fail | large needle | 2022/08/30 05:27:21 | 2022/08/30 05:40:49 | 0.22 | OK | 20182 | 重新對針OK | NO | NO | NO | YES | YES | NO
EQPRepairJ22112831 | T6373-35 | CWC8J0F66-01.002 | OverKill | JB1QNS06 | NA01687G(COG48BA)-A-AT-X1-C | GGT528.1-BC22 | 14 | many o/s-A7048 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/30 06:03:03 | 2022/08/30 06:30:36 | 0.46 | OK | 15492 | 1.針位貼上，二分法調針位(0，-2)、RANGE:18 2.P/M(146，4)OK | NO | NO | NO | YES | YES | NO
EQPRepairJ22112839 | T6373-42 | CWC8J0G16-01.001 | OverKill | JB1HKS03 | NA01687G(COG48BA)-A-AT-N2-C | GGU455.1-BC2 | 13 | too many fail bin 29 od15 no particle-A6943 | CPAUTO | Over Kill | Function誤宰 | 2022/08/30 06:53:23 | 2022/08/30 08:04:58 | 1.19 | OK | 17680 | POLISH,POSITION(7,2),MARK(36,0) | NO | NO | NO | YES | YES | NO
EQPRepairJ22112840 | T6373-07 | CWC8J0C06-01 | OverKill | JB1TAS26 | PA5208KF-LT | J6CHC | 08 | TOO MANY O/S-A6810 | CPAUTO | Over Kill | O/D不足 | 2022/08/30 06:53:33 | 2022/08/30 07:26:30 | 0.55 | OK | 13749 | 1.SLOT #08 測試中o/s過多，測試OD 維持5um. 2.輪工確認 針況 針位(119.0)無異常，加針壓od:5->10um 後pass 穩定 續觀察. | NO | NO | NO | YES | YES | NO
EQPRepairJ22112857 | T6373-143 | CWC4J0A49-01.002 | Prober Down | TB1R7S31 | PA5258D-LT | PXSN6.1 | 17 | AlarmCode:8732 AlarmMessage:It failed in lock/unlock of the manual ring lever. | CPAUTO | Prober | Sensor Error | 2022/08/30 08:05:40 | 2022/08/30 08:52:20 | 0.78 | OK | 16631 | 1.unlock延遲，更換氣壓缸ok，setup pass 156.4 | NO | NO | NO | YES | YES | NO
EQPRepairJ22112867 | T6373-141 | CWC8J0F66-01 | OverKill | JB1HKS07 | NA01687G(COG48BA)-A-AT-X1-C | GGT528.1-BC1 | 09 | #02,#03 EDA SHOW 2H-20021 | CPAUTO | Over Kill | 2H | 2022/08/30 08:35:43 | 2022/08/30 09:24:40 | 0.82 | OK | 13749 | Remote--- 1.Slot #02-#03 因BIN29 Function Fail 過多Lot Yield Show 2H. 2.確認EDA 同LOT 拆批至不同機卡ND3-42 測試與回測 皆有相同情形. 3.且不同LOT:GGS268.1-S 在不同機卡測試 也曾有部分Wafer 發生BIN29 Fail 過多穩定. 4.確認針況明顯黑點+金絲沾黏 故先行更換備卡#S07->#S09 PASS PM(36.0) Slot #09. 5.請TA Run corr Check. | NO | NO | NO | YES | YES | NO
EQPRepairJ22112872 | T6373-33 | CWC6J0F73-01.001 | OverKill | TB4OSS05 | PA1521D-50D | U41NK.00 | 21 | bin 31-A6962 | CPAUTO | Over Kill | Function誤宰 | 2022/08/30 09:12:27 | 2022/08/30 10:16:57 | 1.08 | OK | 19771 | 1.易跳bin31,清調針無效 2.更換備卡相同 3.工程品,有指定配件,再次更換回原卡續測 | NO | NO | NO | YES | YES | NO
EQPRepairJ22112878 | T6373-43 | CWC5J0L13-01.002 | Tester Down | JB1UJS19 | PA5129LS_12U80H-LT | P01PY.1 | 07 | AlarmCode:669 AlarmMessage:GP-IB COMMUNICATION TIMEOUT ERROR !! | CPAUTO | Product Issue | RTE | 2022/08/30 10:21:17 | 2022/08/30 10:25:15 | 0.07 | OK | 08774 | 無異常訊息，重新連線觀察 | NO | NO | NO | YES | YES | NO
EQPRepairJ22112881 | T6373-33 | CWC6J0F73-01.001 | OverKill | TB4OSS05 | PA1521D-50D | U41NK.00 | 21 | FAIL BIN31-17436 | CPAUTO | Over Kill | Function誤宰 | 2022/08/30 10:43:10 | 2022/08/30 11:05:33 | 0.37 | OK | 15897 | #21(46.14)重測片.此批有Bin31 Function issue疑似產品問題.續測待換批驗證 | NO | NO | NO | YES | YES | NO
EQPRepairJ22112892 | T6373-141 |  | Corr Fail | JB1HKS09 | NA01687G(COG48BA)-A-AT-X1-C | GGT528.1-BC1 |  | Corr Fail | A6916 | Corr | Corr Fail | 2022/08/30 11:53:28 | 2022/08/30 12:02:46 | 0.16 | OK | 13749 | Corr Fail----系統比對Corr 良率比 Corr Wafer 良率好 | NO | NO | NO | YES | YES | NO
EQPRepairJ22112902 | T6373-35 | CWC8J0F66-01.002 | OverKill | JB1QNS06 | NA01687G(COG48BA)-A-AT-X1-C | GGT528.1-BC22 | 19 | #17,#18 EDA SHOW 2H No corr-200021 | CPAUTO | Over Kill | 2H | 2022/08/30 12:48:15 | 2022/08/30 13:11:33 | 0.39 | OK | 16631 | 1.調整針壓，機台清針，點測BIN6 PASS穩定，續測觀察 | NO | NO | NO | YES | YES | NO
EQPRepairJ22112905 | T6373-47 | CWC8J0F65-01 | OverKill | JB1HKS10 | NA01687G(COG48BA)-A-AT-N2-C | GGT209.1-BC1 | 20 | O/S-17436 | CPAUTO | Over Kill | O/D不足 | 2022/08/30 12:58:13 | 2022/08/30 13:07:52 | 0.16 | OK | 20231 | 1.OD10-->13 pass，PM(145,0) | NO | NO | NO | YES | YES | NO
EQPRepairJ22112908 | T6373-128 | CWC7J0M32-01.003 | OverKill | JB1UJS16 | PA5129LS_12U80H-LT | P060Q.1 | 10 | TOO MANY O/S-A6671 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/30 13:16:58 | 2022/08/30 13:32:12 | 0.25 | OK | 08774 | O/S誤宰 1.執行soaktime 重新對針 機台清針 35.5 | NO | NO | NO | YES | YES | NO
EQPRepairJ22112912 | T6373-143 |  | Other | TB1R7S31 | PA5258D-LT | PXTP3.1 |  | CANNOT OPEN CBVP MASS PRODUCTION-A6513 | CPAUTO | Product Issue | 自動化異常 | 2022/08/30 13:27:22 | 2022/08/30 14:25:36 | 0.97 | OK | 16631 | 1.自動化異常，指令無效，tester shutdown無效，待常日協助確認。 | NO | NO | NO | YES | YES | NO
EQPRepairJ22112915 | T6373-42 | CWC8J0G16-01.001 | OverKill | JB1HKS03 | NA01687G(COG48BA)-A-AT-N2-C | GGU455.1-BC2 | 16 | bin 29 | A6921 | Over Kill | Function誤宰 | 2022/08/30 13:40:10 | 2022/08/30 14:15:43 | 0.59 | OK | 20231 | 1.原配件#03殘金下卡磨針歸庫，更換備卡#07上機pass，溫針重對針，rg:20，PM(144,0) | NO | NO | NO | YES | YES | NO
EQPRepairJ22112916 | T6373-35 | CWC8J0F66-01.002 | OverKill | JB1QNS06 | NA01687G(COG48BA)-A-AT-X1-C | GGT528.1-BC22 | 19 | too many BIN6 | A6668 | Over Kill | Function誤宰 | 2022/08/30 13:41:34 | 2022/08/30 14:45:48 | 1.07 | OK | 16631 | 1.換卡#04 170.3 | NO | NO | NO | YES | YES | NO
EQPRepairJ22112922 | T6373-103 |  | Other | JB1CES10 | NA01248F(COG4)-A-AT-X2-C | NSWFY-B2 |  | PLEASE SAVE 3D CLEAN MODE-a6513 | CPAUTO | Prober | set 3d clean | 2022/08/30 14:09:43 | 2022/08/30 14:15:22 | 0.09 | OK | 07910 | set 3d clean ok | NO | NO | NO | YES | YES | NO
EQPRepairJ22112926 | T6373-143 |  | Product Issue | TB1R7S31 | PA5258D-LT | PXTP3.1 |  | Product Issue---- 1.自動化異常，指令無效，tester shutdown無效，待常日協助確認。 | 16631 | Product Issue | 自動化異常 | 2022/08/30 14:26:13 | 2022/08/30 15:07:29 | 0.69 | OK | 16631 | 1.常日協助處理ok | NO | NO | NO | YES | YES | NO
EQPRepairJ22112931 | T6373-15 | CWC8J0124-01.001 | Corr Fail | JB1HKS08 | NA01687G(COG48BA)-A-AT-N2-C | GGS268.1-B2 | 14 | ok-A6668 | CPAUTO | Corr | Other | 2022/08/30 14:49:58 | 2022/08/30 14:52:06 | 0.04 | OK | 20231 | 1.手動比對pass | NO | NO | NO | YES | YES | NO
EQPRepairJ22112935 | T6373-15 |  | Tester Down | JB6H1S04 | NC01690I(COG46BA)-A-AT-C | GGW030.1-BC1211 |  | cannot open mass production on cbvp-A6668 | CPAUTO | Product Issue | 自動化異常 | 2022/08/30 15:07:18 | 2022/08/30 15:36:16 | 0.48 | OK | 08774 | 開啟ok | NO | NO | NO | YES | YES | NO
EQPRepairJ22112939 | T6373-47 | CWC8J0G48-01.002 | OverKill | JB1HKS10 | NA01687F(COG48BA)-A-AT-C | GGS941.1-22BC12 | 12 | fail bin become o/s-A6894 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/30 15:30:56 | 2022/08/30 15:56:48 | 0.43 | OK | 08774 | Remote 處理---- 1.執行soaktime 重新對針 機台清針 | NO | NO | NO | YES | YES | NO
EQPRepairJ22112942 | T6373-33 | CWC8J0815-01.001 | Other | JB1TAS13 | PA5208K-LT | QT7YW | 14 | WAFER ID CLEAR BUT CANNOT AUTO READ | A6513 | Prober | wafer ID 無法自動辨識 | 2022/08/30 15:47:23 | 2022/08/30 15:55:25 | 0.13 | OK | 19771 | 手動調整後可自動辨識 | NO | NO | NO | YES | YES | NO
EQPRepairJ22112945 | T6373-07 | CWC8J0C06-01 | OverKill | JB1TAS26 | PA5208KF-LT | J6CHC | 11 | #09,#10 EDA SHOW 2R(o/s)-20021 | CPAUTO | Over Kill | 2R | 2022/08/30 15:54:36 | 2022/08/30 18:06:48 | 2.20 | OK | 13749 | 1.EDA SHOW 2R O/S 針尖黑點 更換備卡#s26->#s23. 2.換卡後上機固定lcd pin:462 fail，抬頭更換pogo pin無效. 3.更換第二張配件#S22 PASS 預溫 重對針OK Slot #11 (46.6) | NO | NO | NO | YES | YES | NO
EQPRepairJ22112947 | T6373-13 | CWC8J0H99-01 | Setup Fail | JB8WXS08 | ND01733G(COF94AA)-A-C | H2UN66.1-1B | 01 | setup o/s 調針清針重moumt無效 下卡磨針無備卡 | 05671 | Probe Card | 針位異常 | 2022/08/30 16:06:33 | 2022/08/30 16:34:33 | 0.47 | OK | 11385 | P/C down----改機s/u open不過，配件下卡磨針，配件室通知針位不佳，送修 | NO | NO | NO | YES | YES | NO
EQPRepairJ22112953 | T6373-13 |  | Probe Card Down | JB1PNS11 | NB01761A(COF94AA)-A-C | H2NF30.1-B1C12 |  | P/C down----改機s/u open不過，配件下卡磨針，配件室通知針位不佳，送修 | 11385 | Probe Card | 針位異常 | 2022/08/30 16:35:02 | 2022/08/30 22:22:24 | 5.79 | OK | 13597 | P/C down----改機s/u open不過，配件下卡磨針，配件室通知針位不佳，送修 1.送修回來上機pass，range 10 um 2.請產線重新改機run corr | NO | NO | NO | YES | YES | NO
EQPRepairJ22112965 | T6373-143 | CWC7J0523-01 | Setup Fail | TB1HVS21 | PA1333R-25E | RDKZJ9000 | 04 | setup fail due to Bin13(Bin Leakage), change P/C TB1VS21 TB1HVS20-A6513 | CPAUTO | Setup Fail | Function Fail | 2022/08/30 17:56:47 | 2022/08/30 18:24:05 | 0.45 | OK | 16631 | 1.重新對針bin24 function，run card上標示bin24.25.27問題客戶指示先行續測，交回setup 04 | NO | NO | NO | YES | YES | NO
EQPRepairJ22112968 | T6373-34 | CWC8J0957-01 | OverKill | TB1YLS39 | PA6485E(CP1-CP3) | P0CSG.1 | 03 | O/S | A6962 | Over Kill | O/S誤宰 | 2022/08/30 18:00:34 | 2022/08/30 18:27:43 | 0.45 | OK | 07910 | 重新對針 pass p/m(219.1)#3 RY13 | NO | NO | NO | YES | YES | NO
EQPRepairJ22112969 | T6373-01 | CEC8J0085-01 | Other | JB1TAS20 | PA5208K-LT | QTC9L | 08 | AE CHANGE P/C-20022 | CPAUTO | Other | Other | 2022/08/30 18:12:03 | 2022/08/30 18:37:24 | 0.42 | OK | 13749 | 1.設備啟家指示協助更換配件測試JB1TAS20->JB1TAS25 OK. 2.上機預溫重對針、點測PASS溫針退針壓 PM(45.0) X=-5 Y=4 Slot #08. 3.依指示下修機台清針參數 3D Clean WA:OD20/2000ea/10下 CL:OD20/300ea/10下 | NO | NO | NO | YES | YES | NO
EQPRepairJ22112975 | T6373-143 | CWC7J0523-01 | Setup Fail | TB1HVS20 | PA1333R-25E | RDKZJ9000 | 04 | CONTACT BIN24-PASS, CONFIRMED BY ENGR BUT WHEN RUNCORR SHOW TOO MANY BIN24-NOT PASS, MBD AS PER ENGR. COMMAND | A6513 | Setup Fail | Function Fail | 2022/08/30 18:38:20 | 2022/08/30 19:10:25 | 0.53 | OK | 16631 | 1.run corr全面bin24，更換備卡#21可pass，但依舊有bin24.25.27問題，run card標示bin24.25.27問題客戶指示續測，此批為回測重測bin3~32,55，非全面重測，先不run corr，待觀察重測狀況 | NO | NO | NO | YES | YES | NO
EQPRepairJ22112976 | T6373-129 | CWC8J0G95-01.002 | OverKill | JB1QNS01 | NA01687G(COG48BA)-A-AT-N2-C | GGU012.1-B1C22 | 12 | SHOW I O/S MODE TEST | A6571 | Over Kill | show I | 2022/08/30 18:39:03 | 2022/08/30 18:41:26 | 0.04 | OK | 11385 | EDA I 續測 | NO | NO | NO | YES | YES | NO
EQPRepairJ22112981 | T6373-47 | CWC8J0G48-01 | OverKill | JB1HKS10 | NA01687F(COG48BA)-A-AT-C | GGS941.1-22BC11 | 25 | o/s-20022 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/30 19:01:37 | 2022/08/30 19:27:20 | 0.43 | OK | 15895 | remote----機台清針，+OD5，P/M35.4 | NO | NO | NO | YES | YES | NO
EQPRepairJ22112994 | T6373-34 | CWC8J0957-01 | Corr Fail | TB1YLS39 | PA6485E(CP1-CP3) | P0CSG.1 | 04 | T32 corr. compare failed-A6053 | CPAUTO | Corr | Corr Fail | 2022/08/30 20:22:58 | 2022/08/30 20:31:48 | 0.15 | OK | 17542 | Remote處理 ----- Corr 良率比 Corr Wafer 好，前流程 Fail Bin 回收成 Pass，手動比對 Corr Pass | NO | NO | NO | YES | YES | NO
EQPRepairJ22113023 | T6373-34 | CWC8J0957-01 | OverKill | TB1YLS39 | PA6485E(CP1-CP3) | P0CSG.1 | 05 | CONTINUOUS EDA SHOW 2R O/S-A6881 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/30 22:55:41 | 2022/08/30 23:00:57 | 0.09 | OK | 03426 | Show R 針位調整 觀察 | NO | NO | NO | YES | YES | NO
EQPRepairJ22113034 | T6373-34 | CWC8J0957-01 | OverKill | TB1YLS39 | PA6485E(CP1-CP3) | P0CSG.1 | 06 | too many o/s..bin4-A7025 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/31 00:00:30 | 2022/08/31 00:18:50 | 0.31 | OK | 17542 | 機台清針，起測一段pass穩定 | NO | NO | NO | YES | YES | NO
EQPRepairJ22113064 | T6373-135 | CWC7J0L13-01 | Prober Down | TB1R8S76 | PA5223Q_12U60H-LT | P06UM.1 | 13 | MODEL NOT FOUND | A7164 | Prober | Wafer Edge/Mode Not Found | 2022/08/31 02:23:54 | 2022/08/31 03:01:41 | 0.63 | OK | 15895 | Mode Not Found，手動抓取，check die p/m ok | NO | NO | NO | YES | YES | NO
EQPRepairJ22113069 | T6373-147 | CWC8J0746-01 | OverKill | JB1UJS18 | PA5129LS_12U80H-LT | P07HA.1 | 02 | run corr too many o/s | A7160 | Over Kill | O/S誤宰 | 2022/08/31 02:58:17 | 2022/08/31 03:50:53 | 0.88 | OK | 15895 | corr fail o/s過多，機台清針pass | NO | NO | NO | YES | YES | NO
EQPRepairJ22113091 | T6373-33 | CWC8J0815-01.001 | OverKill | JB1TAS13 | PA5208K-LT | QT7YT | 19 | o/s too many | A6724 | Over Kill | O/S誤宰 | 2022/08/31 05:13:50 | 2022/08/31 05:40:56 | 0.45 | OK | 13597 | 機台清針PASS | NO | NO | NO | YES | YES | NO
EQPRepairJ22113094 | T6373-38 | CMC8J0082-01 | OverKill | JB1GKS10 | PA6543B | J6630 | 16 | CONTINUOUS EDAN SHOW 2H WAIT CORR WAFER-A6881 | CPAUTO | Over Kill | Show H | 2022/08/31 05:33:24 | 2022/08/31 05:37:02 | 0.06 | OK | 03426 | Show H---[CUSTOMER_HOLD_CRITERIA] [HW_BIN60<98%](Bin 61 < 1109EA) 1070(Hold Wafer) -----請產線續測 | NO | NO | NO | YES | YES | NO
EQPRepairJ22113097 | T6373-34 | CWC8J0957-01 | Corr Fail | TB1YLS39 | PA6485E(CP1-CP3) | P0CSG.1 | 07 | Corr Fail | A6881 | Corr | Corr Fail | 2022/08/31 05:43:37 | 2022/08/31 05:49:53 | 0.10 | OK | 17542 | Remote處理 ----- Corr Fail，系統比對異常，手動比對 Corr Pass | NO | NO | NO | YES | YES | NO
EQPRepairJ22113116 | T6373-141 | CWC8J0G48-01.001 | OverKill | JB1HKS09 | NA01687F(COG48BA)-A-AT-C | GGS941.1-22BC2 | 21 | EDA SHOW 2H CORR PASSED-A6963 | CPAUTO | Over Kill | Show H | 2022/08/31 06:55:32 | 2022/08/31 06:56:50 | 0.02 | OK | 03426 | Show H 請產線續測 | NO | NO | NO | YES | YES | NO
EQPRepairJ22113137 | T6373-128 | CWC7J0M32-01.003 | OverKill | JB1UJS16 | PA5129LS_12U80H-LT | P060Q.1 | 19 | TOO MANY O/S, PLEASE CHECK HAVE PARTICLES-A6986 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/31 08:46:50 | 2022/08/31 08:58:58 | 0.20 | OK | 20231 | 1.清針起測一段穩定PASS，pm(115,0) | NO | NO | NO | YES | YES | NO
EQPRepairJ22113138 | T6373-07 | CWC8J0C06-01 | OverKill | JB1TAS22 | PA5208KF-LT | J6CHC | 17 | O/S-20022 | CPAUTO | Over Kill | 機台清針後Pass | 2022/08/31 08:49:24 | 2022/08/31 09:19:14 | 0.50 | OK | 13749 | 3D Clean 清針PASS | NO | NO | NO | YES | YES | NO
EQPRepairJ22113156 | T6373-29 | CMC8J0019-01 | Other | TB1IMS14 | NA01404E-A-AT-X2-C | GFWM86.1-2BCC | 25 | T44 corr compare fail-20022 | CPAUTO | Corr | Corr Fail | 2022/08/31 10:11:45 | 2022/08/31 10:21:56 | 0.17 | OK | 16631 | 1.Corr Fail，比對pass | NO | NO | NO | YES | YES | NO
EQPRepairJ22113157 | T6373-147 | CWC8J0746-01 | OverKill | JB1UJS18 | PA5129LS_12U80H-LT | P07HA.1 | 13 | o/s-14212 | CPAUTO | Over Kill | O/S誤宰 | 2022/08/31 10:12:34 | 2022/08/31 11:05:43 | 0.89 | OK | 08774 | 下卡磨針 35.5 116.0 | NO | NO | NO | YES | YES | NO
EQPRepairJ22113170 | T6373-43 |  | Other | TB1R8S67 | PA5258D-LT | PXKQ9.1 |  | 3D PLS 3Q | 14212 | Other | Other | 2022/08/31 11:21:44 | 2022/08/31 11:31:11 | 0.16 | OK | 20231 | 1.設定ok | NO | NO | NO | YES | YES | NO